Contents

- About this document 	7

- Communication format 	8

  - Binary telegram (CoLa B) 	8

  - ASCII telegram (CoLa A) 	9

  - Variable types 	10

  - Command basics 	10

  - Log in: Required user level 	11

- Workflows 	12

  - Parameterize the scan 	12

  - Set timestamp/data angle 	12

  - Telegrams that are valid for all sensors 	12

- Telegrams 	14

  - Log in 	14

  - Basic Settings 	16

    - Set frequency and angular resolution/measurement

sectors 	17

    - Read for frequency and angular resolution 	28

    - Alignment mode (one layer activation for adjustment) 	33

    - Set scan configuration 	34

    - Activate standby mode 	42

    - Start measurement 	44

    - Stop measurement 	45

    - Autostart measurement 	47

    - Laser Control 	48

    - Activate/deactivate field application 	50

    - Application selection and switching 	51

    - Read Application selection and switching 	52

    - Load factory defaults 	53

    - Load application defaults 	55

    - Change password 	57

    - Check password 	59

    - Reboot device 	60

    - Set contamination measurement settings 	62

    - Set contamination indication settings 	64

    - Read contamination measurement settings 	65

    - Read contamination indication settings 	67

    - Read contamination measurement detailed values 	69

    - Read contamination indication data 	70

    - Read contamination indication result 	71

    - Save parameters permanently 	73

    - Set to run 	75

  - Measurement output telegram 	76

    - Configure the data content for the scan 	77

    - Configure measurement angle of the scandata for

output 	81

    - Read for actual output range 	84

    - Configure the distance scaling of the scandata for the

output 	87

    - Read the distance scaling for the distance values of the scandata 	88

    - Poll one telegram 	90

    - Send data permanently 	92

  - Time stamp 	112

    - Set time stamp 	112

    - Read time stamp and status of the measurement

function 	114

    - Read device time 	116

    - Set NTP (Network Time Protocol) parameters 	117

  - Filter 	128

    - Set particle filter 	129

    - Set mean filter 	130

    - Set n-pulse to 1-pulse filter (Echo filter) 	131

    - Set echo filter 	132

    - Set and read fog filter 	134

    - Set sensitivity fog filter (LMS5xx, LMS/MRS1000) 	137

    - Activate/deactivate "fog filter operating radius active" 	138

    - Execute the "Teach from 90° angle” button 	140

    - Setting of the value "Fog filter operating radius active

up to” 	141

    - Enable/disable digital nearfield filter 	142

    - Set digital nearfield filter sector selection 	144

    - Activate Median filter 	145

    - Activate Edge filter 	146

    - Set Edge filter 	147

    - Set Cubic Area filter 	149

    - Activate Gloss compensation filter 	151

    - Activate long range mode 	153

    - Read status of long range mode 	154

  - Encoder 	156

    - Set increment source 	156

    - Set encoder settings 	158

    - Set encoder resolution 	159

    - Set fixed speed 	161

    - Read speed threshold 	162

    - Read encoder speed 	163

    - Reset encoder values 	165

  - Inputs and Outputs 	166

    - Read state of the ports 	166

    - Read Port Configration of all I/OS (LRS4000 : Ports 3-8) 	169

    - Read state of the inputs 	172

    - Read state of the outputs 	174

    - Receive outputstate by event 	175

    - Set output state 	180

    - Change output 6/3 function 	182

    - Change output 1 function 	183

    - Change output 1 logic state 	185

    - Change output 2 function 	186

    - Change output 2 logic state 	187

    - Set synchronization mode 	189

    - Set synchronization phase 	190

    - Change input 4 function 	191

    - Set debouncing time for input x 	192

    - Read status of external sync signal 	193

  - Status 	195

    - Read contamination status of the LMS 	195

    - Read firmware version 	197

    - Read the device state 	199

    - Status commands for LD-XXX and NAV310 	201

    - Read device order number 	207

    - Read device type 	208

    - Read operating hours 	210

    - Read power on counter 	212

    - Read temperature 	213

    - Set device name 	215

    - Read device name 	216

    - Read angle compensation sine 	218

    - Reset output counter 	219

    - Read heating state 	221

    - FieldEvalLoggingList by Event 	222

    - FieldEvalLoggingList on Request 	225

  - Interfaces 	227

    - Set IP address 	228

    - Read IP address 	230

    - Set Ethernet gateway 	231

    - Read Ethernet gateway 	233

    - Set IP mask 	234

    - Read IP mask 	236

    - Set baud rate for host interface 	238

    - Read baud rate of host interface 	239

    - Set interface type 	241

    - Read interface type 	242

    - Set Host port number 	244

    - Set Host port Command Language (CoLa dialect) 	246

    - Set Motor synchronization 	248

    - Set Phase Shift of Slave for motor synchronization 	249

    - Enable/Disable Front Panel 	251

    - Set function front panel 	254

    - Set front LEDs 	256

    - Set function of LED1 	258

    - Set function of LED2 	259

    - Switch on/off LED1 or LED2 	261

    - Set 7-segment display to specific symbol or number 	263

  - Application 	266

    - Request status change of monitoring fields on event 	267

    - Individual request of monitoring fields to their status

changes - ECR xy 	270

    - Request SOPAS field data structure 	276

    - Request perpendicular distance once 	280

    - Request perpendicular distance continiously on event 	284

    - Request latest field infringement info 	289

    - Request field infringement info continiously on event 	292

    - Fieldset selection method 	295

    - Read active fieldset 	296

    - Write active fieldset 	298

- Diagnostics 	300

  - SOPAS error codes 	300

  - Additional information 	302

- List of tables 	303

- About this document

Please read this chapter carefully before beginning to use the telegram listing.

The document shows how to send telegrams via a terminal program using the SICK protocol CoLa A (ASCII and hexadecimal values, with TCP port 2111 or 2112) or CoLa B (binary/hexadecimal values, with TCP port 2112 only) to the laserscanners LMS1xx, LMS5xx, TiM2xx, TiM5xx (TiM55x, TiM56x, TiM57x), TiM7xx, LMS1000, MRS1000, MRS6000, NAV310, LD-OEM15xx, LD-LRS36xx, LMS4000 and LRS4000. This comprises

the query of the current device state or certain parameter values, how to modify parameter values and the way in which the device confirms or responds to commands/telegrams.

The devices generally support automatic IP address discovery. Default IP address is:

  - LMSxxx: 192.168.0.1

  - TiMxxx: 192.168.0.1

  - MRSxxxx: 192.168.0.1

  - NAV310: 192.168.1.10

  - LD-XXXxxxx: 192.168.1.10

  - LMS4000: 192.168.0.1

  - LRS4000: 192.168.0.1

Subnet mask is 255.255.255.0. IP ports:

  - 2111: CoLa A (fixed) (for LMS4000 fixed CoLa B)

  - 2112: CoLa A (can be switched to CoLa B)

  - 2213: UDP

The document does not or only in a few exceptional cases differentiate between individual device versions or sub product families such as LMS5xx Lite and LMS5xx PRO. Most parameter changes also require certain user levels. Additionally, commands may change during the product lifecycle and development process with a new firmware.

This telegram listing is based on the following firmware statuses (or newer):

  - LMS1xx: V1.80 (V1.21 for LMS12x/13x)

  - LMS5xx: V2.10 (V2.10 for LMS531)

  - TiM2xx: V1.00

  - TiM5xx: V2.51

  - TiM7xx: V2.51

  - MRS1000: 1.0.0 (1.0.0.0R)

  - LMS1000: 1.4.0 (1.4.0.0R)

  - NAV310: V1.03

  - LD-OEM15xx: V1.12 (V1.32 for OEM1500)

  - LD-LRS36xx: V1.12 (V1.32 for LRS3600)

  - LMS4000: V1.5 (1.5.0.0R)

  - LRS4000: V1.1.0 (1.1.0.120C)

If commands do not seem to work, please verify that your device version supports this functionality, that the minimum required user level has been selected and check on updates of this documentation.

NOTE

In case you prefer to use complete drivers instead of single telegrams, the following options are availabe:

C++ drivers: https://github.com/SICKAG/sick_scan_xd  ROS drivers: https://github.com/SICKAG/sick_scan_xd  ROS2 drivers: https://github.com/SICKAG/sick_scan_xd

- Communication format

  - Binary telegram (CoLa B)

The binary telegram is the basic protocol of the scanner (CoLa B ). All values are in hexadecimal code and grouped into pairs of two digits (= 1 byte). The string consists of four parts: header, data length, data and checksum (CS).

The header indicates with 4 × STX (02 02 02 02) the start of the telegram.

The data length defines the size of the data part (command part) by indicating the number of digit pairs in the third part. The size of the data length itself is 4 bytes, which means that the data part might have a maximum of 16 8= 4,294,967,295 digit pairs.

The data part comprises the actual command with letters and characters converted to Hex (according to the ASCII chart) and the parameters of either decimal numbers converted to Hex or fixed Hex values with a specific, intrinsic meaning (no conversion). There is always a blank (20) between the command and the parameters, but not between the different parameter values.

The checksum finally serves to verify that the telegram has been transferred correctly. The length of the checksum is 1 byte, CRC8. It is calculated with XOR.

Example: Binary telegram

| 02 02 02 02 | 00 00 00 17 | 73 4D 4E 20 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 03 F4 72 47 44 | B3 |
| --- | --- | --- | --- |
| Header | Length | Data | CS |

Table 1: Example: Binary telegram

This is an example telegram for setting the user level "Authorized Client ”:

- Header = 02 02 02 02

- Length = 23 digit pairs (17h)

- Data:

  - 73 4D 4E 20 = sMN = start of Sopas command (and blank)

  - 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 = Set Access Mode = the actual command for setting the user level (and blank)

  - 03 = fixed Hex value meaning user level "Authorized Client”

  - F4 72 47 44 = fixed Hex value, serving as password for the selected user level "Authorized Client”

- Checksum = B3 from XOR calculation

  - ASCII telegram (CoLa A)

The ASCII telegram is an alternative to the binary telegram. Due to the variable string length of ASCII telegrams, the Binary telegram is recommended when using scanners with a PLC.

The ASCII telegram has the advantage that commands can be written in plaintext. The string consists only of two parts: the framing and the data part.

The framing indicates with <STX> and <ETX> the start and stop of each telegram.

The data part comprises the actual command with letters and characters (plaintext), parameter values either in decimal (special indicator required) or in hexadecimal (ex- ample: a frequency of 25Hz = +2500 (decimal) = 09C4 (Hex)) and fixed hexadecimal values with a specific, intrinsic meaning. As leading zeros are being deleted, there is always a blank required between all command parts and parameter parts.

NOTE

The device will comfirm parameter values always in hexadecimal code, regardless of the code sent.

As further alternative within CoLa A, depending on the preferences of the user, all values can be written directly in Hex. This means however a 1:1 conversion of all letters and characters including numbers and fixed hexadecimal values via the ASCII chart.

Example: ASCII telegram

| ASCII | <STX> | sMN {SPC} SetAccessMode {SPC} 03 {SPC} F4724744 | <ETX> |
| --- | --- | --- | --- |
| Hex | 02 | 73 4D 4E 20 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 30 33 20 46 34 37 32 34 37 34 34 | 03 |
|  | Start | Data | Stop |

Table 2: Example: ASCII telegram

This is again an example telegram for setting the user level "Authorized Client”. As only fixed hexadecimal parameter values are needed, the option to use parameter values in decimal code with special indicator cannot be applied here:

- Framing = <STX> = telegram start = 02 (Hex)

- Data:

  - sMN = start of Sopas command (and blank) = 73 4D 4E 20 (Hex)

  - SetAccessMode = the actual command for setting the user level (and blank) = 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 (Hex)

  - 03 = fixed Hex value meaning user level "Authorized Client” (and blank) = 30 33 20 (Hex)

  - F4 72 47 44 = fixed Hex value, serving as password for the selected user level "Authorized Client” = 46 34 37 32 34 37 34 34 (Hex)

- Framing = <ETX> = telegram stop = 03 (Hex)

  - Variable types

| Variable type | Length (byte) | Value range | Sign |
| --- | --- | --- | --- |
| Bool_1 | 1 | 0 or 1 | No |
| Uint_8 | 1 | 0... 255 | No |
| Int_8 | 1 | -128... +127 | Yes |
| Uint_16 | 2 | 0... 65,535 | No |
| Int_16 | 2 | -32,768... +32,767 | Yes |
| Uint_32 | 4 | 0... 4,294,967,295 | No |
| Int_32 | 4 | -2,147,483,648... +2,147,483,647 | Yes |
| Enum_8 | 1 | Certain values defined in a list of Choices (0 &hellip; 255) | No |
| Enum_16 | 2 | Certain values defined in a list of Choices (0 &hellip; 65535) | No |
| String | Context- dependent | Strings are not terminated in zeroes |  |
| Real |  | Float nach IEEE754 (see www.h- schmidt.net/FloatConverter/IEEE754d e.html) |  |

Data length is always given in Bytes!

  - Command basics

| Description | Value ASCII | Value Hex | Value Binary |
| --- | --- | --- | --- |
| Start of text | <STX> | 02 | 02 02 02 02 + given length |
| End of text | <ETX> | 03 | Calculated checksum |
| Read | sRN | 73 52 4E |  |
| Write | sWN | 73 57 4E |  |
| Method | sMN | 73 4D 4E |  |
| Event | sEN | 73 45 4E |  |
| Answer | sRA | 73 52 41 |  |
|  | sWA | 73 57 41 |  |
|  | sAN | 73 41 4E |  |
|  | sEA | 73 45 41 |  |
|  | sSN | 73 53 4E |  |
| Space | {SPC} | 20 | 20 |

If values are divided into two parts (e.g. measurement data), they are documented according to LSB 0 (e.g. 00 07), output however is according to MSB (e.g. 07 00).

  - Log in: Required user level

| Task | Required user level |
| --- | --- |
| Change sensor parameters | Authorized Client |
| Requests or queries (e.g. for measurement data or device state) | None |
| Manage password | Service |

In general, every sWN command for changing paramters requires to log in to the device first (see chapter 4.1 ). When being logged in, any desired parameter valid for this user level can be changed. All changes become active only after having logged off again from the device via the sMN Run command (see chapter 4.2.26 ).

In this document, a required, specific user level is indicated in the telegram structure head line.

- Workflows

  - Parameterize the scan

    - Log in: sMN SetAccessMode (see 4.1, page 14)

    - Set frequency and resolution: sMN mLMPsetscancfg (see 4.2.1, page 17)

    - Configure scandata content: sWN LMDscandatacfg (see 4.3.1, page 77)

    - Configure scandata output: sWN LMPoutputRange (see 4.3.2 page 81)

    - Store parameters: sMN mEEwriteall (see 0, page 72)

    - Log out: sMN Run (see 4.2.26, page 75)

    - Request scan:

sRN LMDscandata (see 4.3.4, page 87)

sEN LMDscandata (see 4.3.7, page 92) (Device output &hellip;)

More detailed command descriptions can be found in the course of this document.

Example: Sequence for LD-OEM1501, NAV310, LD-LR3601, LD-LR3611 to configure 2 sectors and get measurement scans

Sector configuration: 	Resolution: 10Hz; 0,125°;

Sector 1: 0°&hellip; 44 °(0h... 6 B6C0h);

Sector 2: 45°&hellip; 180 ° (6 DDD6h &hellip; 1 B7740h)

- Stop measurement: sMN LMCstopmeas sAN LMCstopmeas 0

- Log in: sMN SetAccessMode (see 4.1, page 14)

3 	Set Sectors : LCMstate001B7740 04E2 000000 0000000 04E2 000000 000000

sAN mLMPsetscancfg 0 3E8 2 4E2 0 6B6C0 4E2 6DDD6 1B7740 4E2 0 0 4E2 0 0

- Store parameters: sMN mEEwriteall (see 0, page 72)

- Log out: sMN Run (see 4.2.26, page 75)

- Start Measurement: sMN LMCstartmeas sAN LMCstartmeas 0

- Request scan:

sRN LMDscandata (see 4.3.4, page 87)

sEN LMDscandata (see 4.3.7, page 92) (Device output &hellip;)

  - Set timestamp/data angle

- Log in: sMN SetAccessMode (see 4.1, page 14)

- Sopas command: sMN LSPsetdatetime (see 4.4.1, page 112)

- Log out: sMN Run (see 4.2.26, page 75)

  - Telegrams that are valid for all sensors

Here you find list of telegrams that can be used with every of the above mentioned LiDAR sensor families. Entries marked with a asterisk (*) are valid for almost all sensors. For a detailed description, please refer directly to chapter 4.

| 4.1 | Log in |
| --- | --- |
| 4.2.2 | Read for frequency and angular resolution |
| 4.2.6 | Start measurement * |
| 4.2.7 | Stop measurement * |

| 4.2.13 | Load factory defaults * |
| --- | --- |
| 4.2.14 | Load application defaults * |
| 4.2.16 | Check password * |
| 4.2.17 | Reboot device * |
| 4.2.25 | Save parameters permanently |
| 4.2.26 | Set to run |
| 4.3.1 | Configure the data content for the scan |
| 4.3.2 | Configure measurement angle of the scandata for output * |
| 4.3.3 | Read for actual output range * |
| 4.3.6 | Poll one telegram |
| 4.3.7 | Send data permanently |
| 4.5.1 | Set particle filter * |
| 4.7.3 | Read state of the inputs |
| 4.7.4 | Read state of the outputs |
| 4.7.5 | Receive outputstate by event |
| 4.7.6 | Set output state * |
| 4.8.2 | Read firmware version |
| 4.8.3 | Read the device state |
| 4.8.5 | Read device order number |
| 4.8.6 | Read device type |
| 4.8.7 | Read operating hours * |
| 4.8.8 | Read power on counter * |
| 4.8.10 | Set device name * |
| 4.8.11 | Read device name * |
| 4.8.13 | Reset output counter * |
| 4.9.1 | Set IP address |
| 4.9.2 | Read IP address |
| 4.9.3 | Set Ethernet gateway |
| 4.9.4 | Read Ethernet gateway |
| 4.9.5 | Set IP mask |
| 4.9.6 | Read IP mask |
| 4.9.11 | Set Host port number * |
| 4.9.12 | Set Host port Command Language (CoLa dialect) * |

- Telegrams

  - Log in

NOTES

 Pl ease not e that f or TiMx xx an d LMS 40 00th e laser I s shut off a fter a succ esful log in and no measurement data is created any more. The laser is turned on again after log out (sMN Run).

 An automated hash-value calculator can be found in SOPAS ET under menu "password”. Required userlevel "Service ”.

| Telegram structure: sMN SetAccessMode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | User level | String | 13 | All | SetAccessMode | 53 65 74 41 63 63 65 73 73 4D 6F 64 65 |
| User level | Select user level | Int_8 | 1 | All | Maintenance: 02 Authorized client: 03 Service: 04 | Maintenance: 02 Authorized client: 03 Service: 04 |
| Password | Hash value for the selected user level | Uint_32 | 4 | All | Maintenance: B21ACE26 Authorized client: F4724744 Service: 81BE23AA | Maintenance: B2 1A CE 26 Authorized client: F4 72 47 44 Service: 81 BE 23 AA |

Table 3: Telegram structure: sMN SetAccessMode

Example: sMN SetAccessMode

Log in as "Authorized client” with password "F4724744 ”.

| CoLa A | ASCII | <STX> sMN {SPC} SetAccessMode {SPC} 03 {SPC} F4724744 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 30 33 20 46 34 37 32 34 37 34 34 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 4D 4E 20 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 03 F4 72 47 44 B3 |

Table 4: Example: sMN SetAccessMode

| Telegram structure: sAN SetAccessMode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | User level | String | 13 | All | SetAccessMode | 53 65 74 41 63 63 65 73 73 4D 6F 64 65 |
| Change user level | Changed level | Bool_1 | 1 | All | Error: 0 Success: 1 | Error: 00 Success: 01 |

Table 5: Telegram structure: sAN SetAccessMode

Example for LMS100: sAN SetAccessMode

| CoLa A | ASCII | <STX> sAN {SPC} SetAccessMode {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 41 4E 20 53 65 74 41 63 63 65 73 73 4D 6F 64 65 20 01 38 |

Table 6: Example for LMS100: sAN SetAccessMode

  - Basic Settings

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.2.1 | Set frequency and angular resolution/measurement sectors | x | x |  |  |  | x | x | x |  |  |  |  |  |
| 4.2.2 | Read for frequency and angular resolution | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.2.3 | Alignment mode (one layer activation for adjustment) |  |  |  |  |  |  |  |  | x |  |  |  |  |
| 4.2.4 | Set scan configuration |  | x |  |  |  | x | x | x |  |  |  |  | x |
| 4.2.5 | Activate standby mode | x | x |  |  |  |  |  |  | x | x |  | x | x |
| 4.2.6 | Start measurement | x | x |  | x | x |  |  |  | x | x | x | x | x |
| 4.2.7 | Stop measurement | x | x |  | x | x |  |  |  | x | x | x | x |  |
| 4.2.8 | Autostart measurement | x | x |  |  |  |  | x | x |  |  | x |  |  |
| 4.2.9 | Laser Control |  |  |  |  |  |  |  |  |  |  |  | x |  |
| 4.2.10 | Activate/deactivate field application |  |  |  |  |  |  | x | x |  |  |  |  |  |
| 4.2.11 | Application selection and switching |  |  |  |  |  |  |  |  | x | x |  |  |  |
| 4.2.12 | Read Application selection and switching |  |  |  |  |  |  |  |  | x | x |  |  |  |
| 4.2.13 | Load factory defaults | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.2.14 | Load application defaults | x | x |  | x | x |  |  |  | x | x | x | x | x |
| 4.2.15 | Change password | x | x |  |  |  |  |  |  |  |  | x | x | x |
| 4.2.16 | Check password | x | x | x | x | x |  |  |  |  |  | x | x | x |
| 4.2.17 | Reboot device | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.2.18 | Set contamination measurement settings | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.2.19 | Set contamination indication settings | x | x |  |  |  |  |  |  | x | x |  |  |  |
| 4.2.20 | Read contamination measurement settings | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.2.21 | Read contamination indication settings |  |  |  |  |  |  |  |  | x | x |  |  |  |
| 4.2.22 | Read contamination measurement detailed values | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.2.23 | Read contamination indication data |  |  |  |  |  |  |  |  | x | x |  |  |  |
| 4.2.24 | Read contamination indication result |  |  |  |  |  |  |  |  | x | x |  |  |  |
| 4.2.25 | Save parameters permanently | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.2.26 | Set to run | x | x | x | x | x | x | x | x | x | x | x | x | x |

    - Set frequency and angular resolution/measurement sectors

NOTES

 	Please note that the new values will be activated only after log out (from the user level), when re-entering the Run mode (see Table 124 on page 75 ).

Coordination system:

LMS5xx (-5° to 185°)  Start angle and stop angle are fixed values and not changeable for LMS5xx, only in the data output!This also applies for LMS1xx series.

LMS1xx and TiMxxx (-45° to 225°)

MRS1000 (-47,5° to 227,5°) and LMS1000 (-48° to 228°)

LRS4000 (+180° to - 180°)

The LD series is available in two versions having a different rotation direction and coordinate system:

LD-OEM1501, NAV310, LD-LR36X1 (0° to 360°)

LD-OEM1500 and LD-LR36X0 (-90° to +270°)

For sending the sector configuration there follow these rules:

 	Send the sectors in their ascending sequence.

 	For LD and NAV products: Send allways the definition for all sectors (unused sector as "{SPC} 0 {SPC} 0 ".)

 	For LMS products: They have only one measurement sector, send only the first one and leave the rest away.

For more details on sector configuration see examples below. For complete workflow see example in section 3, page 12.

| Telegram structure: sMN mLMPsetscancfg (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Configuration of scan frequency and angular resolution | String | 14 | All | mLMPsetscancfg | 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 |
| Scan frequency | [1/100 Hz] | Uint_32 | 4 | LMS1xx | 25Hz: +2500d (9C4h) 50Hz: +5000d (1388h) | 25Hz: 00 00 09 C4 50Hz: 00 00 13 88 |
|  |  |  |  | LMS5xx | 25Hz: +2500d (9C4h) | 25Hz: 00 00 09 C4 |
|  |  |  |  |  | 35Hz: +3500d (DACh) | 35Hz: 00 00 0D AC |
|  |  |  |  |  | 50Hz: +5000d (1388h) | 50Hz: 00 00 13 88 |
|  |  |  |  |  | 75Hz: +7500d (1A0Bh) | 75Hz: 00 00 1A 0B |
|  |  |  |  |  | 100Hz: +10000d (2710h) | 100Hz: 00 00 27 10 |
|  |  |  |  | MRS1000 | 50Hz: +5000d (1388h) | 50Hz: 00 00 13 88 |
|  |  |  |  |  | 25Hz: +2500d (9C4h) | 25Hz: 00 00 09 C4 |
|  |  |  |  |  | 12,5 Hz: +1250d (4E2h) | 12,5Hz: 00 00 04 E2 |

| Telegram structure: sMN mLMPsetscancfg (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  | LMS1000 | 150Hz: +15000d (3A98h) 75Hz: +7500d (1D4Ch) 37,5 Hz: +3750d (0EA6h) | 150Hz: 00 00 3A 98 75Hz: 00 00 1D 4C 37,5 Hz: 00 00 0E A6 |
|  |  |  |  | NAV310 LD-OEM15xx | 5Hz... 20Hz: 500d... 2000d (1 F4h... 7 D0h) | 5Hz... 20Hz: 00 00 01 F4... 00 00 07 D0 |
|  |  |  |  | LD-LRS36xx | 5Hz... 15Hz: + 500d... + 1500d (1 F4h &hellip; 5DCh) | 5Hz... 15Hz: 00 00 01 F4... 00 00 05 DC |
| Number of active sectors | Indicates the number of active sectors (e. g. NAV310 with 2 active sectors out of available 4) | Int_16 | 2 | LMS1xx LMS5xx LMS1000 MRS1000 | +1 (0001h) | 0001 |
|  |  |  |  | NAV310 LD-OEM15xx LD-LRS36xx | +1... +4 (0001... 0004h) | 0001... 0100 (binary) |

| Telegram structure: sMN mLMPsetscancfg (Authorized client) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | Angular resolution | [1/10000 °] | Uint_32 4 |  | LMS1xx | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  | Same value for each |  |  |  | 0.5: + 5000d (1388h) | 0.5: 00 00 13 88 |
|  |  | sector required. |  |  |  |  |  |
|  |  |  |  |  | LMS5xx | 0.042°: +0417 d (01A1h) | 0,042: 00 00 01 A1 |
|  |  |  |  |  |  | 0.083°: +0833 d (0341h) | 0,083: 00 00 03 41 |
|  |  |  |  |  |  | 0.1667: + 1667d (683h) | 0.1667: 00 00 06 83 |
|  |  |  |  |  |  | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  |  |  |  |  | 0.333: + 3333d (D05h) | 0.333: 00 00 0D 05 |
|  |  |  |  |  |  | 0.5: + 5000d (1388h) | 0.5: 00 00 13 88 |
|  |  |  |  |  |  | 0.667: + 6667d (1 A0Bh) | 0.667: 00 00 1A 0B |
|  |  |  |  |  |  | 1: + 10000d (2710h) | 1: 00 00 27 10 |
|  |  |  |  |  | MRS1000 | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  |  |  |  |  | 0.125: + 1250d (4 E2h) | 0,125: 00 00 04 E2 |
|  |  |  |  |  |  | 0.0625: + 0625d (271h) | 0,0625: 00 00 02 71 |
|  |  |  |  |  | LMS1000 | 0.75: + 7500d (1 D4Ch) | 0.75: 00 00 1D 4C |
|  |  |  |  |  |  | 0.375: + 3750d (0 EA6h) | 0.375: 00 00 0E A6 |
|  |  |  |  |  |  | 0.1875: + 1875d (0753h) | 0.1875: 00 00 07 53 |
|  |  |  |  |  | NAV310 | 0.125 °&hellip; 1 °: | 0.125 °&hellip; 1 °: |
|  |  |  |  |  | LD-OEM15xx | + 1250d °&hellip; + 10000d | 00 00 04 E2... |
|  |  |  |  |  | LD-LRS36xx | (4E2h °&hellip; 2710h) | 00 00 27 10 |
|  | Start angle | [1/10000 °] | Int_32 | 4 | LMS1xx | -450000d (FFF92230h) | FF F9 22 30 |
|  |  | Value for start angle |  |  | LMS5xx | -50000d (FFFF3CB0h) | FF FF 3C B0 |
|  |  | musst always be |  |  |  |  |  |
|  |  |  |  |  | MRS1000 | -475000d (FFF8C088h) | FF F8 C0 88 |
|  |  | greater than Stop |  |  |  |  |  |
|  |  | angle of previous |  |  |  |  |  |
|  |  |  |  |  | LMS1000 | -480000d (FFF8AD00h) | FF F8 AD 00 |
|  |  | sector. Set to 0 if sector is inactive (not used). |  |  |  |  |  |
|  |  |  |  |  | NAV310 LD-OEM15x1 LD-LRS36x1 | 0... + 3600000d (0h... 36 EE80h) | 00 00 00 00 00... 00 36 EE 80 |
|  |  | Values for LMSxxx are fixed. |  |  |  |  |  |
|  |  |  |  |  | LD-OEM15x0 LD-LRS36x0 | -900000d... + 2700000d (FFF24460h &hellip; 41 EB0h) | FF F2 44 60&hellip; 00 04 1E B0 |
|  | Stop angle | [1/10000 °] | Int_32 | 4 | LMS1xx | +2250000d (225510h) | 00 22 55 10 |
|  |  | Value for stop angle |  |  | LMS5xx | + 1850000d (1 C3A90h) | 00 1C 3A 90 |
|  |  | musst always be |  |  |  |  |  |
|  |  |  |  |  | MRS1000 | + 2275000d (22 B6B8h) | 00 22 B6 B8 |
|  |  | greater than start |  |  |  |  |  |
|  |  | angle of previous |  |  |  |  |  |
|  |  |  |  |  | LMS1000 | + 2280000d (22 CA40h) | 00 22 CA 40 |
|  |  | sector. Set to 0 if sector is inactive (not used). |  |  |  |  |  |
|  |  |  |  |  | NAV310 LD-OEM15x1 LD-LRS36x1 | 0... + 3600000d (0h &hellip; 36 EE80h) | 00 00 00 00 00... 00 36 EE 80 |
|  |  | Values for LMSxxx are fixed. |  |  |  |  |  |
|  |  |  |  |  | LD-OEM15x0 LD-LRS36x0 | -900000d... + 2700000d (FFF24460h &hellip; 41 EB0h) | FF F2 44 60&hellip; 00 04 1E B0 |

Table 7: Telegram structure: sMN mLMPsetscancfg

Remark for measurement data output with LMS5xx at 0,083° and 0,042°:

With angular resolution of 0,083° or 0,042°, it is recommended to compensate the systematic error of the angular position. To calibrate the LMS5 in the system a correction table should be established in the customer's system software.

This table ensures the mapping of an indicated angle in the scan data telegram on an angle in the real system environment.

The calibration could be done, e.g. with a test object or with a reference scan on the system structure with known geometry. The table should comprise the complete field of view which is relevant for the application.

Example for LMS1xx

1 measurement sector

of 270°

Example for LMS1xx

ATTENTION: Scan angle can not be changed here, only in the data output!This applies for LMS1xx and LMS5xx series.

Scan frequency = 50Hz

Sectors = 1 sector (This value is always 1 for these devices) Angular resolution = 0, 5°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} +5000 {SPC} +1 {SPC} +5000 {SPC} -450000 {SPC} +2250000 <ETX> |
| --- | --- | --- |
|  |  | Alternatively: |
|  |  | <STX> sMN {SPC} mLMPsetscancfg {SPC} 1388 {SPC} 1 {SPC} 1388 {SPC} FFF92230 {SPC} 225510 <ETX> |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 2B 35 30 30 20 2B 31 20 2B 35 30 30 30 20 2D 34 35 30 30 30 20 2B 32 32 35 30 30 30 30 03 |
|  |  | Alternatively: |
|  |  | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 31 33 38 38 20 31 20 31 33 38 38 20 |
|  |  | 46 46 4639 32 32 33 30 20 32 32 35 35 31 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 25 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 13 88 00 01 00 00 13 88 FF F9 22 30 00 22 55 10 21 |

Start angle of sector = -45°(Fix values, angle not changeable) Stop angle of sector = 225° (Fix values, angle not changeable)

Table 8: Example: sMN mLMPsetscancfg for LMS1xx with 1 measurement sector of 270°

Example for LD-xxx###1

1 measurement sector

of 360°

Examples for LD-OEM1501, NAV310, LD-LR36X1

Scan frequency = 8Hz Sectors = 1 sector

Angular resolution = 0,25° Start angle of sector = 0° Stop angle of sector = 360°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} 0320 {SPC} 01 {SPC} 09C4 {SPC} 0 {SPC} 0036EE80 {SPC} 09C4 {SPC} 0 {SPC} 0 {SPC} 09C4 {SPC} 0 {SPC} 0 {SPC} 09C4 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 30 33 32 30 20 30 31 20 30 39 43 34 20 30 20 30 30 33 36 45 45 38 30 20 30 39 43 34 20 30 20 30 20 30 39 43 34 20 30 20 30 20 30 39 43 34 20 30 20 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 55 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 03 20 00 01 00 00 09 C4 |
|  |  | 00 00 00 00 00 36 EE 80 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 09 C4 00 |
|  |  | 00 00 00 00 00 00 00 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 00 00 E4 |

Table 9: Example: sMN mLMPsetscancfg for LD-XXX###1 with 1 measurement sector of 360°

Example for LD-XXX###1

1 measurement sector

of 270°

Scan frequency = 10Hz Sectors = 1 sector

Angular resolution = 0,50° Start angle of sector = +45° Stop angle of sector = +315°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} +1000 {SPC} +1 {SPC} +5000 {SPC} +450000 {SPC} +315000 {SPC} +5000 { |
| --- | --- | --- |
|  |  | SPC} 0 {SPC} 0 {SPC} +5000 {SPC} 0 {SPC} 0 {SPC} +5000 {SPC} 0 {SPC} 0 <ETX> |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 2B 31 30 30 20 2B 31 20 2B 35 30 30 2B 34 |
|  |  | 35 30 30 30 30 20 2B 33 31 35 30 30 20 2B 35 30 30 20 30 30 30 30 30 30 30 30 20 30 30 30 30 30 30 20 |
|  |  | 2B 35 30 30 30 30 30 30 30 20 30 30 30 30 30 30 30 20 2B 35 30 30 20 30 30 30 30 30 30 20 30 30 30 |
|  |  | 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 49 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 03 E8 00 |
|  |  | 01 00 00 13 88 00 06 DD DE 00 30 10 B0 00 00 13 88 00 00 00 00 00 00 00 00 00 00 13 88 00 00 00 00 00 00 00 00 |
|  |  | 00 00 13 88 00 00 00 00 00 00 00 00 2C |

Table 10: Example: sMN mLMPsetscancfg for LD-XXX###1 with 1 measurement sector of 270°

Example for LD-xxx###1

2 measurement sectors

Scan frequency = 8Hz Sectors = 2 sectors

Sector 1 = +70°&hellip; +90 °

Sector 2 = +120°&hellip; +150° Angular resolution = 0,25°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} 0320 {SPC} 02 {SPC} 09C4 {SPC} +700000 {SPC} +900000 {SPC} 09C4 {SPC } +1200000 {SPC} +1500000 {SPC} 09C4 {SPC} 0{ SPC} 0 {SPC} 09C4 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 30 33 32 30 20 30 32 2 0 30 39 43 34 20 2B 37 30 30 30 30 20 2B 39 30 30 30 30 20 30 39 43 34 20 2B 31 32 30 30 30 30 30 20 2B 31 35 30 30 30 30 30 30 39 43 34 20 30 20 30 20 30 39 43 34 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 49 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 00 00 03 20 00 02 00 00 09 C4 00 0A AE 60 00 0D BB A0 00 00 09 C4 00 12 4F 80 00 16 E3 60 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 00 00 E8 |

Table 11: Example: sMN mLMPsetscancfg for LD-XXX###1 with 2 measurement sectors

Example for LD-xxx###1

4 measurement sectors

Scan frequency = 8Hz Sectors = 4 sectors

Sector 1 = +320°&hellip; +45

Sector 2 = +70°&hellip; +90 °

Sector 3 = +120°&hellip; +150 °

Sector 4 = +210°&hellip; +240° Angular resolution = 0,25°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} 0320 {SPC} 04 {SPC} 09C4 {SPC} +3200000 {SPC} +450000 {SPC} 09C4 {SP C} +700000 {SPC} +900000 {SPC} 09C4 {SPC} +1200000 {SPC} +1500000 {SPC} 09C4 {SPC} +2100000 {SPC} +240 0000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 30 33 32 30 20 30 34 20 30 39 43 34 20 2B 33 32 30 30 30 30 20 2B 34 35 30 30 20 30 39 43 34 20 2B 37 30 30 30 30 20 2B 39 30 30 30 30 30 20 30 39 43 34 20 2B 31 32 30 30 30 20 2B 31 35 30 30 30 30 30 30 39 20 43 34 20 2B 32 31 30 30 30 20 2B 32 34 30 30 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 49 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 03 20 00 04 00 00 09 C4 00 30 D4 00 00 06 DD D0 00 00 09 C4 00 0A AE 60 00 0D BB A0 00 00 09 C4 00 12 4F 80 00 16 E3 60 00 00 09 C4 00 20 0B 20 00 24 9F 00 B1 |

Table 12: Example: sMN mLMPsetscancfg for LD-XXX###1 with 4 measurement sectors

Example for LD-xxx###0

1 measurement sector

of 360°

Examples for LD-OEM1500 and LD-LR36X0

Scan frequency = 8Hz Sectors = 1 sector

Angular resolution = 0,25° Start angle of sector = -90° Stop angle of sector = +270°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} 0320 {SPC} 01 {SPC} 09C4 {SPC} ­900000 {SPC} +2700000 {SPC} 09C4 {SPC } 00000000 {SPC} 000000 {SPC} 09C4 {SPC} 000000 {SPC} 0000000 {SPC} 09C4 {SPC} 000000 {SPC} 000000 <ETX > |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 30 33 32 30 20 30 31 20 30 39 43 34 20 2D 39 30 30 30 30 20 2B 32 37 30 30 30 30 20 30 39 43 34 20 30 30 30 30 30 30 30 30 20 30 30 30 30 30 30 20 30 39 43 34 20 30 30 30 30 30 30 20 30 30 30 30 30 30 30 20 30 39 43 34 20 30 30 30 30 30 30 20 30 30 30 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 49 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 03 20 00 01 00 00 09 C4 FF F2 44 60 00 29 32 E0 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 00 00 00 00 A3. |

Table 13: Example: sMN mLMPsetscancfg for LD-XXX###0 with 1 measurement sector of 360°

Example for LD-xxx###0

1 measurement sector

of 270°

Scan frequency = 10Hz Sectors = 1 sector

Angular resolution = 0,50° Start angle of sector = -45° Stop angle of sector = +225°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} +1000 {SPC} +1 {SPC} +5000 {SPC} ­450000 {SPC} +225000 {SPC} +5000 { SPC} 0 {SPC} 0 {SPC} +5000 {SPC} 0 {SPC}0{SPC} +5000 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 2B 31 30 30 20 2B 31 20 2B 35 30 30 30 20 2D 34 35 30 30 20 2B 32 32 35 30 30 30 20 2B 35 30 30 20 2D 34 35 30 30 30 30 20 2B 32 32 35 30 30 20 2B 35 30 30 20 30 20 30 20 2B 35 30 30 20 30 20 30 20 2B 35 30 30 30 20 30 20 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 49 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 03 E8 00 01 00 00 13 88 FF F9 22 30 00 22 55 10 00 00 13 88 00 00 00 00 00 00 00 00 00 00 13 88 00 00 00 00 00 00 00 00 00 00 13 88 00 00 00 00 00 00 00 00 CA |

Table 14: Example: sMN mLMPsetscancfg for LD-XXX###0 with 1 measurement sector of 270°

Example for LD-xxx###0

2 measurement sectors

Scan frequency = 8Hz Sectors = 2 sectors

Sector 1 = +120°&hellip; +150 °

Sector 2 = +180°&hellip; +200° Angular resolution = 0,25°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} 320 {SPC} 2 {SPC} 9C4 {SPC} +1200000 {SPC} +1500000 {SPC} 9C4 {SPC} +1800000 {SPC} +2000000 {SPC} 9C4 {SPC} 0 {SPC} 0 {SPC} 9C4 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 33 32 30 20 32 20 39 43 34 20 2B 31 32 30 30 30 30 30 20 2B 31 35 30 30 30 30 20 39 43 34 20 2B 31 38 30 30 30 20 2B 32 30 30 30 30 30 30 20 39 43 34 20 30 20 30 20 39 43 34 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 49 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 03 20 00 02 00 09 C4 00 12 4F 80 00 16 E3 60 00 00 09 C4 00 1B 77 40 00 1E 84 80 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 00 09 C4 00 00 00 00 00 00 00 00 00 00 00 00 00 0C |

Table 15: Example: sMN mLMPsetscancfg for LD-XXX###0 with 2 measurement sectors

Example for LD-xxx###0

4 measurement sectors

Scan frequency = 8Hz Sectors = 4 sectors

Sector 1 = +230°&hellip; -50 °

Sector 2 = +30°&hellip; +60 °

Sector 3 = +120°&hellip; +150 °

Sector 4 = +210°&hellip; +200° Angular resolution = 0,25°

| CoLa A | ASCII | <STX> sMN {SPC} mLMPsetscancfg {SPC} 320 {SPC} 4 {SPC} 9C4 {SPC} +2300000 {SPC} ­500000 {SPC} 9C4 {SPC} +30 0000 {SPC} +600000 {SPC} 9C4 {SPC} +1200000 {SPC} +1500000 {SPC} 9C4{ SPC} +1800000 {SPC} +2000000 <ET X> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 33 32 30 20 34 20 39 43 34 20 2B 32 33 30 30 30 30 30 20 2D 35 30 30 30 30 20 39 43 34 20 2B 33 30 30 30 20 2B 36 30 30 30 30 20 39 43 34 20 2B 31 32 30 30 30 20 2B 31 35 30 30 30 30 30 20 39 43 34 20 2B 31 38 30 30 30 30 30 20 2B 32 30 30 30 30 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 49 73 4D 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 03 20 00 04 00 00 09 C4 00 23 18 60 FF F8 5E E0 00 00 09 C4 00 04 93 E0 00 09 27 C0 00 00 09 C4 00 12 4F 80 00 16 E3 60 00 00 09 C4 00 1B 77 40 00 1E 84 80 71 |

Table 16: Example: sMN mLMPsetscancfg for LD-XXX###0 with 4 measurement sectors

| Telegram structure: sAN mLMPsetscancfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Info of scan frequency and angular resolution | String | 14 | All | mLMPsetscancfg | 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 |
| Status code | Accepted when value is 0 | Enum_8 | 1 | All | No error: 0 Frequency error: 1 Resolution error: 2 Resolution and scanarea error: 3 Scanarea error: 4 Other errors: 5 | No error: 00 Frequency error: 01 Resolution error: 02 Resolution and scan area error: 03 Scanarea error: 04 Other errors: 05 |
| Scan frequency | [1/100 Hz] | Uint_32 | 4 | LMS1xx | 25Hz: +2500d (9C4h) 50Hz: +5000d (1388h) | 25Hz: 00 00 09 C4 50Hz: 00 00 13 88 |
|  |  |  |  | LMS5xx | 25Hz: +2500d (9C4h) 35Hz: +3500d (DACh) 50Hz: +5000d (1388h) 75Hz: +7500d (1A0Bh) 100Hz: +10000d (2710h) | 25Hz: 00 00 09 C4 35Hz: 00 00 0D AC 50Hz: 00 00 13 88 75Hz: 00 00 1A 0B 100Hz: 00 00 27 10 |
|  |  |  |  | NAV310 LD-OEM 15xx | 5Hz... 20Hz: + 500d... + 2000d (1 F4h... 7 D0h) | 5Hz... 20Hz: 00 00 01 F4... 00 00 07 D0 |
|  |  |  |  | LD-LRS 36xx | 5Hz... 15Hz: + 500d... + 1500d (1 F4h &hellip; 5DCh) | 5Hz... 15Hz: 00 00 01 F4... 00 00 05 DC |
| Number of active sectors | Indicates the number of active sectors | Int_16 | 2 | LMS1xx LMS5xx | 1 (0001h) | 0001 |
|  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | 1... 4 (0001h... 0004h) | 0001... 0100 (binary) |

| Telegram structure: sAN mLMPsetscancfg |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Per sector (active and inactive sectors) | Angular resolution | [1/10000 °] | Uint_32 | 4 | LMS1xx | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  |  |  |  |  | 0.5: + 5000d (1388h) | 0.5: 00 00 13 88 |
|  |  |  |  |  | LMS5xx | 0.042: + 0417d (1 A1h) | 0.042: 00 00 01 A1 |
|  |  |  |  |  |  | 0.083: + 0833d (341h) | 0.083: 00 00 03 41 |
|  |  |  |  |  |  | 0.1667: + 1667d (683h) | 0.1667: 00 00 06 83 |
|  |  |  |  |  |  | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  |  |  |  |  | 0.333: + 3333d (D05h) | 0.333: 00 00 0D 05 |
|  |  |  |  |  |  | 0.5: + 5000d (1388h) | 0.5: 00 00 13 88 |
|  |  |  |  |  |  | 0.667: + 6667d (1 A0Bh) | 0.667: 00 00 1A 0B |
|  |  |  |  |  |  | 1: + 10000d (2710h) | 1: 00 00 27 10 |
|  |  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | 0.125 °&hellip; 1 °: + 1250d °&hellip; + 10000d (4E2h °&hellip; 2710h) | 0.125 °&hellip; 1: 00 00 04 E2... 00 00 27 10 |
|  | Start angle | [1/10000 °] | Int_32 | 4 | LMS1xx | -450000d (FFF92230h) | FF F9 22 30 |
|  |  |  |  |  | LMS5xx | -50000d (FFFF3CB0h) | FF FF 3C B0 |
|  |  |  |  |  | NAV310 LD-OEM 15x1 LD-LRS 36x1 | 0... + 3600000d (0h... 36 EE80h) | 00 00 00 00 00... 00 36 EE 80 |
|  |  |  |  |  | LD-OEM 15x0 LD-LRS 36x0 | -900000d... + 2700000d (FFF24460h &hellip; 41 EB0h) | FF F2 44 60&hellip; 00 04 1E B0 |
|  | Stop angle | [1/10000 °] | Int_32 | 4 | LMS1xx | +2250000d (225510h) | 00 22 55 10 |
|  |  |  |  |  | LMS5xx | + 1850000d (1 C3A90h) | 00 1C 3A 90 |
|  |  |  |  |  | NAV310 LD-OEM 15x1 LD-LRS 36x1 | 0... + 3600000d (0h &hellip; 36 EE80h) | 00 00 00 00 00... 00 36 EE 80 |
|  |  |  |  |  | LD-OEM 15x0 LD-LRS 36x0 | -900000d... + 2700000d (FFF24460h &hellip; 41 EB0h) | FF F2 44 60&hellip; 00 04 1E B0 |

Table 17: Telegram structure: sAN mLMPsetscancfg

Example: sAN mLMPsetscancfg

| CoLa A | ASCII | <STX> sAN {SPC} mLMPsetscancfg {SPC} 0 {SPC} 1388 {SPC} 1 {SPC} 1388 {SPC} FFF92230 {SPC} 225510 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 30 20 31 33 38 38 20 31 20 31 33 38 38 20 46 46 46 39 32 32 33 30 20 32 32 35 35 31 30 03 |

| CoLa B | Binary | 02 02 02 00 00 00 26 73 41 4E 20 6D 4C 4D 50 73 65 74 73 63 61 6E 63 66 67 20 00 00 00 13 88 00 01 00 00 13 88 FF F9 22 30 00 22 55 10 2D |
| --- | --- | --- |

Table 18: Example: sAN mLMPsetscancfg

    - Read for frequency and angular resolution

| Telegram structure: sRN LMPscancfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Info of scan frequency and angular resolution | String | 10 | All | LMPscancfg | 4C 4D 50 73 63 61 6E 63 66 67 |

Table 19: Telegram structure: sRN LMPscancfg

Example for LMS100: sRN LMPscancfg

| CoLa A | ASCII | <STX> sRN {SPC} LMPscancfg <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 4D 50 73 63 61 6E 63 66 67 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 52 4E 20 4C 4D 50 73 63 61 6E 63 66 67 63 |

Table 20: Example for LMS100: sRN LMPscancfg

| Telegram structure: sRA LMPscancfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Info of scan | String | 10 | All | LMPscancfg | 4C 4D 50 73 63 61 6E 63 |
|  | frequency and angular resolution |  |  |  |  | 66 67 |
| Scan frequency | [1/100 Hz] | Uint_32 | 4 | LMS1xx | 25Hz: +2500d (9C4h) 50Hz: +5000d (1388h) | 25Hz: 00 00 09 C4 50Hz: 00 00 13 88 |
|  |  |  |  | LMS5xx | 25Hz: +2500d (9C4h) | 25Hz: 00 00 09 C4 |
|  |  |  |  |  | 35Hz: +3500d (DACh) | 35Hz: 00 00 0D AC |
|  |  |  |  |  | 50Hz: +5000d (1388h) | 50Hz: 00 00 13 88 |
|  |  |  |  |  | 75Hz: +7500d (1A0Bh) | 75Hz: 00 00 1A 0B |
|  |  |  |  |  | 100Hz: +10000d (2710h) | 100Hz: 00 00 27 10 |
|  |  |  |  | TiM24x | 14.5Hz: +1450d (5AAh) | 14.5Hz: 00 00 05 AA |
|  |  |  |  | TiMxxx | 15Hz: +1500d (5DCh) | 15Hz: 00 00 05 DC |
|  |  |  |  | NAV310 | 5Hz... 20Hz: | 5Hz... 20Hz: |
|  |  |  |  | LD-OEM | + 500d... + 2000d | 00 00 01 F4... |
|  |  |  |  | 15xx | (1 F4h... 7 D0h) | 00 00 07 D0 |
|  |  |  |  | LD-LRS | 5Hz... 15Hz: | 5Hz... 15Hz: |
|  |  |  |  | 36xx | + 500d... + 1500d (1 F4h &hellip; 5DCh) | 00 00 01 F4... 00 00 05 DC |
|  |  |  |  | MRS1000 | 50Hz: +5000d (1388h) | 50Hz: 00 00 13 88 |
|  |  |  |  |  | 25Hz: +2500d (9C4h) | 25Hz: 00 00 09 C4 |
|  |  |  |  |  | 12,5 Hz: +1250d (4E2h) | 12,5Hz: 00 00 04 E2 |
|  |  |  |  | LMS1000 | 150Hz: +15000d (3A98h) | 150Hz: 00 00 3A 98 |
|  |  |  |  |  | 75Hz: +7500d (1D4Ch) | 75Hz: 00 00 1D 4C |
|  |  |  |  |  | 37,5 Hz: +3750d (0EA6h) | 37,5 Hz: 00 00 0E A6 |
|  |  |  |  | MRS6000 | 10Hz: +1000d (3E8h) | 10Hz: 00 00 03 E8 |
|  |  |  |  | LRS4000 | 12.5Hz:+ 1250d (4B0) | 12,5H z:00 00 04 B0 |
|  |  |  |  |  | 25Hz: +2500d (9C4) | 25Hz: 00 00 09 C4 |

| Telegram structure: sRA LMPscancfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Number of sectors | Indicates the number of sectors. The subsequent values will be transmitted 1&hellip; 4 accordingly. | Int_16 | 2 | LMS1xx LMS5xx TiMxxx MRS1000 LMS1000 MRS6000 | Sector 1: 0001h | Sector 1: 0001 |
|  |  |  |  | LRS4000 |  |  |
|  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | Sector 1: 0001h Sector 2: 0002h Sector 3: 0003h Sector 4: 0004h | Sector 1: 0001 Sector 2: 0010 Sector 3: 0011 Sector 4: 0100 |
| Angular resolution | [1/10000 °] | Uint_32 | 4 | LMS1xx | 0.25 °: + 2500d (9 C4h) 0.5: + 5000d (1388h) | 0.25: 00 00 09 C4 0.5: 00 00 13 88 |
|  |  |  |  | LMS5xx | 0.042: + 0417d (1 A1h) | 0.1667: 00 00 06 83 |
|  |  |  |  |  | 0.083: + 0833d (341h) | 0.25: 00 00 09 C4 |
|  |  |  |  |  | 0.1667: + 1667d (683h) | 0.333: 00 00 0D 05 |
|  |  |  |  |  | 0.25 °: + 2500d (9 C4h) | 0.5: 00 00 13 88 |
|  |  |  |  |  | 0.333: + 3333d (D05h) | 0.667: 00 00 1A 0B |
|  |  |  |  |  | 0.5: + 5000d (1388h) | 1: 00 00 27 10 |
|  |  |  |  |  | 0.667: + 6667d (1 A0Bh) |  |
|  |  |  |  |  | 1: + 10000d (2710h) |  |
|  |  |  |  | TiMxxx | 0.333: + 3333d (D05h) | 0.333: 00 00 0D 05 |
|  |  |  |  |  | 1: + 10000d (2710h) | 1: 00 00 27 10 |
|  |  |  |  | NAV310 | 0.125 °&hellip; 1 °: | 0.125 °&hellip; 1: |
|  |  |  |  | LD-OEM | + 1250d °&hellip; + 10000d | 00 00 04 E2... |
|  |  |  |  | 15xx | (4E2h °&hellip; 2710h) | 00 00 27 10 |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 36xx |  |  |
|  |  |  |  | MRS1000 | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  |  |  |  | 0.125: + 1250d (4 E2h) | 0,125: 00 00 04 E2 |
|  |  |  |  |  | 0.0625: + 0625d (271h) | 0,0625: 00 00 02 71 |
|  |  |  |  | LMS1000 | 0.75: + 7500d (1 D4Ch) | 0.75: 00 00 1D 4C |
|  |  |  |  |  | 0.375: + 3750d (0 EA6h) | 0.375: 00 00 0E A6 |
|  |  |  |  |  | 0.1875: + 1875d (0753h) | 0.1875: 00 00 07 53 |
|  |  |  |  | MRS6000 | 0.13: + 1300d (515h) | 0.13: 00 00 05 15 |

| Telegram structure: sRA LMPscancfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  | LRS4000 | @ 12.5Hz: | @ 12,5Hz: |
|  |  |  |  |  | 0.0200: + 200d (C8h) | 0.0200: 00 00 C8 |
|  |  |  |  |  | 0.0400: + 400d (190h) | 0.0400: 00 00 01 90 |
|  |  |  |  |  | 0.0600: + 600d (258h) | 0.0600: 00 00 02 58 |
|  |  |  |  |  | 0.1000: + 1000d (3 E8h) | 0.1000: 00 00 03 E8 |
|  |  |  |  |  | 0.1200: + 1200d (4 B0h) | 0.1200: 00 00 04 B0 |
|  |  |  |  |  | - | - - |
|  |  |  |  |  | @ 25Hz: | @ 25Hz: |
|  |  |  |  |  | 0.0400: + 400d (190h) | 0.0400: 00 00 01 90 |
|  |  |  |  |  | 0.0800 + 800d (320h) | 0.0800: 00 00 03 20 |
|  |  |  |  |  | 0.1200: + 1200d (4 B0h) | 0.1200: 00 00 04 B0 |
|  |  |  |  |  | 0.2000: + 2000d (7 D0h) | 0.2000: 00 00 07 D0 |
|  |  |  |  |  | 0.2400: + 2400d (960h | 0.2400: 00 00 09 60 |
| Start angle | [1/10000 °] | Int_32 | 4 | TiM24x | -1200000d... + 1200000d (FFEDB080h... 1 D4C0h) | FF ED B0 80&hellip; 00 01 D4 C0 |
|  |  |  |  | LMS1xx | -450000d... + 2250000d | FF F9 22 30&hellip; |
|  |  |  |  | TiMxxx | (FFF92230h &hellip; 225510h) | 00 22 55 10 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  | NAV310 | 0... + 3600000d | 00 00 00 00 00... |
|  |  |  |  | LD-OEM | (0h &hellip; 36 EE80h) | 00 36 EE 80 |
|  |  |  |  | 15x1 |  |  |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 36x1 |  |  |
|  |  |  |  | LD-OEM | -900000d... + 2700000d | FF F2 44 60&hellip; |
|  |  |  |  | 15x0 | (FFF24460h... 41 EB0h) | 00 04 1E B0 |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 36x0 |  |  |
|  |  |  |  | MRS1000 | -475000d (FFF8C088h) | FF F8 C0 88 |
|  |  |  |  | LMS1000 | -480000d (FFF8AD00h) | FF F8 AD 00 |
|  |  |  |  | MRS6000 | + 300000d (493 E0h) | 00 04 93 E0 |
|  |  |  |  | LRS4000 | -1800000d (FFE488C0h) | FF E4 88 C0 |

| Telegram structure: sRA LMPscancfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Stop angle | [1/10000 °] | Int_32 | 4 | TiM24x | -1200000d... + 1200000d (FFEDB080h... 1 D4C0h) | FF ED B0 80&hellip; 00 01 D4 C0 |
|  |  |  |  | LMS1xx | -450000d... + 2250000d | FF F9 22 30&hellip; |
|  |  |  |  | TiMxxx | (FFF92230h &hellip; 225510h) | 00 22 55 10 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  | NAV310 | 0... + 3600000d | 00 00 00 00 00... |
|  |  |  |  | LD-OEM | (0h &hellip; 36 EE80h) | 00 36 EE 80 |
|  |  |  |  | 15x1 |  |  |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 36x1 |  |  |
|  |  |  |  | LD-OEM | -900000d... + 2700000d | FF F2 44 60&hellip; |
|  |  |  |  | 15x0 | (FFF24460h... 41 EB0h) | 00 04 1E B0 |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 36x0 |  |  |
|  |  |  |  | MRS1000 | + 2275000d (22 B6B8h) | 00 22 B6 B8 |
|  |  |  |  | LMS1000 | + 2280000d (22 CA40h) | 00 22 CA 40 |
|  |  |  |  | MRS6000 | + 1500000d (16 E360h) | 00 16 E3 60 |
|  |  |  |  | LRS4000 | + 1800000d (1 B7740h) | 00 1B 77 40 |

Table 21: Telegram structure: sRA LMPscancfg

Example: sRA LMPscancfg

| CoLa A | ASCII | <STX> sRA {SPC} LMPscancfg {SPC} 1388 {SPC} 1 {SPC} 1388 {SPC} FFF92230 {SPC} 225510 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 4D 50 73 63 61 6E 63 66 67 31 33 38 38 20 31 20 31 33 38 38 20 46 46 46 39 32 32 33 30 20 32 32 35 35 31 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 21 73 52 41 20 4C 4D 50 73 63 61 6E 63 66 67 00 00 13 88 00 01 00 00 13 20 88FF F9 22 30 00 22 55 10 3E |

Table 22: Example: sRA LMPscancfg

    - Alignment mode (one layer activation for adjustment)

| Telegram structure: sWN MMAlignmentMode (Service) (sMN SetAccessMode 04 81BE23AA) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set device to alignment mode | String | 15 | All | MMAlignmentMode | 4D 4D 41 6C 69 67 6E 6D 65 6E 74 4D 6F 64 65 |
| Layer activation |  |  |  |  | all Layer: 0 red Layer -2.5°: 1 blue Layer 0°: 2 green Layer +2.5°: 3 yellow Layer +5: 4 | 00 01 02 03 04 |

Table 23: Telegram structure: sWN MMAlignmentMode

Example: sWN MMAlignmentMode 2

| CoLa A | ASCII | <STX> sWN {SPC} MMAlignmentMode {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4D 4D 41 6C 69 67 6E 6D 65 6E 74 4D 6F 64 65 20 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 57 4E 20 4D 4D 41 6C 69 67 6E 6D 65 6E 74 4D 6F 64 65 20 02 14 |

Table 24: Example: sWN MMAlignmentMode

| Telegram structure: sWA MMAlignmentMode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set device to alignment mode | String | 15 | All | MMAlignmentMode | 4D 4D 41 6C 69 67 6E 6D 65 6E 74 4D 6F 64 65 |

Table 25: Telegram structure: sWA MMAlignmentMode

Example: sWA MMAlignmentMode

| CoLa A | ASCII | <STX> sWA {SPC} MMAlignmentMode <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4D 4D 41 6C 69 67 6E 6D 65 6E 74 4D 6F 64 65 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 57 41 20 4D 4D 41 6C 69 67 6E 6D 65 6E 74 4D 6F 64 65 20 19 |

Table 26: Example: sWA MMAlignmentMode

    - Set scan configuration

Sets the device to an defined scan configuration, consisting of scan frequency, angular resolution, sector definition and interlace mode.

| Telegram structure: sMN mCLsetscancfglist |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set scan configuration | String | 17 | All | mCLsetscancfglist | 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 |
| Mode | Interlace mode (see table below) | Enum_8 | 1 | All | + 1d, + 2d, + 3d... (01h, 02h, 03h...) | 01, 02, 03... |

Table 27: Telegram structure: sMN mCLsetscancfglist

Interlace mode

The interlace mode allows to achieve a higher angular resolution by combining scans with lower resolution. The individual scans are shifted to each other.

The command mCLsetscancfglist selects combinations of scan resolution, scan frequency and resolution. If the scan area will not match to the application then an adjustment is possible by the command "mLMPsetscancfg” (see section 4.2.1 "Set frequency and angular resolution/measurement sectors" on page 17 ).

LD Series switching table:

LD-OEM15xx, LD-LRS36xx, NAV310

| Mode | Inter- laced | Scan freq. | Result. scan freq. | Reso- lution | Total Resol. | Field of view | Sector | LRS 3601 3611 | OEM 1501 | NAV 310 | LRS 3600 3610 | OEM 1500 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 0x | 8Hz | 8Hz | 0.25 ° | 0.25 ° | 360 ° | 0... 360 ° | x | x | x | (x) | (x) |
| 2 | 0x | 15Hz | 15Hz | 0.5 ° | 0.5 ° | 360 ° | 0... 360 ° | x | x | x | (x) | (x) |
| 3 | 0x | 10Hz | 10Hz | 0.25 ° | 0.25 ° | 300 ° | 30... 330 ° | x | x | x | x | x |
| 4 | 0x | 5Hz | 5Hz | 0.125 ° | 0.125 ° | 300 ° | 30... 330 ° | x | x | x | x | x |
| 5 | 0x | 6Hz | 6Hz | 0.1875 ° | 0.1875 ° | 360 ° | 0... 360 ° | x | x | x | (x) | (x) |
| 6 | 0x | 8Hz | 8Hz | 0.25 ° | 0.25 ° | 359.5 ° | 0.25 ° ... 359.25 ° |  |  |  | x | X |
| 8 | 0x | 15Hz | 15Hz | 0.375 ° | 0,375 ° | 300 ° | 30... 330 ° | x | X | x | x | x |
| 9 | 0x | 15Hz | 15Hz | 0.5 ° | 0.5 ° | 359 ° | 0.5... &hellip;. 359.5 ° |  |  |  | x | x |
| 21 | 0x | 20Hz | 20Hz | 0.5 ° | 0.5 ° | 300 ° | 30... 330 ° |  | X | x |  | x |
| 22 | 0x | 20Hz | 20Hz | 0.75 ° | 0.75 ° | 360 ° | 0... 360 ° |  | x | x |  | (x) |
| 44 | 4x | 10Hz | 2.5Hz | 0.25 ° | 0.0625 ° | 300 ° | 30... 330 ° | x | x |  | (x) | (x) |
| 46 | 4x | 16Hz | 4Hz | 0.5 ° | 0.125 ° | 300 ° | 30... 330 ° |  | x |  |  | (x) |

Table 28: Interlace mode for sMN mCLsetscancfglist for LD series

(x): Only at raw data scan (field application)

LMS5xx Variant switching table:

Default is 190° field of view from -5 to 185°. Filtering have to be set via data transmission output.

| Mode in Hex | Mode in Dec | Interlaced | Scan freq. | Resolution |
| --- | --- | --- | --- | --- |
| 0 | 0 | 0x | 25Hz | 0,167 ° |
| 1 | 1 | 0x | 25Hz | 0,25 ° |
| 2 | 2 | 0x | 35Hz | 0,25 ° |
| 3 | 3 | 0x | 35Hz | 0,5 ° |
| 4 | 4 | 0x | 50Hz | 0,333 ° |
| 5 | 5 | 0x | 50Hz | 0,5 ° |
| 6 | 6 | 0x | 75Hz | 0,5 ° |
| 7 | 7 | 0x | 75Hz | 1 ° |
| 8 | 8 | 0x | 100Hz | 0,667 ° |
| 9 | 9 | 0x | 100Hz | 1 ° |
| A | 10 | 2x | 50Hz | 0,167 ° |
| B | 11 | 2x | 75Hz | 0,25 ° |
| C | 12 | 2x | 100Hz | 0,167 ° |
| D | 13 | 2x | 100Hz | 0,333 ° |
| E | 14 | 2x | 100Hz | 0,5 ° |
| F | 15 | 4x | 25Hz | 0,083 ° |
| 10 | 16 | 4x | 25Hz | 0,042 ° |

Table 29: Interlace mode for sMN mCLsetscancfglist for LMS5xx

Remark for measurement data output with LMS5xx at 0,083° and 0,042°:

With angular resolution of 0,083° or 0,042°, it is recommended to compensate the systematic error of the angular position. To calibrate the LMS5 in the system a correction table should be established in the customer's system software.

This table ensures the mapping of an indicated angle in the scan data telegram on an angle in the real system environment.

The calibration could be done, e.g. with a test object or with a reference scan on the system structure with known geometry. The table should comprise the complete field of view which is relevant for the application.

LRS4000 switching table:

Default is 360° field of view from -180 to +180°

| Mode | Mode Name | Inter- laced | Scan freq. | Result. scan freq. | Reso- lution | Total Resol. | Field of view |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 11 | 12.5Hz & 0.020 ° | 0x | 12.5Hz | 12.5Hz | 0.040 ° | 0.020 ° | 288 ° |
| 1 | 12.5Hz & 0.040 ° | 0x | 12.5Hz | 12.5Hz | 0.040 ° | 0.040 ° | 360 ° |
| 2 | 12.5Hz & 0.060 ° | 0x | 12.5Hz | 12.5Hz | 0.060 ° | 0.060 ° | 360 ° |
| 4 | 12.5Hz & 0.100 ° | 0x | 12.5Hz | 12.5Hz | 0.100 ° | 0.100 ° | 360 ° |
| 5 | 12.5Hz & 0.120 ° | 0x | 12.5Hz | 12.5Hz | 0.120 ° | 0.120 ° | 360 ° |
| 71 | 25Hz & 0.040 ° | 0x | 25Hz | 25Hz | 0.080 ° | 0.040 ° | 288 ° |
| 61 | 25Hz & 0.080 ° | 0x | 25Hz | 25Hz | 0.080 ° | 0.080 ° | 360 ° |
| 62 | 25Hz & 0.120 ° | 0x | 25Hz | 25Hz | 0.120 ° | 0.120 ° | 360 ° |
| 64 | 25Hz & 0.200 ° | 0x | 25Hz | 25Hz | 0.200 ° | 0.200 ° | 360 ° |
| 65 | 25Hz & 0.240 ° | 0x | 25Hz | 25Hz | 0.240 ° | 0.240 ° | 360 ° |

Table 30: Interlace mode for sMN mCLsetscancfglist for LRS4000

Example: Set scan configuration 1: sMN mCLsetscancfglist 1

| CoLa A | ASCII | <STX> sMN {SPC} mCLsetscancfglist {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 17 20 73 4D 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 01 0F |

Table 31: Example: Set scan configuration 1: sMN mCLsetscancfglist 1

| Telegram structure: sAN mCLsetscancfglist |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Confirm scan configuration | String | 17 | All | mCLsetscancfglist | 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 |
| Status code | Wrong setting | Enum_8 | 1 | All | OK: 0 Error frequency: 1 | OK: 00 Error frequency: 01 |
|  |  |  |  |  | Error resolution: 2 | Error resolution: 02 |
|  |  |  |  |  | Err. res. and freq.: 3 | Err. res. and freq.: 03 |
|  |  |  |  |  | Err. scan field: 4 | Err. scan field: 04 |
|  |  |  |  |  | Error: 5 | Error: 05 |

Table 32: Telegram structure: sAN mCLsetscancfglist

Example: sAN mCLsetscancfglist OK

| CoLa A | ASCII | <STX> sAN {SPC} mCLsetscancfglist {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 41 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 00 10 |

Table 33: Example: sAN mCLsetscancfglist OK

Example LMS5xx

Note

After sending this telegram, it will take 30 seconds to process the new configuration in the sensor.

| Telegram structure: sMN mCLsetscancfglist (user level 'authorized client' required) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sMN | 73 4D 4E |
| Command | Set scan configuration | String | 17 | All | mCLsetscancfglist | 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 |
| Mode | Scan configuration | Enum_8 | 1 | All | 0d = 25Hz, 0.167 ° | 00 |
|  |  |  |  |  | 1d = 25Hz, 0.25 ° | 01 |
|  |  |  |  |  | 2d = 35Hz, 0.25 ° | 02 |
|  |  |  |  |  | 3d = 35Hz, 0.5 ° | 03 |
|  |  |  |  |  | 4d = 50Hz, 0.333 ° | 04 |
|  |  |  |  |  | 5d = 50Hz, 0.5 ° | 05 |
|  |  |  |  |  | 6d = 75Hz, 0.5 ° | 06 |
|  |  |  |  |  | 7d = 75Hz, 1.0 ° | 07 |
|  |  |  |  |  | 8d = 100Hz, 0.667 ° | 08 |
|  |  |  |  |  | 9d = 100Hz, 1.0 ° | 09 |
|  |  |  |  |  | 10d = 50Hz, 0.167° interl. | 0A |
|  |  |  |  |  | 11d = 75Hz, 0.25° interl. | 0B |
|  |  |  |  |  | 12d = 100Hz, 0.167° interl. | 0C |
|  |  |  |  |  | 13d = 100Hz, 0.333° interl. | 0D |
|  |  |  |  |  | 14d = 100Hz, 0.5° interl. | 0E |
|  |  |  |  |  | 15d = 25Hz, 0.083° interl. | 0F |
|  |  |  |  |  | 16d = 100Hz, 0.042° interl. | 10 |

Table 34: Telegram structure: sMN mCLsetscancfglist

LMS5xxx Example: sMN mCLsetscancfglist 5

| CoLa A | ASCII | <STX> sMN {SPC} mCLsetscancfglist {SPC} 5 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 35 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 4D 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 05 0A |

Table 35: Example: sMN mCLsetscancfglist 5

| Telegram structure: sAN mCLsetscancfglist (user level 'authorized client' required) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Confirm scan configuration | String | 17 | All | mCLsetscancfglist | 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 |
| Status code | Result | Enum_8 | 1 | All | 0d = OK 1d = Frequency error | 00 01 |
|  |  |  |  |  | 2d = Resolution error | 02 |
|  |  |  |  |  | 3d = Frequency and resolution combination error | 03 |
|  |  |  |  |  | 4d = Range error | 04 |
|  |  |  |  |  | 5d = General error | 05 |

Table 36: Telegram structure: sAN mCLsetscancfglist

Example: sAN mCLsetscancfglist 0

| CoLa A | ASCII | <STX> sAN {SPC} mCLsetscancfglist {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 41 4E 20 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 00 03 |

Table 37: Example: sAN mCLsetscancfglist 0

Example LRS4000

Note

After sending this telegram, it will take 30 seconds to process the new configuration in the sensor.

| Telegram structure: sMN mCLsetscancfglist (user level 'authorized client' required) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set scan configuration | String | 17 | All | mCLsetscancfglist | 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 |
| Mode | Scan configuration (see table below) | Enum_8 | 1 | All | 11d = 12.5Hz & 0.020 ° 1d = 12.5Hz & 0.040 ° | 0B 01 |
|  |  |  |  |  | 2d = 12.5Hz & 0.060 ° | 02 |
|  |  |  |  |  | 4d = 12.5Hz & 0.100 ° | 04 |
|  |  |  |  |  | 5d = 12.5Hz & 0.120 ° | 05 |
|  |  |  |  |  | 71d = 25Hz & 0.040 ° | 47 |
|  |  |  |  |  | 61d = 25Hz & 0.080 ° | 3D |
|  |  |  |  |  | 62d = 25Hz & 0.120 ° | 3E |
|  |  |  |  |  | 64d = 25Hz & 0.200 ° | 40 |
|  |  |  |  |  | 65d = 25Hz & 0.240 ° | 41 |

Table 38: Telegram structure: sMN mCLsetscancfglist

Example: Set scan configuration 1: sMN mCLsetscancfglist 1

| CoLa A | ASCII | <STX> sMN {SPC} mCLsetscancfglist {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 17 20 73 4D 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 01 0F |

Table 39: Example: Set scan configuration 1: sMN mCLsetscancfglist 1

| Telegram structure: sAN mCLsetscancfglist |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Confirm scan configuration | String | 17 | All | mCLsetscancfglist | 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 |
| Status code | Wrong setting | Enum_8 | 1 | All | OK: 0 Error frequency: 1 | OK: 00 Error frequency: 01 |
|  |  |  |  |  | Error resolution: 2 | Error resolution: 02 |
|  |  |  |  |  | Err. res. and freq.: 3 | Err. res. and freq.: 03 |
|  |  |  |  |  | Err. Range error: 4 | Err. scan field: 04 |
|  |  |  |  |  | General Error: 5 | Error: 05 |

Table 40: Telegram structure: sAN mCLsetscancfglist

Example: sAN mCLsetscancfglist OK

| CoLa A | ASCII | <STX> sAN {SPC} mCLsetscancfglist {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 41 4E 20 6D 43 4C 73 65 74 73 63 61 6E 63 66 67 6C 69 73 74 20 00 10 |

Table 41: Example: sAN mCLsetscancfglist OK

    - Activate standby mode

Shut off the laser in order to extend the lifetime of laser diode.

LMS1xx, MR1000, LMS1000 and LMS4000: The motor keeps on turning. LMS5xx: Reduce motor speed to 25Hz

| Telegram structure: sMN LMCstandby (All = Authorized client; LMS4000 = Operator) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set device to standby | String | 10 | All | LMCstandby | 4C 4D 43 73 74 61 6E 64 62 79 |

Table 42: Telegram structure: sMN LMCstandby

Example: sMN LMCstandby

| CoLa A | ASCII | <STX> sMN {SPC} LMCstandby <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 4C 4D 43 73 74 61 6E 64 62 79 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0E 73 4D 4E 20 4C 4D 43 73 74 61 6E 64 62 79 65 |

Table 43: Example: sMN LMCstandby

| Telegram structure: sAN LMCstandby |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Set device to standby | String | 10 | All | LMCstandby | 4C 4D 43 73 74 61 6E 64 62 79 |
| Status code | Accepted when value is 0 | Enum_8 | 1 | All | No error: 0 | No error: 00 |

Table 44: Telegram structure: sAN LMCstandby

Example: sAN LMCstandby

| CoLa A | ASCII | <STX> sAN {SPC} LMCstandby {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 4C 4D 43 73 74 61 6E 64 62 79 30 03 |

| CoLa B | Binary | 02 02 02 00 00 00 10 73 41 4E 20 4C 4D 43 73 74 61 6E 64 62 79 20 00 49 |
| --- | --- | --- |

Table 45: Example: sAN LMCstandby

    - Start measurement

Start the laser and (unless in Standby mode) the motor of the device

| Telegram structure: sMN LMCstartmeas (All = Authorized client; LMS4000 = Operator) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Start measurement | String | 12 | All | LMCstartmeas | 4C 4D 43 73 74 61 72 74 6D 65 61 73 |

Table 46: Telegram structure: sMN LMCstartmeas

Example: sMN LMCstartmeas

| CoLa A | ASCII | <STX> sMN {SPC} LMCstartmeas <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 4C 4D 43 73 74 61 72 74 6D 65 61 73 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 10 73 4D 4E 20 4C 4D 43 73 74 61 72 74 6D 65 61 73 68 |

Table 47: Example: sMN LMCstartmeas

| Telegram structure: sAN LMCstartmeas |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Start measurement | String | 12 | All | LMCstartmeas | 4C 4D 43 73 74 61 72 74 6D 65 61 73 |
| Status code | Accepted when value is 0 | Enum_8 | 1 | All | No error: 0 Not allowed: 1 | No error: 00 Not allowed: 01 |

Table 48: Telegram structure: sAN LMCstartmeas

Example: sAN LMCstartmeas

| CoLa A | ASCII | <STX> sAN {SPC} LMCstartmeas {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 4C 4D 43 73 74 61 72 74 6D 65 61 73 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 41 4E 20 4C 4D 43 73 74 61 72 74 6D 65 61 73 20 00 44 |

Table 49: Example: sAN LMCstartmeas

    - Stop measurement

LMS1xx/MRS1000/LMS1000/TiM5xx/TiM7xx/LMS4000: Shut off the laser and stop the motor. LMS5xx: Shut off the laser and the motor is running at the set up frequency.

| Telegram structure: sMN LMCstopmeas (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Stop measurement | String | 11 | All | LMCstopmeas | 4C 4D 43 73 74 6F 70 6D 65 61 73 |

Table 50: Telegram structure: sMN LMCstopmeas

Example: sMN LMCstopmeas

| CoLa A | ASCII | <STX> sMN {SPC} LMCstopmeas <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 4C 4D 43 73 74 6F 70 6D 65 61 73 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 4D 4E 20 4C 4D 43 73 74 6F 70 6D 65 61 73 10 |

Table 51: Example: sMN LMCstopmeas

| Telegram structure: sAN LMCstopmeas |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Stop measurement | String | 11 | All | LMCstopmeas | 4C 4D 43 73 74 6F 70 6D 65 61 73 |
| Status code | Accepted when value is 0 | Enum_8 | 1 | All | No error: 0 Not allowed: 1 | No error: 00 Not allowed: 01 |

Table 52: Telegram structure: sAN LMCstopmeas

Example: sAN LMCstopmeas

| CoLa A | ASCII | <STX> sAN {SPC} LMCstopmeas {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 4C 4D 43 73 74 6F 70 6D 65 61 73 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 41 4E 20 4C 4D 43 73 74 6F 70 6D 65 61 73 20 00 3C |

Table 53: Example: sAN LMCstopmeas

    - Autostart measurement

| Telegram structure: sWN LMPautostartmeas (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Autostart measurement | String | 16 | All | LMPautostartmeas | 4C 4D 50 61 75 74 6F 73 74 61 72 74 6D 65 61 73 |
| Status code | Activate / Deactivate Autostart | Bool_1 | 1 | All | Autostart off: 0 Autostart on: 1 | 00 01 |

Table 54: Telegram structure: sWN LMPautostartmeas

Example: sWN LMPautostartmeas 1

| CoLa A | ASCII | <STX> sWN {SPC} LMPautostartmeas {SPC} 1 <ETX> |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 50 61 75 74 6F 73 74 61 72 74 6D 65 61 73 31 03 |  |  |  |  |  |  |
| CoLa B | Binary | 02 02 02 02 00 00 00 16 | 73 57 4E | 20 | 4C 4D 50 61 75 74 6F 73 74 61 72 74 6D 65 61 73 | 20 | 01 | 4F |

Table 55: Example: sWN LMPautostartmeas 1

This parameter defines whether the scanner will start directly rotate and measure when powering up or remain in idle mode. The changed setting (saved with the command sMN mEEWriteall) will be then be active with the next power-up cycle.

In case autostart is deactivated the sensor will remain in idle mode as follows: LMS1xx: Laser: off / motor speed: stop

LMS5xx : 	Laser: off / motor speed: 25Hz LD-OEM15XX : Laser: off / motor speed: stop LD-LRS36XX: Laser: off / motor speed: stop MRS6000: 	Laser off / motor speed: stop

| Telegram structure: sWA LMPautostartmeas |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Autostart measure- ment | String | 14 | All | LMPautostartmeas | 4C 4D 43 73 74 61 72 74 6D 65 61 73 |

Table 56: Telegram structure: sWA LMDautostartmeas

Example: sWA LMPautostartmeas

| CoLa A | ASCII | <STX> sWA {SPC} LMPautostartmeas <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 4D 43 73 74 61 72 74 6D 65 61 73 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 57 41 20 4C 4D 50 61 75 74 6F 73 74 61 72 74 6D 65 61 73 20 41 |

Table 57: Example: sWA LMPautostartmeas

    - Laser Control

Define if laser is always on or rather switched on and off by specific trigger signal. Also select delay times and timeout.

| Telegram structure: sWN IOlasc (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Autostart measurement | String | 6 | All | IOlasc | 49 4F 6C 61 73 63 |
| Trigger Source | Select Trigger Source | Uint_8 | 1 | All | Free running: 0 Software: 1 Input 1: 2 Input 2: 3 Input 1 or 2: 4 | 00 01 02 03 04 |
| Delay on Start | Delay on Start in ms | Uint_16 | 2 | All | 0... +65535d (0 &hellip; FFFFh) | 00 00... FF FF |
| Delay on Stop | Delay on Stop in ms | Uint_16 | 2 | All | 0... +65535d (0 &hellip; FFFFh) | 00 00... FF FF |

| Telegram structure: sWN IOlasc (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Laser Timeout | Laser shuts down after defined time in s, if trigger signal does not disappear | Uint_16 | 2 | All | 0 = inactive 1... +65535d (0 &hellip; FFFFh) | 00 00 00 01... FF FF |
| Delay settings | Reserved | Uint_8 | 1 | All | Always: 0 | 00 |

Table 58: Telegram structure: sWN IOlasc

Example: sWN IOlasc 1 +500 0 0 0

| CoLa A | ASCII | <STX> sMN {SPC} IOlasc {SPC} 1 {SPC} +500 {SPC} 0 {SPC} 0 {SPC} 0 <ETX> |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Hex | 02 73 57 4E 20 49 4F 6C 61 73 63 20 31 20 2B 35 30 30 20 30 20 30 20 30 03 |  |  |  |  |
| CoLa B | Binary | 02 02 02 02 | 00 00 00 16 | 73 57 4E | 20 49 4F 6C 61 73 63 20 01 01 F4 00 00 00 00 00 00 00 00 00 00 20 | A5 |

Table 59: Example: sWN IOlasc with Software Trigger and 0.5 s delay on start.

| Telegram structure: sWA IOlasc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Autostart measure- ment | String | 14 | All | IOlasc | 49 4F 6C 61 73 63 |

Table 60: Telegram structure: sWA IOlasc

Example: sWA IOlasc

| CoLa A | ASCII | <STX> sWA {SPC} IOlasc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 49 4F 6C 61 73 63 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0B 73 57 41 20 49 4F 6C 61 73 63 20 7E |

Table 61: Example: sWA IOlasc

    - Activate/deactivate field application

With the aid of the integrated field application, the LD-OEM1500/LD-LRS3600 evaluates up to four evaluation fields within its scan area.

| Telegram structure: sWN CLApplication (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Activate/deactivate field application | String | 13 | All | CLApplication | 43 4C 41 70 70 6C 69 63 61 74 69 6F 6E |
| Mode | Application | Enum_ 16 | 2 | All | Scan only: 00 Field application: 11 | Scan only: 00 00 Field application: 00 11 |

Table 62: Telegram structure: sWN CLApplication

Example: Activate the field application: sWN CLApplication 11

| CoLa A | ASCII | <STX> sWN {SPC} CLApplication {SPC} 11 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 43 4C 41 70 70 6C 69 63 61 74 69 6F 6E 20 31 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 57 4E 20 43 4C 41 70 70 6C 69 63 61 74 69 6F 6E 20 00 11 1F |

Table 63: Example: Activate the field application: sWN CLApplication 11

| Telegram structure: sWA CLApplication |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Activate/deactivate field application | String | 13 | All | CLApplication | 43 4C 41 70 70 6C 69 63 61 74 69 6F 6E |

Table 64: Telegram structure: sWA CLApplication

Example: sWA CLApplication correct and accepted

| CoLa A | ASCII | <STX> sWA {SPC} CLApplication <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 43 4C 41 70 6C 69 63 61 74 69 6F 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 57 41 20 43 4C 41 70 70 6C 69 63 61 74 69 6F 6E |

Table 65: Example: sWA CLApplication correct and accepted

    - Application selection and switching

Selection between the field application (default) and the ranging application in the device.

| Telegram structure: sWN SetActiveApplications (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Selects all currently active applications of the scanner | String | 13 | All | SetActiveApplications | 43 4C 41 70 70 6C 69 63 61 74 69 6F 6E |
| Array lenght |  |  |  | All | 0 .. 1 | 00... 01 |
| Identifier | Application | String |  |  | FEVL (Field Application) RANG (Ranging) | 46 45 56 4C 52 41 4E 47 |
| Active |  | Bool_1 |  |  | False = 0 True = 1 | False = 00 True = 01 |

Table 66: Telegram structure: sWN SetActiveApplications

Example: Activate the field application: sWN CLApplication 11

| CoLa A | ASCII | <STX> sWN {SPC} SetActiveApplications {SPC} 1 {SPC} FEVL {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 73 57 4E 20 53 65 74 41 63 74 69 76 65 41 70 70 6C 69 63 61 74 69 6F 6E 73 20 31 20 46 45 56 4C 20 31 |
| CoLa B | Binary | 02 02 02 02 00 00 00 22 73 57 4E 20 53 65 74 41 63 74 69 76 65 41 70 6C 69 63 61 74 69 6F 6E 73 20 31 20 46 45 56 4C 20 31 34 |

Table 67: Example: Activate the field application: : sWN SetActiveApplications 1 FEVL 1

| Telegram structure: sWA SetActiveApplications |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Selects all currently active applications of the scanner | String |  | All | SetActiveApplications | 53 65 74 41 63 74 69 76 65 41 70 70 6C 69 63 61 74 69 6F 6E 73 |

Table 68: Telegram structure: sWA SetActiveApplications

Example: sWA CLApplication correct and accepted

| CoLa A | ASCII | <STX> sWA {SPC} SetActiveApplications <ETX> |
| --- | --- | --- |
|  | Hex | 73 57 41 20 53 65 74 41 63 74 69 76 65 41 70 6C 69 63 61 74 69 6F 6E 73 |
| CoLa B | Binary | 02 02 02 02 00 00 00 19 73 57 41 20 53 65 74 41 41 63 74 69 76 65 41 70 70 6C 69 63 61 74 69 6F 6E 73 02 |

Table 69: Example: sWA SetActiveApplications correct and accepted

    - Read Application selection and switching

| Telegram structure: sRN SetActiveApplications |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Info of scan frequency and angular resolution | String | 10 | All | SetActiveApplications | 53 65 74 41 63 74 69 76 65 41 70 70 6C 69 63 61 74 69 6F 6E 73 |

Table 70: Telegram structure: sRN SetActiveApplications

Example for MRS1000: sRN SetActiveApplications

| CoLa A | ASCII | <STX> sRN {SPC} SetActiveApplications <ETX> |
| --- | --- | --- |
|  | Hex | 73 52 4E 20 53 65 74 41 63 74 69 76 65 41 70 70 6C 69 63 61 74 69 6F 6E 73 |
| CoLa B | Binary | 02 02 02 02 00 00 00 19 73 52 4E 20 53 65 74 41 63 74 69 76 65 41 70 6C 69 63 61 74 69 6F 6E 73 08 |

Table 71: Example for MRS1000: sRN SetActiveApplications

| Telegram structure: sRA SetActiveApplications |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Info of scan frequency and angular resolution | String | 10 | All | SetActiveApplications | 73 52 4E 20 53 65 74 41 63 74 69 76 65 41 70 70 6C 69 63 61 74 69 6F 6E 73 |

    - Load factory defaults

NOTE

The Factory-Reset (Load factory defaults) deletes the entire parametrization of the device. All parameters, settings and system applications will be set to default.

| Telegram structure: sMN mSCloadfacdef (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | Not possible |
| Command | Load factory defaults | String | 13 | All | mSCloadfacdef | Not possible |

Table 72: Telegram structure: sMN mSCloadfacdef

Example: sMN mSCloadfacdef

| CoLa A | ASCII | <STX> sMN {SPC} mSCloadfacdef <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 53 43 6C 6F 61 64 66 61 63 64 65 66 03 |
| CoLa B | Binary | Not possible |

Table 73: Example: sMN mSCloadfacdef

| Telegram structure: sAN mSCloadfacdef |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | Not possible |
| Command | Load factory defaults | String | 13 | All | mSCloadfacdef | Not possible |

Table 74: Telegram structure: sAN mSCloadfacdef

Example: sAN mSCloadfacdef

| CoLa A | ASCII | <STX> sAN {SPC} mSCloadfacdef <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 53 43 6C 6F 61 64 66 61 63 64 65 66 03 |
| CoLa B | Binary | Not possible |

Table 75: Example: sAN mSCloadfacdef

    - Load application defaults

NOTE

The Application-Reset (Load application defaults) deletes only the user parametrization of the Fields, Evaluation cases (EVC) and parameters under the header "Application”. Other parameters like Interface settings, Echo Filter, etc. remain unaffected.

| Telegram structure: sMN mSCloadappdef (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | Not possible |
| Command | Load application defaults | String | 13 | All | mSCloadappdef | Not possible |

Table 76: Telegram structure: sMN mSCloadappdef

Example: sMN mSCloadappdef

| CoL a A | ASCII | <STX> sMN {SPC} mSCloadappdef <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 53 43 6C 6F 61 64 70 70 64 65 66 03 |
| Co La B | Binary | Not possible |

Table 77: Example: sMN mSCloadappdef

| Telegram structure: sAN mSCloadappdef |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | Not possible |
| Command | Load application defaults | String | 13 | All | mSCloadappdef | Not possible |

Table 78: Telegram structure: sAN mSCloadappdef

Example: sAN mSCloadappdef

| CoLa A | ASCII | <STX> sAN {SPC} mSCloadappdef <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 53 43 6C 6F 61 64 61 70 70 64 65 66 03 |
| CoLa B | Binary | Not possible |

Table 79: Example: sAN mSCloadappdef

    - Change password

NOTE

If logged in with a higher level you may set the password for lower levels as well.

| Telegram structure: sMN SetPassword (the same User level or higher) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set password request | String | 13 | All | SetPassword | 53 65 74 50 61 73 73 77 6F 72 64 |
| User level | User level that the password will be applied | Int_8 | 1 | All | Maintenance: 02 Authorized client: 03 Service: 04 | Maintenance: 02 Authorized client: 03 Service: 04 |
| Password | Hash value of the new password | Uint_32 | 4 | All | <Hash value> | <Hash value> |

Table 80: Telegram structure: sMN SetPassword

Example: sMN SetPassword

Set password for Authorized user to "testtest ”.

| CoLa A | ASCII | <STX> sMN {SPC} SetPassword {SPC} 03 {SPC} 19 20 E4 C9 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 53 65 74 50 61 73 73 77 6F 72 64 20 30 33 20 19 20 E4 C9 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 4D 4E 20 53 65 74 50 61 73 73 77 6F 72 64 20 30 33 20 19 20 E4 C9 3A |

Table 81: Example: sMN SetPassword

Calculating the hash value of the password

 	Login SOPAS with user level "Service ”.

 	Select [Device] > Password > Calculate Hash value.

| Telegram structure: sAN SetPassword |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | Not possible |
| Command | Set password requested | String | 13 | All | SetPassword | 53 65 74 50 61 73 73 77 6F 72 64 |
| Success | Confirmation | Int_8 | 1 | All | 0: Failed 1: Success | 0: Failed 1: Success |

Table 82: Telegram structure: sAN SetPassword

Example: sAN SetPassword

| CoLa A | ASCII | <STX> sAN {SPC} SetPassword {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 53 65 74 50 61 73 73 77 6F 72 64 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 41 4E 20 53 65 74 50 61 73 73 77 6F 72 64 20 31 30 |

Table 83: Example: sAN SetPassword

    - Check password

| Telegram structure: sMN CheckPassword (the same User level or higher) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Check password request | String | 13 | All | CheckPassword | 43 68 65 63 6B 50 61 73 73 77 6F 72 64 |
| User level | User level to check the password | Int_8 | 1 | All | Maintenance: 02 Authorized client: 03 Service: 04 | Maintenance: 02 Authorized client: 03 Service: 04 |
| Password | Hash value of the password to be checked | Uint_32 | 4 | All | <Hash value> | <Hash value> |

Table 84: Telegram structure: sMN CheckPassword

Example: sMN CheckPassword

Check password "testtest” for Authorized user.

| CoLa A | ASCII | <STX> sMN {SPC} CheckPassword {SPC} 03 {SPC} 19 20 E4 C9 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 43 68 65 63 6B 50 61 73 73 77 6F 72 64 20 30 33 20 19 20 E4 C9 03 |
| CoLa B | Binary | 02 02 02 00 00 00 19 73 4D 4E 20 43 68 65 63 6B 50 61 73 73 77 6F 72 64 20 30 33 20 19 20 E4 C9 0E |

Table 85: Example: sMN CheckPassword

| Telegram structure: sAN CheckPassword |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | Not possible |
| Command | Check password requested | String | 13 | All | CheckPassword | 43 68 65 63 6B 50 61 73 73 77 6F 72 64 |
| Success | Confirmation | Int_8 | 1 | All | 0: Failed 1: Success | 0: Failed 1: Success |

Table 86: Telegram structure: sAN CheckPassword

Example: sAN CheckPassword

| CoLa A | ASCII | <STX> sAN {SPC} CheckPassword {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 43 68 65 63 6B 50 61 73 73 77 6F 72 64 20 30 31 03 |
| CoLa B | Binary | 02 73 41 4E 20 43 68 65 63 6B 50 61 73 73 77 6F 72 64 31 03 |

Table 87: Example: sAN CheckPassword

    - Reboot device

This command includes saving all parameters.

| Telegram structure: sMN mSCreboot (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Reboot device | String | 9 | All | mSCreboot | 6D 53 43 72 65 62 6F 6F 74 |

Table 88: Telegram structure: sMN mSCreboot

Example: sMN mSCreboot

| CoLa A | ASCII | <STX> sMN {SPC} mSCreboot <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 53 43 72 65 62 6F 6F 74 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0D 73 4D 4E 20 6D 53 43 72 65 62 6F 6F 74 2C |

Table 89: Example: sMN mSCreboot

| Telegram structure: sAN mSCreboot |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Reboot device | String | 9 | All | mSCreboot | 6D 53 43 72 65 62 6F 6F 74 |

Table 90: Telegram structure: sAN mSCreboot

Example: sAN mSCreboot

| CoLa A | ASCII | <STX> sAN {SPC} mSCreboot <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 53 43 72 65 62 6F 6F 74 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0E 73 41 4E 20 6D 53 43 72 65 62 6F 6F 74 00 |

Table 91: Example: sAN mSCreboot

    - Set contamination measurement settings

| Telegram structure: sWN LCMcfg (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command |  | String | 6 | All | LCMcfg | 4C 43 4D 63 66 67 |
| Strategy | Strategy code | Enum_8 | 1 | All | Inactive: 0 | Inactive: 00 |
|  |  |  |  |  | High available: 1 | High available: 01 |
|  |  |  |  |  | Available: 2 | Available: 02 |
|  |  |  |  |  | Sensitive: 3 | Sensitive: 03 |
|  |  |  |  |  | Semi-sensitive: 4 | Semi-sensitive: 04 |
| Response time | Time lapse | Uint_16 | 4 | All | + 1d... + 60d (01h &hellip; 3Ch) | 00 01... 00 3C |
| Threshold warning | Threshold value | Uint_16 | 4 | All | 0d... + 100d (00h &hellip; 64h) | 00 00... 00 64 |
| Threshold error | Threshold value | Uint_16 | 4 | All | 0d... + 100d (00h &hellip; 64h) | 00 00... 00 64 |

Table 92: Telegram structure: sWN LCMcfg

Example: sWN LCMcfg

| CoLa A | ASCII | <STX> sWN {SPC} LCMcfg {SPC} 1 {SPC} +30 {SPC} +65 {SPC} +45 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 43 4D 63 66 67 20 31 20 2B 33 30 20 2B 36 35 20 2B 34 35 03 |
| CoLa B | Binary | 02 02 02 02 00 00 18 73 57 4E 20 4C 43 4D 63 66 67 20 01 00 1E 00 41 00 2D 39 |

Table 93: Example: sWN LCMcfg

| Telegram structure: sWA LCMcfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command |  | String | 6 | All | LCMcfg | 4C 43 4D 63 66 67 |

Table 94: Telegram structure: sWA LCMcfg

Example: sWA LCMcfg

| CoLa A | ASCII | <STX> sWA {SPC} LCMcfg <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 43 4D 63 66 67 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0B 73 57 41 20 4C 43 4D 63 66 67 45 |

Table 95: Example: sWA LCMcfg

    - Set contamination indication settings

| Telegram structure: sWN ContaminationConfig (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command |  | String | 19 | All | ContaminationConfig | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 |
| Strategy | Strategy code | Enum_8 | 1 | All | Inactive: 0 High available: 1 Available: 2 Semi-sensitive: 3 Sensitive: 4 | Inactive: 00 High available: 01 Available: 02 Semi-sensitive: 03 Sensitive: 04 |
| Response time | Time lapse | Uint_16 | 2 | All | + 1d... + 60d (01h &hellip; 3Ch) | 00 01... 00 3C |
| Threshold warning | Threshold value | Enum_8 | 1 | All | LOW: 0 MID: 1 HIGH: 2 | LOW: 00 MID: 01 HIGH 02 |
| Cover | Weatherhood constalation | Enum_8 | 1 | All | No Weatherhood: 0 Weatherh. 210 no Flaps: 1 Weatherh. 210 Flap left: 2 Weatherh. 210 Flap right: 3 Weatherh. 210 both Flaps:4 Weatherhhood 275: 5 | No Weatherhood: 00 Weatherh. 210 no Flaps: 01 Weatherh. 210 Flap left: 02 Weatherh. 210 Flap ri: 03 Weatherh. 210 B. Flaps: 04 Weatherhhood 275: 05 |
| Enable Warning | Warning monitoring | Bool_1 |  |  | Off: 0 On: 1 | OFF (False): 00 ON (True): 01 |
| Enable Error | Error monitoring | Bool_1 |  |  | Off: 0 On: 1 | OFF (False): 00 ON (True): 01 |

Table 96: Telegram structure: sWN ContaminationConfig

Example: sWN ContaminationConfig

| CoLa A | ASCII | <STX> sWN {SPC} ContaminationConfig {SPC} 1 {SPC} 3 {SPC} 2 {SPC} 0 {SPC} 1 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 20 31 20 33 20 32 20 30 20 31 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 1F 73 57 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 20 01 00 03 02 00 01 01 00 |

Table 97: Example: sWN ContaminationConfig

| Telegram structure: sWA ContaminationConfig |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command |  | String | 19 | All | ContaminationConfig | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 |

Table 98: Telegram structure: sWA ContaminationConfig

Example: sWA ContaminationConfig

| CoLa A | ASCII | <STX> sWA {SPC} ContaminationConfig <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 03 |
| C o La B | Binary | 02 02 02 02 00 00 00 18 73 57 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 20 |

Table 99: Example: sWA ContaminationConfig

    - Read contamination measurement settings

| Telegram structure: sRN LCMcfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 6 | All | LCMcfg | 4C 43 4D 63 66 67 |

Table 100: Telegram structure: sRN LCMcfg

Example: sRN ContaminationConfig

| CoLa A | ASCII | <STX> sRN {SPC} LCMcfg <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 43 4D 63 66 67 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 52 4E 20 4C 43 4D 63 66 67 6F |

Table 101: Example: sRN LCMcfg

| Telegram structure: sRA LCMcfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read for settings | String | 6 | All | LCMcfg | 4C 43 4D 63 66 67 |
| Strategy | Strategy code | Enum_8 | 1 | All | Inactive: 0 | Inactive: 00 |
|  |  |  |  |  | High available: 1 | High available: 01 |
|  |  |  |  |  | Available: 2 | Available: 02 |
|  |  |  |  |  | Sensitive: 3 | Sensitive: 03 |
|  |  |  |  |  | Semi-sensitive: 4 | Semi-sensitive: 04 |
| Response time | Time lapse | Uint_16 | 2 | All | + 1d... + 60d (00h &hellip; 3Ch) | 00 00... 00 3C |
| Threshold warning | Threshold value | Uint_16 | 2 | All | 0d... + 100d (00h &hellip; 64h) | 00 00... 00 64 |
| Threshold error | Threshold value | Uint_16 | 2 | All | 0d... + 100d (00h &hellip; 64h) | 00 00... 00 64 |

Table 102: Telegram structure: sRA LCMcfg

Example: sRA LCMcfg

| CoLa A | ASCII | <STX> sRA {SPC} LCMcfg {SPC} 1 {SPC} 1 {SPC} 46 {SPC} 1E <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 43 4D 63 66 67 20 31 20 31 20 34 36 20 31 45 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 41 20 4C 43 4D 63 66 67 20 01 00 01 00 46 00 1E 18 |

Table 103: Example: sRA LCMcfg

    - Read contamination indication settings

| Telegram structure: sRN ContaminationConfig |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 19 | All | ContaminationConfig | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 |

Table 104: Telegram structure: sRN ContaminationConfig

Example: sRN ContaminationConfig

| CoLa A | ASCII | <STX> sRN {SPC} ContaminationConfig <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 52 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 25 |

Table 105: Example: sRN ContaminationConfig

| Telegram structure: sRA ContaminationConfig |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 19 | All | ContaminationConfig | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 |
| Strategy | Strategy code | Enum_8 | 1 | All | Inactive: 0 High available: 1 Available: 2 Semi-sensitive: 3 Sensitive: 4 | Inactive: 00 High available: 01 Available: 02 Semi-sensitive: 03 Sensitive: 04 |
| Response time | Time lapse | Uint_16 | 4 | All | + 1d... + 60d (01h &hellip; 3Ch) | 00 01... 00 3C |
| Threshold warning | Threshold value | Enum_8 | 1 | All | LOW: 0 MID: 1 HIGH: 2 | LOW: 00 MID: 01 HIGH 02 |
| Cover | Weatherhood constalation | Enum_8 | 1 | All | No Weatherhood: 0 Weatherh. 210 no Flaps: 1 Weatherh. 210 Flap left: 2 Weatherh. 210 Flap right: 3 Weatherh. 210 both Flaps:4 Weatherhhood 275: 5 | No Weatherhood: 00 Weatherh. 210 no Flaps: 01 Weatherh. 210 Flap left: 02 Weatherh. 210 Flap ri: 03 Weatherh. 210 B. Flaps: 04 Weatherhhood 275: 05 |
| Enable Warning | Warning monitoring | Bool_1 |  |  | Off: 0 On: 1 | OFF (False): 00 ON (True): 01 |
| Enable Error | Error monitoring | Bool_1 |  |  | Off: 0 On: 1 | OFF (False): 00 ON (True): 01 |

Table 106: Telegram structure: sRA ContaminationConfig

Example: sRA ContaminationConfig

| CoLa A | ASCII | <STX> sRA {SPC} ContaminationConfig {SPC} 1 {SPC} 3 {SPC} 2 {SPC} 0 {SPC} 1 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 3 1 20 33 20 32 20 3 0 20 3 1 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1F 73 52 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 43 6F 6E 66 69 67 20 01 00 03 02 00 01 01 0A |

Table 107: Example: sRA ContaminationConfig

    - Read contamination measurement detailed values

| Telegram structure: sRN CMContLvlM |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 10 | All | CMContLvlM | 43 4D 43 6F 6E 74 4C 76 6C 4D |

Table 108: Telegram structure: sRN CMContLvlM

Example: sRN CMContLvlM

| CoLa A | ASCII | <STX> sRN {SPC} CMContLvlM <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 43 4D 43 6F 6E 74 4C 76 6C 4D 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0E 73 52 4E 20 43 4D 43 6F 6E 74 4C 76 6C 4D 6C |

Table 109: Example: sRN CMContLvlM

| Telegram structure: sRA CMContLvlM |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 10 | All | CMContLvlM | 43 4D 43 6F 6E 74 4C 76 6C 4D |
| Contamina- tion data for different channels | [% of transparency] in order of the different channels | Uint_8 | 1 | LMS1xx | Order of 7 channels: -25.8/12.8/51.4/90/128.6/167.2/205.8 0d... + 100d (00h &hellip; 64h) | Order of 7 channels: -25.8/12.8/51.4/90/128.6/167.2/205.8 00... 64 |
|  |  |  |  | LMS5xx NAV310 LD-OEM 15xx LD-LRS 36xx | Order of 6 channels: 5°/35°/70°/110°/145°/ 175 ° 0d... + 100d (00h &hellip; 64h) | Order of 6 channels: 5°/35°/70°/110°/145°/ 175 ° 00... 64 |

Table 110: Telegram structure: sRA CMContLvlM

Example for LMS5xx: sRA CMContLvlM

5°- to 110°-channel: 100 %, 145°- and 175°-channel only 84 % availability:

| CoLa A | ASCII | <STX> sRA {SPC} CMContLvlM {SPC} 64 {SPC} 64 {SPC} 64 {SPC} 64 {SPC} 54 {SPC} 54 {SPC}<ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 43 4D 43 6F 6E 74 4C 76 6C 4D 20 64 64 64 64 54 54 03 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 52 41 20 43 4D 43 6F 6E 74 4C 76 6C 4D 20 64 64 64 64 54 54 43 |

Table 111: Example for LMS5xx: sRA CMContLvlM

    - Read contamination indication data

| Telegram structure: sRN ContaminationData |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 17 | All | ContaminationData | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 44 61 74 61 |

Table 112: Telegram structure: sRN ContaminationData

Example: sRN ContaminationData

| CoLa A | ASCII | <STX> sRN {SPC} ContaminationData <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 44 61 74 61 |
| CoLa B | Binary | 02 02 02 02 00 00 00 15 73 52 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 44 61 74 61 3F |

Table 113: Example: sRN ContaminationData

| Telegram structure: sRA ContaminationData |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 17 | All | ContaminationData | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 44 61 74 61 |
| Contamina- tion data for different channels | [Status of contamination] in order of the different channels | Enum_8 | 1 | MRS 1000 LMS 1000 | Order of 6 channels: (1 / 2 / 3 / 4 / 5 / 6 ) CM channel: x CM NONE: 0 | Order of 6 channels: (1 / 2 / 3 / 4 / 5 / 6 ) CM channel: x CM NONE: 00 |
|  |  |  |  |  | CM WARN: 1 | CM WARN: 01 |
|  |  |  |  |  | CM ERROR: 2 | CM ERROR: 02 |

Table 114: Telegram structure: sRA ContaminationData

Example: sRA ContaminationData

| CoLa A | ASCII | <STX> sRA {SPC} ContaminationData {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC}<ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 44 61 74 20 30 20 30 20 30 20 30 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 1C 73 52 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 44 61 74 00 00 00 00 20 00 00 10 |

Table 115: Example: sRA ContaminationData

    - Read contamination indication result

| Telegram structure: sRN ContaminationResult |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 19 | All | ContaminationResult | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 52 65 73 75 6C 74 |

Table 116: Telegram structure: sRN ContaminationResult

Example: sRN ContaminationResult

| CoLa A | ASCII | <STX> sRN {SPC} ContaminationResult <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 52 65 73 75 6C 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 52 4E 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 52 65 73 75 6C 74 26 |

Table 117: Example: sRN ContaminationResult

| Telegram structure: sRA ContaminationResult |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 19 | All | ContaminationResult | 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 52 65 73 75 6C 74 |
| Contamina- tion Result | Result of contamination indication | Enum_8 | 1 | MRS 1000 LMS 1000 | Order of 2 results: (Warning / Error ) Result: Warning / Error | Order of 2 results: (Warning / Error ) Result: Warning / Error |
|  |  |  |  |  | False: 0 | False: 00 |
|  |  |  |  |  | True: 1 | True: 01 |

Table 118: Telegram structure: sRA ContaminationResult

Example: sRA ContaminationResult

| CoLa A | ASCII | <STX> sRA {SPC} ContaminationResult {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 52 65 73 75 6C 74 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 1A 73 52 41 20 43 6F 6E 74 61 6D 69 6E 61 74 69 6F 6E 52 65 73 75 6C 74 20 00 00 09 |

Table 119: Example: sRA ContaminationResult

    - Save parameters permanently

| Telegram structure: sMN mEEwriteall (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Store parameters permanently | String | 11 | All | mEEwriteall | 6D 45 45 77 72 69 74 65 61 6C 6C |

Table 120: Telegram structure: sMN mEEwriteall

Example: sMN mEEwriteall

| CoLa A | ASCII | <STX> sMN {SPC} mEEwriteall <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 45 45 77 72 69 74 65 61 6C 6C 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 4D 4E 20 6D 45 45 77 72 69 74 65 61 6C 6C 21 |

Table 121: Example: sMN mEEwriteall

| Telegram structure: sAN mEEwriteall |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Store parameters permanently | String | 11 | All | mEEwriteall | 6D 45 45 77 72 69 74 65 61 6C 6C |
| Status code | Accepted when value is 1 | Bool_1 | 1 | All | Error: 0 Success: 1 | Error: 00 Success: 01 |

Table 122: Telegram structure: sAN mEEwriteall

Example: sAN mEEwriteall

| CoLa A | ASCII | <STX> sAN {SPC} mEEwriteall {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 45 45 77 72 69 74 65 61 6C 6C 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 41 4E 20 6D 45 45 77 72 69 74 65 61 6C 6C 20 01 0C |

Table 123: Example: sAN mEEwriteall

    - Set to run

Log out from device and activate all parameter changes.

| Telegram structure: sMN Run |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Start the device | String | 3 | All | Run | 52 75 6E |

Table 124: Telegram structure: sMN Run

Example: sMN Run

| CoLa A | ASCII | <STX> sMN {SPC} Run <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 52 75 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 07 73 4D 4E 20 52 75 6E 19 |

Table 125: Example: sMN Run

| Telegram structure: sAN Run |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Start the device | String | 3 | All | Run | 52 75 6E |
| Status code | Accepted when value is 1 | Bool_1 | 1 | All | Error: 0 Success: 1 | Error: 00 Success: 01 |

Table 126: Telegram structure: sAN Run

Example: sAN Run

| CoLa A | ASCII | <STX> sAN {SPC} Run {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 52 75 6E 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 09 73 41 4E 20 52 75 6E 20 01 34 |

Table 127: Example: sAN Run

  - Measurement output telegram

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV3 10 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.3.1 | Configure the data content for the scan | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.3.2 | Configure measurement angle of the scandata for output | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.3.3 | Read for actual output range | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.3.4 | Configure the distance scaling of the scandata for the output |  |  |  |  |  |  |  |  |  |  |  |  | x |
| 4.3.5 | Read the distance scaling for the distance values of the scandata |  |  |  |  |  |  |  |  |  |  |  |  | x |
| 4.3.6 | Poll one telegram | x | x | x | x | x | x | x | x | x | x |  | x | x |
| 4.3.7 | Send data permanently | x | x | x | x | x | x | x | x | x | x | x | x | x |

    - Configure the data content for the scan

| Telegram structure: sWN LMDscandatacfg (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Configure scandata | String | 14 | All | LMDscandatacfg | 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 |
| Data channel | Defines the telegram content (DIST) | Uint_8 | 2 | LMS1xx | Output channel 1: 1 0 Output channel 2: 2 0 | 01 00 02 00 |
|  |  |  |  |  | Output channel 1+2: 3 0 | 03 00 |
|  |  |  |  | LMS5xx | Set via Echo Filter, therefore: 0 | 00 00 |
|  |  |  |  |  | Other values than 0 will be ignored. |  |
|  |  |  |  | TiMxxx NAV310 LD-OEM 15xx LD-LRS 36xx | Output channel 1: 1 0 | 01 00 |
|  |  |  |  | MRS1000 LMS1000 | Set via Echo Filter, therefore: 0 | 00 00 |
|  |  |  |  | MRS6000 | Other values than 0 will be ignored. |  |
|  |  |  |  | LRS4000 |  |  |
|  |  |  |  | LMS 4000 | No distance values: 0 0 | 00 00 |
|  |  |  |  |  | Distance values: 1 0 | 01 00 |
| Further | Remission, Angle | Uint_8 | 1 | All | No values: 0 | 00 |
| data channels | offset, Ambient light & Quality data output |  |  |  | RSSI: 1 | 01 |
|  |  |  |  | MRS1000 | No values: 0 | 00 |
|  |  |  |  | LMS1000 | RSSI: 1 | 01 |
|  |  |  |  |  | AINF: 8 | 08 |
|  |  |  |  |  | RSSI & AINF: 9 | 09 |

| Telegram structure: sWN LMDscandatacfg (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  | LMS4000 | No values: 0 Remission only: 1 Angle only: 2 Remission & Angle: 3 Quality only: 4 Remission & Quality: 5 Angle & Quality: 6 Remission & Angle & Quality: 7 | 00 01 02 03 04 05 06 07 |
| Resolution | Resolution of remission data 1) | Enum_8 | 1 | All | 8 Bit: 0 16Bit: 1 | 00 01 |
|  |  |  |  | LMS4000 | Always: 1 | 01 |
| Unit | Unit of remission data | Enum_8 | 1 | All | Digits: 0 | 00 |
|  |  |  |  | LMS4000 | Digits (RSSI): 0 Percent (REFL): 1 | 00 01 |
| Encoder | Encoder data | Uint_8 | 2 | LMS1xx LMS5xx LMS4000 LRS4000 | No encoder: 0 0 Channel 1: 1 0 | 00 00 01 00 |
|  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx TiMxxx MRS1000 LMS1000 MRS6000 | No encoder: 0 0 | No encoder: 00 00 |
| Position | Position values | Bool_1 | 1 | All | No: 0 Yes: 1 | 00 01 |
| Device name | Sends the device name | Bool_1 | 1 | All | No: 0 Yes: 1 | 00 01 |
| Comment | Saved comment | Bool_1 | 1 | All | No: 0 Yes: 1 | 00 01 |
| Time | Sends time information | Bool_1 | 1 | All | No: 0 Yes: 1 | 00 01 |
|  |  |  |  | TiM24x | No: 0 | 00 |

1) 	LMS5xx since V1.10, 8 bit only. ; MRS1000/LMS1000 8bit only

| Telegram structure: sWN LMDscandatacfg (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Output rate | Sends the output rate | Uint_16 | 2 | LMS1xx LMS5xx TiMxxx LMS4000 | All scans: +1d (1h) Each 2nd scan: +2d (2h) Each 50000th scan: + 50000d (C350h) | 00 01 00 02 C3 50 |
|  |  |  |  | MRS1000 LMS1000 | All scans: +1d (1h) | 00 01 |
|  |  |  |  | MRS6000 TiM240 LRS4000 | All scans: +1d (1h) Each 2nd scan: +2d (2h) ... Max: Each 100th Scan: + 100d (64h) | 00 01 00 02 ... 00 64 |
|  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | All scans: +1d (1h) Each 2nd scan: +2d (2h) Each 200th scan: +200d (C8h) | 00 01 00 02 00C8 |

Table 128: Telegram structure: sWN LMDscandatacfg

Example 1: output channel 1, remission, resolution 16Bit, no encoder and all scans

| CoLa A | ASCII | <STX> sWN {SPC} LMDscandatacfg {SPC} 01 {SPC} 00 {SPC} 1 {SPC} 1 {SPC} 0 {SPC} 00 {SPC} 00 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} +1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 30 31 20 30 30 20 31 20 31 20 30 20 30 30 20 30 30 20 30 20 30 20 30 20 30 20 2B 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 20 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 20 01 00 01 01 00 00 00 00 00 00 00 00 01 43 |

Table 129: Example 1: sWN LMDscandatacfg

Example 2: output channel 1, remission, resolution 16Bit, no encoder, each 10th scan

| CoLa A | ASCII | <STX> sWN {SPC} LMDscandatacfg {SPC} 01 {SPC} 00 {SPC} 1 {SPC} 1 {SPC} 0 {SPC} 00 {SPC} 00 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} +10 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 30 31 20 30 30 20 30 20 31 20 30 20 30 30 20 30 30 20 30 20 30 20 30 20 30 20 2B 31 30 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 20 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 20 01 00 00 01 00 00 00 00 00 00 00 00 10 52 |

Table 130: Example 2: sWN LMDscandatacfg

Example 3: output channel 2, resolution 16Bit, encoder active, each 10th scan

| CoLa A | ASCII | <STX> sWN {SPC} LMDscandatacfg {SPC} 02 {SPC} 00 {SPC} 0 {SPC} 1 {SPC} 0 {SPC} 01 {SPC} 00 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} +10 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 30 32 20 30 20 30 20 31 20 30 20 30 31 20 30 20 30 20 30 20 30 20 30 20 2B 31 30 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 20 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 20 02 00 00 01 00 01 00 00 00 00 00 00 4A 63 |

Table 131: Example3: sWN LMDscandatacfg

| Telegram structure: sWA LMDscandatacfg |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Configure scandata | String | 14 | All | LMDscandatacfg | 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 |

Table 132: Telegram structure: sWA LMDscandatacfg

Example: sWA LMDscandatacfg

| CoLa A | ASCII | <STX> sWA {SPC} LMDscandatacfg <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 63 66 67 20 4D |

Table 133: Example: sWA LMDscandatacfg

    - Configure measurement angle of the scandata for output

| Telegram structure: sWN LMPoutputRange (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Change output angle range | String | 14 | All | LMPoutputRange | 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 |
| Status code | Length | Int_16 | 2 | All | 1 | 00 01 |
| Angular resolu- tion 2) | [1/10000 °] | Uint_32 | 4 | LMS1xx | 0.25 °: + 2500d (9 C4h) 0.5: + 5000d (1388h) | 0.25: 00 00 09 C4 0.5: 00 00 13 88 |
|  |  |  |  | LMS5xx | 0.1667: + 1667d (683h) 0.25 °: + 2500d (9 C4h) 0.333: + 3333d (D05h) 0.5: + 5000d (1388h) 0.667: + 6667d (1 A0Bh) 1: + 10000d (2710h) | 0.1667: 00 00 06 83 0.25: 00 00 09 C4 0.333: 00 00 0D 05 0.5: 00 00 13 88 0.667: 00 00 1A 0B 1: 00 00 27 10 |
|  |  |  |  | TiMxxx | 0.333: + 3333d (D05h) 1: + 10000d (2710h) | 0.333: 00 00 0D 05 1: 00 00 27 10 |
|  |  |  |  | TiM24x | 1: + 10000d (2710h) | 1: 00 00 27 10 |
|  |  |  |  | MRS1000 | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  |  |  | LMS1000 | 0.75: + 7500d (1 D4Ch) |  |
|  |  |  |  | MRS 6000 | 0.13: + 1300d (515h) | 0.13: 00 00 05 15 |
|  |  |  |  | LMS 4000 | 0.0833: +833 (341h) | 0.0833: 00 00 03 41 |
|  |  |  |  | LRS4000 | @ 12.5Hz: 0.0200: + 200d (C8h) 0.0400: + 400d (190h) 0.0600: + 600d (258h) | @ 12,5Hz: 0.0200: 00 00 C8 0.0400: 00 00 01 90 0.0600: 00 00 02 58 |

2) 		Note: Angular resolution can not be changed here, it is taken automatically from the basic scan settings!The angular resolution is not exactly 0.1667 degree, and this value should not be used for calculations. The result is an angular resolution of 0,1‾6 or 1/6 of a degree (six measurements per degree). When used for calculations a customer should recover the real value, e.g. by double AngRes = 2.0 / round(2.0 / GivenAngRes).

| Telegram structure: sWN LMPoutputRange (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | 0.1000: + 1000d (3 E8h) 0.1200: + 1200d (4 B0h) | 0.1000: 00 00 03 E8 0.1200: 00 00 04 B0 |
| Start angle | [1/10000 °] | Int_32 | 4 | LMS1xx TiMxxx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  | TiM24x | -1200000d... + 1200000d (FFEDB080h... 124 F80h) | FF ED B0 80&hellip; 00 12 4F 80 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  | MRS1000 | -475000d... + 2275000d (FFF8C088h... 22 B6B8h) | FF F8 C0 88&hellip; 00 22 B6 B8 |
|  |  |  |  | LMS1000 | -480000d (FFF8AD00h) |  |
|  |  |  |  | MRS 6000 | 30000d... 1500000d (493 E0h... 16 E360h) | 00 04 93 E0... 00 16 E3 60 |
|  |  |  |  | LMS 4000 | +550000d&hellip; +1250000d (86470h &hellip; 1312D0h) | 00 08 64 70... 00 13 12 D0 |
|  |  |  |  | LRS4000 | -1800000d ( - FFE488C0h) | FF E4 88 C0 |
| Stop angle | [1/10000 °] | Int_32 | 4 | LMS1xx TiMxxx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  | TiM24x | -1200000d... + 1200000d (FFEDB080h... 124 F80h) | FF ED B0 80&hellip; 00 12 4F 80 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  | MRS1000 | -475000d... + 2275000d (FFF8C088h... 22 B6B8h) | FF F8 C0 88&hellip; 00 22 B6 B8 |
|  |  |  |  | LMS1000 | + 2280000d (22 CA40h) |  |
|  |  |  |  | MRS 6000 | 30000d... 1500000d (493 E0h... 16 E360h) | 00 04 93 E0h... 00 16 E3 60h |
|  |  |  |  | LMS 4000 | +550000d&hellip; +1250000d (86470h &hellip; 1312D0h) | 00 08 64 70... 00 13 12 D0 |
|  |  |  |  | LRS4000 | + 1800000d (1 B7740h) | 00 1B 7740 |

Table 134: Telegram structure: sWN LMPoutputRange

Example: sWN LMPoutputRange 0,50°resolution, 0°-90°

| CoLa A | ASCII | <STX> sWN {SPC} LMPoutputRange {SPC} 1 {SPC} 1388 {SPC} 0 {SPC} DBBA0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 20 31 20 31 33 38 38 20 30 20 44 42 42 41 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 21 73 57 4E 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 00 01 00 00 13 88 00 00 00 00 00 0D BB A0 F7 |

Table 135: Example: sWN LMPoutputRange 0,50°resolution, 0°-90°

| Telegram structure: sWA LMPoutputRange |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Store parameters | String | 14 | All | LMPoutputRange | 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 |

Table 136: Telegram structure: sWA LMPoutputRange

Example: sWA LMPoutputRange

| CoLa A | ASCII | <STX> sWA {SPC} LMPoutputRange <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 41 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 20 74 |

Table 137: Example: sWA LMPoutputRange

    - Read for actual output range

| Telegram structure: sRN LMPoutputRange |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Output range | String | 14 | All | LMPoutputRange | 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 |

Table 138: Telegram structure: sRN LMPoutputRange

Example: sRN LMPoutputRange

| CoLa A | ASCII | <STX> sRN {SPC} LMPoutputRange <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 12 73 52 4E 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 5E |

Table 139: Example: sRN LMPoutputRange

| Telegram structure: sRA LMPoutputRange |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Output range | String | 14 | All | LMPoutputRange | 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 |

| Telegram structure: sRA LMPoutputRange |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Number of sectors | Indicates the number of sectors. The subsequent values will be transmitted 1&hellip; 4 accordingly. | Int_16 | 2 | All | Sector 1: 0001h | Sector 1: 0001 |
| Angular resolution | [1/10000 °] | Uint_32 | 4 | LMS1xx | 0.25 °: + 2500d (9 C4h) 0.5: + 5000d (1388h) | 0.25: 00 00 09 C4 0.5: 00 00 13 88 |
|  |  |  |  | LMS5xx | 0.1667: + 1667d (683h) 0.25 °: + 2500d (9 C4h) 0.333: + 3333d (D05h) 0.5: + 5000d (1388h) 0.667: + 6667d (1 A0Bh) 1: + 10000d (2710h) | 0.1667: 00 00 06 83 0.25: 00 00 09 C4 0.333: 00 00 0D 05 0.5: 00 00 13 88 0.667: 00 00 1A 0B 1: 00 00 27 10 |
|  |  |  |  | TiMxxx | 0.333: + 3333d (D05h) 1: + 10000d (2710h) | 0.333: 00 00 0D 05 1: 00 00 27 10 |
|  |  |  |  | TiM24x | 1: + 10000d (2710h) | 1: 00 00 27 10 |
|  |  |  |  | MRS1000 | 0.25 °: + 2500d (9 C4h) | 0.25: 00 00 09 C4 |
|  |  |  |  | LMS1000 | 0.75: + 7500d (1 D4Ch) |  |
|  |  |  |  | MRS 6000 | 0.13: + 1300d (515h) | 0.13: 00 00 05 15 |
|  |  |  |  | LMS 4000 | 0.0833: +833 (341h) | 0.0833: 00 00 03 41 |
|  |  |  |  | LRS4000 | @ 12.5Hz: 0.0200: + 200d (C8h) 0.0400: + 400d (190h) 0.0600: + 600d (258h) 0.1000: + 1000d (3 E8h) 0.1200: + 1200d (4 B0h) @ 25Hz: 0.0400: + 400d (190h) 0.0800 + 800d (320h) 0.1200: + 1200d (4 B0h) 0.2000: + 2000d (7 D0h) 0.2400: + 2400d (960h) | @ 12,5Hz: 0.0200: 00 00 C8 0.0400: 00 00 01 90 0.0600: 00 00 02 58 0.1000: 00 00 03 E8 0.1200: 00 00 04 B0 @ 25Hz: 0.0400: 00 00 01 90 0.0800: 00 00 03 20 0.1200: 00 00 04 B0 0.2000: 00 00 07 D0 0.2400: 00 00 09 60 |

| Telegram structure: sRA LMPoutputRange |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Start angle | [1/10000 °] | Int_32 | 4 | LMS1xx TiMxxx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  | TiM24x | -1200000d... + 1200000d (FFEDB080h... 124 F80h) | FF ED B0 80&hellip; 00 12 4F 80 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  | MRS1000 | -475000d... + 2275000d (FFF8C088h... 22 B6B8h) | FF F8 C0 88&hellip; 00 22 B6 B8 |
|  |  |  |  | LMS1000 | -480000d... + 2280000d (FFF8AD00h... 22 CA40h) |  |
|  |  |  |  | MRS 6000 | 30000d... 1500000d (493 E0h... 16 E360h) | 00 04 93 E0h... 00 16 E3 60h |
|  |  |  |  | LMS 4000 | +550000d&hellip; +1250000d (86470h &hellip; 1312D0h) | 00 08 64 70... 00 13 12 D0 |
|  |  |  |  | LRS4000 | -1800000d (FFE488C0h) | FF E4 88 C0 |
| Stop angle | [1/10000 °] | Int_32 | 4 | LMS1xx TiMxxx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  | TiM24x | -1200000d... + 1200000d (FFEDB080h... 124 F80h) | FF ED B0 80&hellip; 00 12 4F 80 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  | MRS1000 | -475000d... + 2275000d (FFF8C088h... 22 B6B8h) | FF F8 C0 88&hellip; 00 22 B6 B8 |
|  |  |  |  | LMS1000 | -480000d... + 2280000d (FFF8AD00h... 22 CA40h) |  |
|  |  |  |  | MRS 6000 | 30000d... 1500000d (493 E0h... 16 E360h) | 00 04 93 E0h... 00 16 E3 60h |
|  |  |  |  | LMS 4000 | +550000d&hellip; +1250000d (86470h &hellip; 1312D0h) | 00 08 64 70... 00 13 12 D0 |
|  |  |  |  | LRS4000 | + 1800000d (1 B7740h) | 00 1B 7740 |

Table 140: Telegram structure: sRA LMPoutputRange

Example: sRA LMPoutputRange

| CoLa A | ASCII | <STX> sRA {SPC} LMPoutputRange {SPC} 1 {SPC} 1388 {SPC} FFF92230 {SPC} 225510 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 31 20 31 33 38 38 20 46 46 46 39 32 32 33 30 20 32 32 35 35 31 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 21 73 52 41 20 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 00 01 00 00 13 20 88FF F9 22 30 00 22 55 10 98 |

Table 141: Example: sRA LMPoutputRange

    - Configure the distance scaling of the scandata for the output

| Telegram structure: sWN LMDscandatascalefactor (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Change distance scale factor | String | 14 | All | LMDscandatascalefactor | 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 |
| ScaleFactor | [mm/digit] | Real as float accordin | 4 | LRS4000 | Factor × 1: 3F800000h Factor × 2 (values have to be scaled by factor two): 40000000h (Default) MinValue="0.1" MaxValue="16.0" Default "2.0” PhysicalUnit="mm/digit" | 3F 80 00 00 40 00 00 00 00 |
|  |  | g |  |  |  |  |
|  |  | IEEE754 |  |  |  |  |

Table 142: Telegram structure: sWN LMDscandatascalefactor

Example:

ScaleFactor=1 : Distance value represented [16bit] 1mm per digit ; Maximum value 2^16=65535 mm. ScaleFactor=2: Distance value represented [16bit] : 2mm per digit ; Maximum vale / range = 131072mm.

Example: sWN LMDscandatascalefactor 1.0°

| CoLa A | ASCII | <STX> sWN {SPC} LMDscandatascalefactor {SPC} 3F800000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 20 3 F800000 03 |
| CoLa B | Binary | 02 02 73 02 00 00 00 23 73 57 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 20 3F 80 00 00 00 28 |

Table 143: Example: sWN LMDscandatascalefactor 1,0

| Telegram structure: sWA LMDscandatascalefactor |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Change distance scale factor | String | 14 | All | LMDscandatascalefactor | 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 |

Table 144: Telegram structure: sWA LMDscandatascalefactor

Example: sWA LMDscandatascalefactor

| CoLa A | ASCII | <STX> sWA {SPC} LMDscandatascalefactor <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 03 |
| CoLa B | Binary | 02 02 02 02 00 00 1A 73 57 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 7A |

Table 145: Example: sWA LMDscandatascalefactor

    - Read the distance scaling for the distance values of the scandata

| Telegram structure: sRN LMDscandatascalefactor |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Change distance scale factor | String | 14 | All | LMDscandatascalefactor | 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 |

Table 146: Telegram structure: sRN LMDscandatascalefactor

Example: sRN LMDscandatascalefactor

| CoLa A | ASCII | <STX> sRN {SPC} LMDscandatascalefactor ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1A 73 52 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 70 |

Table 147: Example: sRN LMPoutputRange

| Telegram structure: sRA LMDscandatascalefactor |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Distance Scale Factor | String | 14 | All | LMDscandatascalefactor | 4C 4D 50 6F 75 74 70 75 74 52 61 6E 67 65 |
| ScaleFactor | [mm/digit] | Real as float accordin | 4 | All | Factor × 1: 3F800000h Factor × 2 (values have to be scaled by factor two): 40000000h (Default) MinValue="0.1" MaxValue="16.0" Defaul "1.0” PhysicalUnit="mm/digit" | 3F 80 00 00 40 00 00 00 00 |
|  |  | g |  |  |  |  |
|  |  | IEEE754 |  |  |  |  |

Table 148: Telegram structure: sRA LMDscandatascalefactor

Example: sRA LMDscandatascalefactor

| CoLa A | ASCII | <STX> sRA {SPC} LMDscandatascalefactor {SPC}3F800000<ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 20 3F 80 00 00 03 |
| CoLa B | Binary | 02 02 02 00 00 00 23 73 52 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 73 63 61 6C 65 66 61 63 74 6F 72 20 3F 80 00 00 03 22 |

Table 149: Example: sRA LMDscandatascalefactor

    - Poll one telegram

Output of values from last scan.

Asking the device for the measurement values of the last valid scan. The device will respond, even if it is not running at the moment.

NOTE

After changing the scanning frequency, there will be no data telegram or answer from the devices LMS1xx, LMS5xx and TiMxxx for up to 30 seconds. The same applies when the device is powering up or rebooting.

| Telegram structure: sRN LMDscandata |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Only one telegram | String | 11 | All | LMDscandata | 4C 4D 44 73 63 61 6E 64 61 74 61 |

Table 150: Telegram structure: sRN LMDscandata

Example: sRN LMDscandata

| C oLa A | ASCII | <STX> sRN {SPC} LMDscandata <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 52 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 05 |

Table 151: Example: sRN LMDscandata

| Telegram structure: sRA LMDscandata |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Find complete telegram structure of the answer in section 4.3.7 „ Send data permanent "on page 92. |  |  |  |  |  |  |

Table 152: Telegram structure: sRA LMDscandata

Example: sRA LMDscandata

| CoLa A | ASCII | No ASCII answer possible. |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 20 01 20 01 20 89 C9 97 20 00 20 00 20 1A AE 1A B1 20 58 1C BC 15 20 58 1D 15 3D 20 00 20 00 20 07 20 00 20 00 20 13 88 20 15 20 F6 20 F9 20 F5 20 EF 20F6 20 F2 20 EF 20 ED 20 F5 20 E9 20 F2 20 FA 20 FC 20 FF 20 F1 20 F2 20 01 07 20 FC 20 FC 20 01 02 20 FF 20 00 20 00 20 00 20 00 20 00 20 00 03 |
| CoLa B | Binary | Find complete telegram structure of the answer in section 4.3.7 „ Send data permanent "on page 92. |

Table 153: Example: sRA LMDscandata

    - Send data permanently

NOTE

After changing the scanning frequency, there will be no data telegram or answer from the devices LMS1xx, LMS5xx and TiMxxx for up to 30 seconds.The same applies when the device is powering up or rebooting.

| Telegram structure: sEN LMDscandata |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Event | String | 3 | All | sEN | 73 45 4E |
| Command | Data telegram | String | 11 | All | LMDscandata | 4C 4D 44 73 63 61 6E 64 61 74 61 |
| Measure- ment | Start/stop | Enum_8 | 1 | All | Stop: 0 Start: 1 | Stop: 00 Start: 01 |

Table 154: Telegram structure: sEN LMDscandata

Example: sEN LMDscandata

| CoLa A | ASCII | <STX> sEN {SPC} LMDscandata {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 45 4E 20 4C 4D 44 73 63 61 6E 64 61 74 61 20 01 33 |

Table 155: Example: sEN LMDscandata

| Telegram structure: sEA LMDscandata |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sEA | 73 45 41 |
| Command | Data telegram | String | 11 | All | LMDscandata | 4C 4D 44 73 63 61 6E 64 61 74 61 |
| Measure- ment | Start/stop | Enum_8 | 1 | All | Stop: 0 Start: 1 | Stop: 00 Start: 01 |

Table 156: Telegram structure: sEA LMDscandata

Example: Confirmation of sEA LMDscandata

| CoLa A | ASCII | <STX> sEA {SPC} LMDscandata {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 45 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 01 33 |

Table 157: Example: Confirmation of sEA LMDscandata

Telegram stream

The answer to the telegram will be followed by the scandata:

NOTE

Leading zeros of a value will not be displayed in ASCII.

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRA sSN | 73 52 41 73 53 4E |
| Command | Data telegram | String | 11 | All | LMDscandata | 4C 4D 44 73 63 61 6E 64 61 74 61 |
| Version number | For detecting for- mat changes by the version. Ver- sion is always 0 or 1 up to now. | Uint_16 | 2 | All | 1h... FFFFh | 00 01... FF FF |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Device | Device number | Defined with SOPAS | Uint_16 | 2 | All | 0h... FFFFh | 00 00... FF FF |
|  | Serial number | Defined in factory | Uint_32 | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Device status | (See values column) | Uint_8 | 2 × 1 | All | OK: 0 0 Error: 0 1 | 00 00 (0000 0000) 00 01 (0000 0001) |
|  |  |  |  |  | LMS1xx LMS5xx | Pollution warning: 0 2 Pollution error with no device error: 0 4 Pollution error with device error: 0 5 | 00 02 (0000 0010) 00 04 (0000 0100) 00 05 (0000 0101) |
|  |  |  |  |  | LMS4000 | Data output during laser off: 1 0 | 01 00 |
| Status info | Telegram counter | Number of measurement telegrams finished in the scanner and given to the interface. 3) | Uint_16 | 2 | All | 0h... FFFFh | 00 00... FF FF |
|  | Scan counter | Number of scans which were crea- ted in the device; counts how many scans were really done. | Uint_16 | 2 | All | 0h... FFFFh | 00 00... FF FF |
|  | Time since start up in µs | Counting the time since power up the device; starting with 0. In the output telegram this is the time at the zero index before the measurement itself starts. | Uint_32 | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Time of transmis- sion in µs | Time in µs when the complete scan is transmitted to the buffer for data output; starting with 0 at scanner bootup. | Uint_32 | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Status of digital in- puts | Low byte represents input 1. | Uint_8 | 2 × 1 | LMS1xx LMS5xx | All inputs low: 0 0 All inputs high: 3 0 | 00 00 03 00 |
|  |  |  |  |  | LMS4000 | All inputs low: 0 0 IN1 high: 1 0 IN2 high: 2 0 IN1+2 high: 3 0 | 00 00 01 00 02 00 03 00 |

3) 	Does not count how many telegrams were really given out; is relevant if not all scans are delivered from the scan core.

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | TiM24x MRS 1000 LMS1000 MRS6000 LRS4000 | 0 0 | 00 00 |
|  | Status of digital outputs | Low byte repre- sents output 1. | Uint_8 | 2 × 1 | All | All outputs low: 0 0 | 00 00 |
|  |  |  |  |  | TiM5xx, TiM7xx | All outputs high: F 0 | 0F 00 |
|  |  |  |  |  | TiM24x | All outputs high: 2 0 | 02 00 |
|  |  |  |  |  | LMS1xx | All internal outputs high: 7 0 All outputs high (inkl. Ext. Out): 07 FF | 07 00 07FF |
|  |  |  |  |  | LMS5xx | All internal outputs high: 3F 0 All outputs high (inkl. Ext. Out): 3F FF | 3F 00 3F FF |
|  |  |  |  |  | LDxxx | All outputs high: F 0 | 0F 00 |
|  |  |  |  |  | LMS4000 | OUT1 high: 1 0 OUT2 high: 2 0 OUT1+2 high: 3 0 OUT4 high: 8 0 OUT1+4 high: 9 0 OUT2+4 high: A 0 OUT1+2+4 high: B 0 | 01 00 02 00 03 00 08 00 09 00 0A 00 0B 00 |
|  |  |  |  |  | LMS1000, MRS1000, MRS6000 LRS4000 | Always: 0 | 00 00 |
|  | former Reserved now Layer angle. | - | Uint_16 | 2 | All | 0 | 0 |
|  |  |  |  |  | MRS1000 | 0  0 Layer2 FF06  -250Layer3 FA 250 Layer1 FE0C  -500Layer4 (value 1/100) | 00 00 00 00 00 46 46 30 36 00 00 46 41 46 45 30 43 |
|  |  |  | Int_16 |  | MRS6000 | Angle = value / 200 Example: F5B2h -2638/200 = -13.19° Layer 24 EDh237/200= 1.185°Layer 1 Range: -13.19°~1.185° (each layer is 0.625°) | Angle = value / 200 Example: F5 B2Layer 24 00EDLayer 1 |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Frequencies | Scan fre- quency | [1/100 Hz] | Uint_32 | 4 | LMS1xx | 25Hz: +2500d (9C4h) 50Hz: +5000d (1388h) | 09C4 13 88 |
|  |  |  |  |  | LMS5xx | 25Hz: +2500d (9C4h) 35Hz: +3500d (DACh) 50Hz: +5000d (1388h) 75Hz: +7500d (1A0Bh) 100Hz: +10000d (2710h) | 09C4 0D AC 13 88 1A 0B 27 10 |
|  |  |  |  |  | TiMxxx | 15Hz: +1500d (5DCh) | 05DC |
|  |  |  |  |  | TiM24x | 14.5Hz: +1450d (5AAh) | 05AA |
|  |  |  |  |  | NAV310 LD-OEM 15xx | 5Hz... 20Hz: + 500d... + 2000d (1 F4h... 7 D0h) | 01F4... 07 D0 |
|  |  |  |  |  | LD-LRS 36xx | 5Hz... 15Hz: + 500d... +1500 (1 F4h &hellip; 5DCh)) | 01F4... 05 DC |
|  |  |  |  |  | MRS1000 | 50Hz: +5000d (1388h) | 50Hz: 13 88 |
|  |  |  |  |  | LMS1000 | 150Hz: +15000d (3A98h) |  |
|  |  |  |  |  | MRS6000 | 10Hz: +1000d (3E8h) | 10Hz: 03 E8 |
|  |  |  |  |  | LMS4000 | 600Hz: +60000d (EA60h) | 600Hz: EA 60 |
|  |  |  |  |  | LRS4000 | 12.5Hz:+ 1250d (4B0) 25Hz: +2500d (9C4) | 12.5Hz:00 00 04 B0 25Hz: 00 00 09 C4 |
|  | Measure- ment fre- quency | Inverse of the time between two measurement shots (in 100 Hz), Example: 50Hz 0.5° resolution  720 shots/20 ms  36 kHz | Uint_32 | 4 | All | 00000000h&hellip; FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  |  |  |  |  | LMS4000 | 21C0h | 00 00 21 C0 |
| Amount of encoder |  | If 0, then next two values are missing. | Enum_1 6 | 2 | TiMxxx MRS1000 LMS1000 MRS6000 | Always: 0 | Always: 00 00 |
|  |  |  |  |  | LMS1xx LMS5xx | 0... 3 | 00 00... 00 03 |
|  |  |  |  |  | LMS4000 LRS4000 | 0 External encoder: 1 | 00 00 00 01 |
| Values | Encoder position | Info in ticks | Uint_32 | 4 | LMS1xx LMS5xx LMS4000 LRS4000 | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | Encoder speed | Ticks/mm | Uint_16 | 2 | LMS1xx LMS5xx LRS4000 | 0h... FFFFh | 00 00... FF FF |
|  |  |  |  |  | LMS4000 | Always: 0 | Always: 00 00 |
| Amount of 16 bit channels |  | Number of 16 bit channels that provide measured data | Uint_16 | 2 | TiMxxx | Output channel: 1 | Output channel: 01 |
|  |  |  |  |  | LMS1xx | Output channels: 1, 2 or 4 | Output channels: 01, 02 or 04 |
|  |  |  |  |  | LMS5xx | Output channels: 1 or 5 | Output channels: 01 or 05 |
|  |  |  |  |  | MRS1000 LMS1000 | Output channels: 1 or 3 | Output channels: 01 or 03 |
|  |  |  |  |  | MRS6000 | Output channels:1..9 | Output channels:01..09 |
|  |  |  |  |  | LMS4000 | Output channels: 0&hellip;3 | Output channels: 00&hellip;03 |
|  |  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | Depending on amount of sectors and selection of output of distance or distance and remission RSSI Example (2 sectors): If 2 channels: sectors 1 + 2 contain Dist1 If 4 channels: sectors 1+2 contain Dist + RSSI1 | Depending on amount of sectors and selection of output of distance or distance and remission RSSI Example (2 sectors): If 2 channels: sectors 1+2 contain Dist1 If 4 channels: sectors 1+2 contain Dist + RSSI1 |
|  |  |  |  |  | LRS4000 | Output Channels: 1,2,3,6 (1 or 3, with or without RSSI) | Output channel: 01 |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Output channel (16 bit) | Content | Defines the con- tent of the output channel Unit of radial distance values (DIST) is mm | String | 5 | LMS1xx | DIST1: Distance values of first pulse DIST2: Distance valus second pulse RSSI1: Energy values of first pulse RSSI2: Energy values second pulse | 44 49 53 54 31 44 49 53 54 32 52 53 53 49 31 52 53 53 49 32 |
|  |  |  |  |  | LMS5xx (with Software &ge; V1.10 only) | DIST1: Distance values of first pulse DIST2: Distance valus second pulse DIST3: Distance values of third pulse DIST4: Distance valus fourth pulse DIST5: Distance values of fifth pulse RSSI1: Energy values of first pulse RSSI2: Energy values second pulse RSSI3: Energy values of third pulse RSSI4: Energy values of forth pulse RSSI5: Energy values of fifth pulse | 44 49 53 54 31 44 49 53 54 32 44 49 53 54 33 44 49 53 54 34 44 49 53 54 35 52 53 53 49 31 52 53 53 49 32 52 53 53 49 33 52 53 53 49 34 52 53 53 49 35 |
|  |  |  |  |  | TiMxxx | DIST1: Distance values | 44 49 53 54 31 |
|  |  |  |  |  | MRS1000 LMS1000 | DIST1: Distance values DIST2: Distance values DIST3: Distance values RSSI1: Energy values RSSI2: Energy values RSSI3: Energy values | 44 49 53 54 31 44 49 53 54 32 44 49 53 54 33 52 53 53 49 31 52 53 53 49 32 52 53 53 49 33 |

98 	TECHNICAL INFORMATION | DETECTION AND RANGING SOLUTIONS 	8014631/1FHD/2022-04-27 | SICK

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | MRS6000 | DIST1: Distance values DIST2: Distance values DIST3: Distance values DIST4: Distance values RSSI1: Energy values RSSI2: Energy values RSSI3: Energy values RSSI4: Energy values VANGL: Vertical Angle | 44 49 53 54 31 44 49 53 54 32 44 49 53 54 33 44 49 53 54 34 52 53 53 49 31 52 53 53 49 32 52 53 53 49 33 52 53 53 49 34 56 41 4E 47 4C |
|  |  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | DIST1: Distance values RSSI1: Energy values | 44 49 53 54 31 52 53 53 49 31 |
|  |  |  |  |  | LMS4000 | DIST1: Distance values RSSI1: Energy values in digit REFL1: Calibrated energy values in percent ANGL1: Angle Offset values | 44 49 53 54 31 52 53 53 49 31 52 45 46 4C 31 41 4E 47 4C 31 |
|  |  |  |  |  | LRS4000 | DIST1: Distance values DIST2: Distance values DIST3: Distance values RSSI1: Energy values RSSI2: Energy values RSSI3: Energy values |  |
|  | Scale factor | Scale factor or factor of the measurement values (for the LMS5xx this depends on the angular resolution) | Real as float accordin g to IEEE754 | 4 | LMS1xx LMS5xx TiMxxx MRS 1000 LMS 1000 | Factor × 1: 3F800000h Factor × 2: 40000000h | 3F 80 00 00 40 00 00 00 00 |
|  |  |  |  |  | MRS6000 | Factor x 12.5: 41480000h Factor × 1: 3F800000h Factor x -0.00025: B983126Fh | 41 48 00 00 3F 80 00 00 B9 83 12 6F |
|  |  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | Factor × 4: 40800000h | 04 08 00 00 |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | LMS4000 | Factor x 0.1: 3DCCCDDh (DIST1) Factor × 1: 3F800000h (RSSI1) Factor × 1: 3F800000h (REFL1) Factor × 1: 3F800000h (ANGL1) | 3D CC CC CD 3F 80 00 00 3F 80 00 00 3F 80 00 00 |
|  |  |  |  |  | LRS4000 | Factor x 2 40000000h (0.1 .. 16.0 (Default 2.0) | 40 00 00 00 00 |
|  | Scale factor offset | Sets starting point of measurement | Real as float accordin g to IEEE754 | 4 | LMS1xx LMS5xx TiMxxx MRS 1000 LMS 1000 | 0 | 00 00 00 00 00 |
|  |  |  |  |  | MRS6000 | Offset 0 : 0 Offset 1.5: 3FC00000 (1.5 offset for VANGL) | 00 00 00 00 00 3F C0 00 00 (1.5 offset for VANGL) |
|  |  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | 00000000h&hellip; FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  |  |  |  |  | LMS4000 | DIST1: 0 RSSI1: 0 REFL1: 0 ANGL1: C7000000h (-32768=3.2768 °) | 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 C7 00 00 00 00 |
|  | Start angle | [1/10000 °] | Uint_32 | 4 | LMS1xx TiMxxx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | TiM24x | -1200000d... + 1200000d (FFEDB080h... 124 F80h) | FF ED B0 80&hellip; 00 12 4F 80 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  |  | MRS 1000 | -475000d... + 2275000d (FFF8C088h... 22 B6B8h) | FF F8 C0 88&hellip; 00 22 B6 B8 |
|  |  |  |  |  | LMS 1000 | -480000d... + 2280000d (FFF8AD00h... 22 CA40h) | FF F8 AD 00&hellip; 00 22 CA 40 |
|  |  |  |  |  | MRS 6000 | + 30000d... + 1500000d (493 E0h... 16 E360h) | 00 04 93 E0h... 00 16 E3 60h |
|  |  |  |  |  | NAV310 LD-OEM 15x1 LD-LRS 36x1 | 0d... + 3600000d (0h &hellip; 36 EE80h) | 00 00 00 00 00... 00 36 EE 80 |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | LD-OEM 15x0 LD-LRS 36x0 | -900000d... + 2700000d (FFF24460h &hellip; 41 EB0h) | FF F2 44 60&hellip; 00 04 1E B0 |
|  |  |  |  |  | LMS4000 | +550000d&hellip; +1250000d (86470h &hellip; 1312D0h) | 00 08 64 70... 00 13 12 D0 |
|  |  |  |  |  | LRS4000 | 1800000d .. + 1800000d (1 B7740h &hellip; FFE488C0h) | 00 1B 77 40 FF E4 88 C0 |
|  | Size of single angular step | Output format in degree: 1/10000° | Uint_16 | 2 | LMS1xx | + 2500d... + 5000d (9 C4h &hellip; 1388h) | 09C4... 13 88 |
|  |  |  |  |  | LMS5xx | + 1667d... + 10000d (683h... 2710h) | 06 83... 27 10 |
|  |  |  |  |  | TiMxxx | + 333d... + 10000d (D05h &hellip; 2710h) | 0D 05... 27 10 |
|  |  |  |  |  | TiM24x | + 10000d (2710h) | 27 10 |
|  |  |  |  |  | MRS1000 | + 2500d (9 C4h) | 09C4 |
|  |  |  |  |  | LMS1000 | + 7500d (1 D4Ch) | 1D 4C |
|  |  |  |  |  | MRS6000 | + 1300d (= 0.13) (515h) | 05 15 |
|  |  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | 0.125 °... 1° def. 0.25° + 1250d... + 10000d (4 E2h &hellip; 2710h) (Default: 09C4h = 0.25°) | 04E2... 27 10 (Default: 09 C4) |
|  |  |  |  |  | LMS4000 | 0.0833: +833 (341h) | 0.0833: 03 41 |
|  |  |  |  |  | LRS4000 | @ 12.5Hz: 0.0200: + 200d (C8h) 0.0400: + 400d (190h) 0.0600: + 600d (258h) 0.1000: + 1000d (3 E8h) 0.1200: + 1200d (4 B0h) @ 25Hz: 0.0400: + 400d (190h) 0.0800 + 800d (320h) 0.1200: + 1200d (4 B0h) 0.2000: + 2000d (7 D0h) 0.2400: + 2400d (960h) | @ 12,5Hz: 0.0200: 00 00 C8 0.0400: 00 00 01 90 0.0600: 00 00 02 58 0.1000: 00 00 03 E8 0.1200: 00 00 04 B0 @ 25Hz: 0.0400: 00 00 01 90 0.0800: 00 00 03 20 0.1200: 00 00 04 B0 0.2000: 00 00 07 D0 0.2400: 00 00 09 60 |
|  | Amount of data | Defines the num- ber of items on measured output | Uint_16 | 2 | All | 0h... FFFFh | 00 00... FF FF |
|  | Data_1 Data_n | Data stream starting Data_1 to Data_n 4) | Uint_16 | 2 | LMS100 | 0h... 4E20h | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  |  | LMS150 | 0h... C350h | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | 0h... FDE8h | 00 00 00 00 00... 00 00 FD E8 |

4) 	LMS1xx without limit.

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | TiMxxx | 0h... 61A8h | 00 00 00 00 00... 00 00 61 A8 |
|  |  |  |  |  | MRS 1000 LMS 1000 | 0h... FA00h | 00 00 00 00 .. 00 00 FA 00 |
|  |  |  |  |  | MRS 6000 | 0h&hellip;8CA0h(DIST) 0h&hellip;FFFFh(RSSI) EDh&hellip;F5B2h(VANGL) | 00 00&hellip;8C A0 (DIST) 00 00&hellip;FF FF( RSSI) 00ED&hellip;F5 B2 (VANGL) |
|  |  |  |  |  | LMS 4000 | 0h&hellip;FFFFh (DIST1) 0h&hellip;FFFFh (RSSI1) 0h&hellip;FFFFh (REFL1) 0h&hellip;FFFFh (ANGL1) | 00 00 &hellip;FF FF 00 00... FF FF 00 00... FF FF 00 00... FF FF |
|  |  |  |  |  | NAV310 LD-OEM 15xx LD-LRS 36xx | 0h... 0992h | 00 00 00 00 00... 00 00 09 92 |
|  |  |  |  |  | LRS4000 | 0000h... FFFFh | 00 00... FF FF |
| For NAV310/LD-OEM15xx/LRS: The array "Output channel 16 bit "has various dimensions "Amount of 16 Bit Channels”, depending on the amount of sectors and if RSSI (output of remission values) is selected as on or off: - If RSSI was not selected (by LMDscandatacfg); there are 2 channels with the contents - Channel 1: First sector (Test target), content: DIST1 - Channel 2: Second sector (Main profile data), content: DIST1 - If RSSI was selected (by LMDscandatacfg); there are 4 channels with the contents - Channel 1: First sector (Test target), content: DIST1 - Channel 2: First sector (Test target), content: RSSI1 - Channel 3: Second sector (Main profile data), content: DIST1 - Channel 4: Second sector (Main profile data), content: RSSI1 The number behind DIST and RSSI is the order number of the pulse. As the NAV310/LD-OEM15xx/LD-LRS36xx scanner are working with a single pulse measurement, it is always "1 ”. |  |  |  |  |  |  |  |
| Amount of 8 bit channels |  | Amount of 8 bit channels, giving out the measured data | Enum_1 6 | 2 | LMS1xx | Output channels: 1 or 2 | 01 or 02 |
|  |  |  |  |  | LMS5xx | Output channels: 1 or 5 | 01 or 05 |
|  |  |  |  |  | MRS1000 LMS1000 | Output channels: 1 or 3 | 01 or 03 |
|  |  |  |  |  | LMS4000 | Output channels: 0 or 1 | 00 or 01 |
|  |  |  |  |  | TiMxxx NAV310 LD-OEM 15xx LD-LRS 36xx MRS6000 LRS4000 | Output channels: 0 | 00 |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Output channel (8 bit) | Content | Defines the content of the output channel | String | 5 | LMS1xx | DIST1 DIST2 RSSI1 RSSI2 | 44 49 53 54 31 44 49 53 54 32 52 53 53 49 31 52 53 53 49 32 |
|  |  |  |  | 5 | LMS5xx (with Software &ge; V1.10 only) | DIST1 DIST2 DIST3 DIST4 DIST5 RSSI1 RSSI2 RSSI3 RSSI4 RSSI5 | 44 49 53 54 31 44 49 53 54 32 44 49 53 54 33 44 49 53 54 34 44 49 53 54 35 52 53 53 49 31 52 53 53 49 32 52 53 53 49 33 52 53 53 49 34 52 53 53 49 35 |
|  |  |  |  | 5 | TiMxxx | DIST1 RSSI1 | 44 49 53 54 31 52 53 53 49 31 |
|  |  |  |  |  | MRS1000 LMS1000 | DIST1 DIST2 DIST3 RSSI1 RSSI2 RSSI3 AINF1 (Ambient light) | 44 49 53 54 31 44 49 53 54 32 44 49 53 54 33 52 53 53 49 31 52 53 53 49 32 52 53 53 49 33 41 49 4E 46 31 |
|  |  |  |  |  | LMS 4000 | QLTY1 | 51 4C 54 59 31 |
|  | Scale factor | Scale factor or of the measurement values (in LMS5xx depends on the angular resolution) | Real as float accordin g to IEEE754 | 4 | All | Factor × 1: 3F800000h Factor × 2 (values have to be scaled by factor two): 40000000h | 3F 80 00 00 40 00 00 00 00 |
|  | Scale factor offset | Sets starting point of measurement | Real as float accordin g to IEEE754 | 4 | All | 0h | 00 00 00 00 00 |
|  | Start angle | Output format: 1/10000° | Int_32 | 4 | LMS1xx | -450000d... + 2250000d | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... 1850000d | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  |  |  |  | LMS1000 | -480000d... + 2280000d (FFF8AD00h... 22 CA40h) | FF F8 AD 00&hellip; 00 22 CA 40 |
|  |  |  |  |  | MRS1000 | -475000d... + 2275000d (FFF8C088h... 22 B6B8h) | FF F8 C0 88&hellip; 00 22 B6 B8 |
|  |  |  |  |  | LMS4000 | +550000d&hellip; +1250000d (86470h &hellip; 1312D0h) | 00 08 64 70... 00 13 12 D0 |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | Size of single angular step | Output format: 1/10000° | Uint_16 | 2 | LMS1xx | + 1000d... + 10000d | 03E8... 27 10 |
|  |  |  |  |  | LMS5xx | + 1667d... + 10000d | 06 83... 27 10 |
|  |  |  |  |  | LMS 1000 | + 7500d (1 D4Ch) | 1D 4C |
|  |  |  |  |  | MRS 1000 | + 2500d (9 C4h) | 09C4 |
|  |  |  |  |  | LMS 4000 | 0.0833: +833 (341h) | 0.0833: 03 41 |
|  | Amount of data | Amount | Uint_16 | 2 | All | 0000h... FFFFh | 00 00... FF FF |
|  | Data_1 Data_n | Data stream starting Data_1 to Data_n | Uint_8 | 1 | All | 00h... FFh | 00... FF |
|  |  |  |  |  | MRS 1000 LMS 1000 | DIST & RSSI: 0h&hellip; FFh AINF (Intensity): 0: 00 1: 04 2: 08 3: 0C 4: 10 5: 14 6: 18 7: 1C | 00... FF 00 04 08 0C 10 14 18 1C |
|  |  |  |  |  | LMS 4000 | 8-Bit Quality Channel Bit 0 (below lower signal threshold): e.g. 1h (0000 0001d) Bit 1 (above upper signal threshold): e.g. 2h (0000 0010d) Bit 2 (below lower distance threshold): e.g. 4h (0000 0100d) Bit 3 (above upper distance threshold): e.g. 8h (0000 1000d) Bit 4 (normal hit): e.g. 10h (0001 0000d) Bit 4+5 (edge hit liklehood): e.g. likely 20h (0010 0000d) or very likely (30h (0011 0000d) Bit 6 (possible outlier): e.g. 40h (0100 0000d) Bit 7 (partial gloss): e.g. 80h (1000 0000d) | 00... FF E. g. 01 E. g. 02 E. g. 04 E. g. 08 E. g. 10 e.g. 20 or 30 E. g. 40 E. g. 80 |
| Position |  | Output of position data | Enum_1 6 | 2 | All | No position data: 0 | No position data: 00 00 |
| Name |  | Device name | Uint_16 | 2 | All | No name: 0 Name: 1 | No name: 00 00 Name: 00 01 |
| Name information | Length | Length of name | Uint_16 | 2 | All | 0h... Fh | 00... 0F |
|  | Name | Device name in characters | String | 16 | All | 20h... 7Ah | 20... 7A |
| Comment |  | Comment | Uint_16 | 2 | All | No comment: 0 Comment: 1 | No comment: 00 00 Comment: 00 01 |

| Telegram structure: sRA LMDscandata/sSN LMDscandata |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Comment information | Length | Length of comment | Uint_8 | 1 | All | 0h... Fh | 00... 0F |
|  | Com- ment | Transmits a comment in characters | String | 16 | All | 20h... 7Ah | 20... 7A |
| Time |  | Transmits a time stamp | Uint_16 | 2 | All | No time: 0 | No time: 00 00 |
|  |  |  |  |  | LMS1xx, LMS5xx, NAV310, LD-OEM 15xx LD-LRS 36xx LMS 4000 LRS4000 | Time: 1 | Time: 00 01 |
| Time info | Year | 1970 to 2037 | Uint_16 | 2 | All | 7 B2h... 07F5h | 07B2... 07 F5 |
|  | Month | 1 to 12 | Uint_8 | 1 | All | 0h... Ch | 00... 0C |
|  | Day | Day of month 1 to 31 | Uint_8 | 1 | All | 0h... 1Fh | 00... 1F |
|  | Hour | 0 to 23 | Uint_8 | 1 | All | 0h... 17h | 00... 17 |
|  | Minute | 0 to 59 | Uint_8 | 1 | All | 0h... 3Bh | 00... 3B |
|  | Second | 0 to 59 | Uint_8 | 1 | All | 0h... 3Bh | 00... 3B |
|  | Micro- second | 0 to 999999 | Uint_32 | 4 | All | 0h... F423Fh | 00 00 00 00 00... 00 0F 42 3F |
| Event info |  | Display event info | Uint_16 | 2 | All | No info: 0 Transmit info: 1 | No info: 00 00 Transmit info: 00 01 |
| Event Information | Type | Fast digital input | String | 4 | All | FDIN | FDIN |
|  | Encoder position | Position of encod- er when event happened | Uint_32 | 4 | All | 00000000h&hellip; FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Time of event | Time (µs) of encoder when event happened | Uint_32 | 4 | All | 00000000h&hellip; FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Angle of event | Angle of encoder when event happened | Int_32 | 4 | All | 0... 3600000 | 00 00 00 00 00... 00 36 EE 80 |

Table 158: Telegram structure: Datastream of sRA LMDscandata/sSN LMDscandata

NOTE

- The grey written parts are not given out by the sensor.

- The event information is not available with the LMS1xx and with the LMS5xx only with firmware V1.20 or higher.

- The order of events within the data structure is "newest” first.

LMDscandata - reserved values

Valid distance measurement values are values starting from 16d upwards; everything below has the following meaning:

| DIST | RSSI | Description |
| --- | --- | --- |
| 0d | 0h | Invalid measurement value; caused by very low remission (extremely dark object), object distance not within measurement range (too close or too far away) or selected filter settings at device |
| 1d | FFFFh (16Bit output) FFh (8Bit output) | Invalid measurement value, device was dazzled or blinded, e.g. by measuring into the sun |
| 2d | 0h | Implausible measurement values (not used for LMS5xx) |
| 3d | 0h | Value was set to invalid by a filter (Echo Filter, Particle Filter) or GRE (ground reference evaluation, MRS1000) (not used for LMS5xx, TiM) |
| 4d-15d | 0h | Reserved, currently not in use |
| &ge; 16d | > 0h | Valid measurement values |

max. measurement value LMS4000: Dez: 3 150mm --> Hex: 7B0C (scale factor 0.1) max. measurement value TiM5xx: Dez: 10 000mm --> Hex: 2710

max. measurement value TiM57x: Dez: 25 000mm --> Hex: 61A8 max. measurement value LMS1xx: Dez: 20 000mm --> Hex: 4E20 max. measurement value LMS15x: Dez: 50 000mm --> Hex: C350 max. measurement value LMS5xx: Dez: 65 000mm --> Hex: FDE8

max. measurement value LMS5xx: Dez: 80 000mm --> Hex: 9C40 with scale factor 2 --> 13880

Higher measurement values will be given out with a zero, that means no measurement value detected.

Calculation and amount of data for LMS5xx

Example how to calculate the amount of data for a measurement telegram.

Sizes of values and telegram parts:

  - one measurement value: 5 byte (4 byte value itself, 1 byte blank after the value)

  - one RSSI value: 3 byte (2 byte value itself, 1 byte blank after the value)

  - telegram header: 81 byte

  - telegram end: 12 byte

Calculation of number of Measurement values depends always on the resolution: 0.5° = 2 measurements per degree

0.25° = 4 measurements per degree

Always one additional measurement for the last measurement

Number of measurement values =

Number of degrees × measurements per degree +1

Example for measurement of 56° in 0.5° resolution (without RSSI data): 56 × 2+1=113 Measurement values

Amount of Data for this measurement values: 113 × 5 Byte = 565 Byte

Calculation of amount of data per telegram:

Data of one Telegram = Header + Measurements + end of telegram

81Byte +113 Measurements +12 Byte 81 Byte +(113 × 5Byte) +12 Byte =

658Byte per Telegram (= 5264 Bit (658 × 8 Bit))

Possible amout for delivery with special Speed:

Number of telegrams per second = Speed ÷ telegram size

Speed Example:

115200Bit/s = 11520 Byte/s = 11,52 Byte/s

11520 (Byte/s) ÷ 658 Byte = 17,5 Telegrams/s

Telegram size with0,25°resolution: Degrees: 270°

Resolution: 0.25°

Measurement Values = 270 × 4+1=1081

Data per Telegram =

81Byte +(1081 × 5 Byte) +12 Byte = 5498 Byte ( = 43984 Bit)

Telegram size with 0,5° resolution: Degrees: 270°

Resolution: 0.5°

Measurement Values = 270 × 2+1=541 Data per Telegram =

81Byte +(541 × 5 Byte) +12 Byte = 2798 Byte ( = 22384 Bit)

As a result in that configuration a 10 MBit connection will not be enough. With a 100 MBit Hub, 3-4 scanner can be used, with a 1 GBit Hub accordingly more.

Example of a telegram stream

Example: telegram LMS1xx, LMS5xx similar with corresponding values (10°-20° data range)

ASCII

<STX> sRA {SPC} LMDscandata {SPC} 1 {SPC} 1 {SPC} 89A27F {SPC} 0 {SPC} 0 {SPC} 343 {SPC} 347 {SPC} 2747 7BA9 {SPC} 2747813B {SPC} 0 {SPC} 0 {SPC} 7 {SPC} 0 {SPC} 0 {SPC} 1388 {SPC} 168 {SPC} 0 {SPC} 1 {SPC} DIST

1 {SPC} 3F800000 {SPC} 00000000 {SPC} 186A0 {SPC} 1388 {SPC} 15 {SPC} 8A1 {SPC} 8A5 {SPC} 8AB {SPC}

8AC {SPC} 8A6 {SPC} 8AC {SPC} 8B6 {SPC} 8C8 {SPC} 8C2 {SPC} 8C9 {SPC} 8CB {SPC} 8C4 {SPC} 8E4 {SPC} 8E1 { SPC} 8EB {SPC} 8E0 {SPC} 8F5 {SPC} 908 {SPC} 8FC {SPC} 907 {SPC} 906 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0

{SPC} 0 <ETX>

BINARY

02 02 02 00 00 00 83 73 52 41 20 4C 4D 44 73 63 61 6E 64 61 74 61 00 01 00 01 00 89

A2 7F 00 00 03 43 03 47 27 47 7B A9 27 47 81 3B 00 00 07 00 00 00 00 00 00 00 00 00 13 88 00 01 68

00 00 00 01 44 49 53 54 31 3F 80 00 00 00 00 00 00 00 00 01 86 A0 13 88 00 15 08 93 08 95 08

AF 08 B3 08 B0 08 A4 08 B0 08 BF 08 B9 08 BA 08 D0 08 D3 08 CF 08 DE 08 EB 08 E3 08 FE 08

EC 09 03 08 FD 08 FD 00 00 00 00 00 00 00 00 00 00 00 00 2B

| Telegram structure: sRA LMDscandata (Example) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Telegram part |  | Variable | Length | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Frame/header |  |  |  | 02 <STX> | 02 02 02 02 |
| Length |  |  |  |  | 00 00 00 83 |
| Command type |  | String | 3 | sRA{SPC} | 73 52 41 20 |
| Command |  | String | 11 | LMDscandata{SPC} | 4C 4D 44 73 63 61 6E 64 61 74 61 20 |
| Version number |  | Uint_16 | 2 | 1{SPC} | 00 01 |
| vi c e | Device number | Uint_16 | 2 | 1{SPC} | 00 01 |

| Telegram structure: sRA LMDscandata (Example) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Telegram part |  | Variable | Length | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | Serial number | Uint_32 | 4 | 89A27F{SPC} Dec: 9020031 | 00 89 A2 7F |
|  | Device status | Uint_8 | 2 × 1 | 0{SPC}0{SPC} | 00 00 |
| Status information | Telegram counter | Uint_16 | 2 | 343{SPC} Dec: 835 | 03 43 |
|  | Scan counter | Uint_16 | 2 | 347{SPC} Dec: 839 | 03 47 |
|  | Time since start up [µs] | Uint_32 | 4 | 27477BA9{SPC} Dec: 658996137 | 27 47 7B A9 |
|  | Time of transmission [µs] | Uint_32 | 4 | 2747813B{SPC} Dec: 568997563 | 27 47 81 3B |
|  | Status of digital inputs | Uint_8 | 2 × 1 | 0{SPC}0{SPC} | 00 00 |
|  | Status of digital outputs | Uint_8 | 2 × 1 | 7{SPC}0{SPC} Corresponds to status 0111 | 07 00 |
|  | Reserved | Uint_16 | 2 | 0{SPC} | 00 00 |
| Frequen- cies | Scan frequency | Uint_32 | 4 | 1388{SPC} Dec: 5000/100=50Hz | 00 00 13 88 |
|  | Measurement frequency | Uint_32 | 4 | 168{SPC} | 00 00 01 68 |
| Amount of encoder |  | Enum_16 | 2 | 0{SPC} No encoder data | 00 00 |
| Position information | Encoder position | Uint_16 | 2 | Not generated, not existing because amount is 0 | Not generated, not existing be- cause amount is 0 |
|  | Encoder speed | Uint_16 | 2 | Not generated, not existing because amount is 0 | Not generated, not existing because amount is 0 |
| Amount of 16 bit channels |  | Enum_16 | 2 | 1{SPC} | 00 01 |
| Output channel (16 bit) | Content | String | 5 | DIST1{SPC} | 44 49 53 54 31 |
|  | Scale factor accord- ing to IEEE754 | Real | 4 | 3F800000{SPC} Floating Point: Value = 1 | 3F 80 00 00 |
|  | Scale factor offset acco. to IEEE754 | Real | 4 | 0{SPC} Floating Point: Value = 0 | 00 00 00 00 00 |
|  | Start angle | Int_32 | 4 | 186A0{SPC} Dec: 100000/10000 = 10° | 00 01 86 A0 |
|  | Size of single angular step | Uint_16 | 2 | 1388{SPC} Dec: 5000/10000 = 0.5° | 13 88 |
|  | Amount of data | Uint_16 | 2 | 15{SPC} Dec: 21 measurement points | 00 15 |
|  | Data_1... Data_21 | Uint_16 | 2 | 8A1{SPC}8A5{SPC}8AB{SPC}8 AC{SPC}8A6{SPC}8AC{SPC}8B 6{SPC}8C8{SPC}8C2{SPC}8C9{ SPC}8CB{SPC}8C4{SPC}8E4{S PC}8E1{SPC}8EB{SPC}8E0{SP C}8F5{SPC}908{SPC}8FC{SPC} 907{SPC}906{SPC} Measurement data Min. 22 mm: 16h Max. 20000 mm: 4E20h | 08A1 08 A5 08 AB 08 AC 08 A6 08 AC 08 B6 08 C8 08 C2 08C9 08 CB 08 C4 08 E4 08 E1 08 EB 08 E0 08 F5 09 08 08FC 09 07 09 06 |

| Telegram structure: sRA LMDscandata (Example) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Telegram part |  | Variable | Length | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Amount of 8 bit channels |  | Enum_16 | 2 | 0{SPC} No 8 bit data | 00 00 No 8 bit data |
| Output channel (8 bit) | Content | String | 5 | - | - |
|  | Scale factor | Real | 4 | - | - |
|  | Scale factor offset | Real | 4 | - | - |
|  | Start angle | Int_32 | 4 | - | - |
|  | Size of single angular step | Uint_16 | 2 | - | - |
|  | Amount of data | Uint_16 | 2 | - | - |
|  | Data_1 Data_n | Uint_8 | 1 | - | - |
| Position |  | Enum_16 | 2 | 0{SPC} No position data | 00 00 No position data |
| Position information | X position | Real | 4 | - | - |
|  | Y position | Real | 4 | - | - |
|  | Z position | Real | 4 | - | - |
|  | X rotation | Real | 4 | - | - |
|  | Y rotation | Real | 4 | - | - |
|  | Z rotation | Real | 4 | - | - |
|  | Rotations type | Enum_8 | 1 | - | - |
|  | Transmits the name of device | Uint_8 | 1 | - | - |
| Name |  | Enum_16 | 2 | No device name: 0 | 00 00 |
| Comment |  | Enum_16 | 2 | No comment: 0 | 00 00 |

| Telegram structure: sRA LMDscandata (Example) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Telegram part |  | Variable | Length | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Time |  | Enum_16 | 2 | 0{SPC} No time transmitted | 00 00 No time transmitted |
| Time info | Year | Uint_16 | 2 | - | - |
|  | Month | Uint_8 | 1 | - | - |
|  | Day | Uint_8 | 1 | - | - |
|  | Hour | Uint_8 | 1 | - | - |
|  | Minute | Uint_8 | 1 | - | - |
|  | Second | Uint_8 | 1 | - | - |
|  | Microsecond | Uint_32 | 4 | - | - |
| Event info |  | Enum_16 | 2 | 0{SPC} No event info available | 00 00 No event info available |
| Event information | Type | String | 4 | - | - |
|  | Encoder position | Uint_32 | 4 | - | - |
|  | Time of event | Uint_32 | 4 | - | - |
|  | Angle of event | Int_32 | 4 | - | - |
| Frame |  |  |  | 03 <ETX> | 2B Checksum |

Table 159: Example of one telegram stream

  - Time stamp

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.4.1 | Set time stamp | x | x |  |  |  | x | x | x |  |  |  | x | x |
| 4.4.2 | Read time stamp and status of the measurement function | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.4.3 | Read device time |  |  |  |  |  | x | x | x |  |  | x | x | x |
| 4.4.4 | Set NTP (Network Time Protocol) parameters | x | x |  |  |  |  |  |  |  |  |  | x | x |

    - Set time stamp

The data format in the telegram is:

+2009{SPC}+7{SPC}+22{SPC}+12{SPC}+0{SPC}+0{SPC}+0.

The numbers represent year, month, day, hour, minute, second, microsecond).

If plus is used up-front the data it is interpreted as an integer decimal number, without the plus it's the scanner reads the data as hex format.

The answer is always in ASCII format.

Attention: There is no real time clock inside the device. When the scanner is switched off and after a reboot, the time has to be set again.

However, it is possible to analyze the Off-time in order to evade this issue.

| Telegram structure: sMN LSPsetdatetime (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set time stamp | String | 14 | All | LSPsetdatetime | 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 |
| Year |  | Uint_16 | 2 | All | 1970d... + 2099d (07 B2h &hellip; 0833h) | 07 b2... 08 33 |
| Month |  | Uint_8 | 1 | All | 01d... + 12d (01h &hellip; 0Ch) | 01... 0C |
| Day |  | Uint_8 | 1 | All | 01d... + 31d (01h... 1Fh) | 00... 1F |
| Hour |  | Uint_8 | 1 | All | 00d... + 23d | 00... 17 |

| Telegram structure: sMN LSPsetdatetime (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | (00h... 17h) |  |
| Minute |  | Uint_8 | 1 | All | 00d... + 59d (00h &hellip; 3Bh) | 00... 3B |
| Second |  | Uint_8 | 1 | All | 00d... + 59d (00h &hellip; 3Bh) | 00... 3B |
| Micro- second |  | Uint_32 | 4 | All | 00000000d... +00999999d (00000000h &hellip; 000F423Fh) | 00 00 00 00 00... 00 0F 42 3F |

Table 160: Telegram structure: sMN LSPsetdatetime

Example 1: sMN LSPsetdatetime

| CoLa A | ASCII | <STX> sMN {SPC} LSPsetdatetime {SPC} 7D9 {SPC} 2 {SPC} 11 {SPC} 10 {SPC} 22 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 20 37 44 39 20 32 20 31 31 20 31 30 20 32 32 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1E 73 4D 4E 20 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 20 07 D9 02 11 10 22 00 00 00 00 00 A3 |

Table 161: Example 1: sMN LSPsetdatetime

Example 2: sMN LSPsetdatetime

| CoLa A | ASCII | <STX> sMN{ SPC} LSPsetdatetime {SPC} +2010 {SPC} +01 {SPC} +26 {SPC} +10 {SPC} +35 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 20 2B 32 30 31 30 20 2B 30 20 2B 32 36 20 2B 31 30 20 2B 33 35 20 2B 30 20 2B 30 30 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1E 73 4D 4E 20 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 20 07 DA 01 1A 0A 23 00 00 00 00 00 A3 |

Table 162: Example 2: sMN LSPsetdatetime

| Telegram structure: sAN LSPsetdatetime |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Set time stamp | String | 14 | All | LSPsetdatetime | 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 |
| Status code | Code number | Enum_8 | 1 | All | Success: 1 | Success: 01 |

Table 163: Telegram structure: sAN LSPsetdatetime

Example 1, 2: sAN LSPsetdatetime

| CoLa A | ASCII | <STX> sAN {SPC} LSPsetdatetime {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 41 4E 20 4C 53 50 73 65 74 64 61 74 65 74 69 6D 65 20 01 51 |

Table 164: Example 1, 2: sAN LSPsetdatetime

Activate time stamp in the output string format or on SOPAS page "data processing ”.

    - Read time stamp and status of the measurement function

Command: sRN STlms

| Telegram structure: sRN STlms |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Status and time | String | 5 | All | STlms | 53 54 6C 6D 73 |

Table 165: Telegram structure: sRN STlms

Example: sRN STlms

| CoLa A | ASCII | <STX> sRN {SPC} STlms <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 53 54 6C 6D 73 03 |
| CoLa B | Binary | 02 02 02 00 00 09 73 52 4E 20 53 54 6C 6D 73 3A |

Table 166: Example: sRN STlms

Answer: sRA STlms

| Telegram structure: sRA STlms |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Status and time | String | 5 | All | STlms | 53 54 6C 6D 73 |
| Status code | Status of the measurement function. | Enum_ 16 | 2 | All | Boot cycle: 0&hellip; 5 Ready: 6 Measurement active: 7 FW Update: 8 Error: 10 | Boot cycle: 00 00&hellip; 00 05 Ready: 00 06 Measurement active: 00 07 FW Update: 00 08 Error: 00 0A |
| Temp. out of range | Device running in temp. range or not | Uint_8 | 1 | All | False (in range) = 0 True (out of range) = 1 | False (in range) = 00 True (out of range) = 01 |
| Length of time para- meter |  | Uint_16 | 2 | All | 0d... +65535d (00h &hellip; FFFFh) | 00 00... FF FF |
| Time | HH HH | Uint_16 | 2 | All | 0d... 99d | 00 00... 00 63 |
|  | : | Uint_8 | 1 | All | : | 3A |
|  | MM MM | Uint_16 | 2 | All | 0d... 99d | 00 00... 00 63 |
|  | : | Uint_8 | 1 | All | : | 3A |
|  | SS SS | Uint_16 | 2 | All | 0d... 99d | 00 00... 00 63 |
| Length of date para- meter |  | Uint_16 | 2 | All | 0d... +65535d (00h &hellip; FFFFh) | 00 00... FF FF |
| Date | DD DD | Uint_16 | 2 | All | 0d... 99d | 00 00... 00 63 |
|  | . | Uint_8 | 1 | All | . | 2E |
|  | MM MM | Uint_16 | 2 | All | 0d... 99d | 00 00... 00 63 |
|  | . | Uint_8 | 1 | All | . | 2E |
|  | YY YY YY YY | Uint_32 | 4 | All | 0d... 9999d | 00 00 00 00 00... 00 00 27 0F |
| LED1 |  | Uint_16 | 2 | All | Inactive: 0 Active: 1 | Inactive: 00 00 Active: 00 01 |
| LED2 |  | Uint_16 | 2 | All | Inactive: 0 Active: 1 | Inactive: 00 00 Active: 00 01 |
| LED3 |  | Uint_16 | 2 | All | Inactive: 0 Active: 1 | Inactive: 00 00 Active: 00 01 |
| Reserved |  | Uint_16 | 3 × 2 | All | 0 0 0 0 | 00 00 00 00 00 00 00 |

Table 167: Telegram structure: sRA STlms

Example: sRA STlms

| CoLa A | ASCII | <STX> sRA {SPC} STlms {SPC} 7 {SPC} 0 {SPC} 8 {SPC} 16:36:54 {SPC} 10 {SPC} 17.03.2030 {SPC} 0 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | Not available |
| CoLa B | Binary | 02 02 02 02 00 00 2F 73 52 41 20 53 54 6C 6D 73 20 00 07 00 00 08 00 10 3A 00 24 3A 00 36 00 0A 00 11 2E 00 03 2E 00 00 07 EE 00 00 00 00 00 00 00 00 00 00 00 00 17 |

Table 168: Example: sRA STlms

    - Read device time

Command to read the actual time of the internal clock (ms). The timer is 32 counter with a resolution of 1 ms.

| Telegram structure: sRN DeviceTime |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 10 | All | DeviceTime | 44 65 76 69 63 65 54 69 6D 65 |

Table 169: Telegram structure: sRN DeviceTime

Example: sRN DeviceTime

| CoLa A | ASCII | <STX> sRN {SPC} DeviceTime <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 44 65 76 69 63 65 54 69 6D 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 52 4E 20 44 65 76 69 63 65 54 69 6D 65 42 |

Table 170: Example: sRN DeviceTime

| Telegram structure: sRA DeviceTime |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 10 | All | DeviceTime | 44 65 76 69 63 65 54 69 6D 65 |
| Device time | Time | Uint_32 | 4 | All | 0d... + 9999d (0h &hellip; 270Fh) | 00 00 00 00 00... 00 00 27 0F |

Table 171: Telegram structure: sRA DeviceTime

Example: sRA DeviceTime 0

| CoLa A | ASCII | <STX> sRA {SPC} DeviceTime {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 44 65 76 69 63 65 54 69 6D 65 00 03 |
| CoLa B | Binary | 02 02 02 00 00 10 73 52 41 20 44 65 76 69 63 65 54 69 6D 65 00 00 00 00 6D |

Table 172: Example: sRA DeviceTime 0

    - Set NTP (Network Time Protocol) parameters

Set time synchronization

| Telegram structure: sWN TSCRole (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set NTP role | String | 7 | All | TSCRole | 54 53 43 52 6F 6C 65 |
| Variable data | NTP role | Uint_8 | 1 | All | None: 0 Client: 1 Server: 2 | None: 00 Client: 01 Server: 02 |

Table 173: Telegram structure: sWN TSCRole

Example: sWN TSCRole

| CoLa A | ASCII | <STX> sWN {SPC} TSCRole {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 54 53 43 52 6F 6C 65 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 57 4E 20 54 53 43 52 6F 6C 65 20 01 1B |

Table 174: Example: sWN TSCRole

| Telegram structure: sWA TSCRole |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set NTP role | String | 7 | All | TSCRole | 54 53 43 52 6F 6C 65 |

Table 175: Telegram structure: sWA TSCRole

Example: sWA TSCRole

| CoLa A | ASCII | <STX> sWA {SPC} TSCRole <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 54 53 43 52 6F 6C 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 41 20 54 53 43 52 6F 6C 65 20 15 |

Table 176: Example: sWA TSCRole

Set time synchronization interface

| Telegram structure: sWN TSCTCInterface (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set time synchroni- zation interface | String | 14 | All | TSCTCInterface | 54 53 43 54 43 49 6E 74 65 72 66 61 63 65 |
| Variable data | Time synchroniza- tion interface data | Uint_8 | 1 | All | Ethernet: 0 CAN: 1 | Ethernet: 00 CAN: 01 |

Table 177: Telegram structure: sWN TSCTCInterface

Example: sWN TSCTCInterface

| CoLa A | ASCII | <STX> sWN {SPC} TSCTCInterface {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 54 53 43 54 43 49 6E 74 65 72 66 61 63 65 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 57 4E 20 54 53 43 54 43 49 6E 74 65 72 66 61 63 65 20 00 7C |

Table 178: Example: sWN TSCTCInterface

| Telegram structure: sWA TSCTCInterface |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set time synchroni- zation | String | 14 | All | TSCTCInterface | 54 53 43 54 43 49 6E 74 65 72 66 61 63 65 |

Table 179: Telegram structure: sWA TSCTCInterface

Example: sWA TSCTCInterface

| CoLa A | ASCII | <STX> sWA {SPC} TSCTCInterface <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 54 53 43 54 43 49 6E 74 65 72 66 61 63 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 41 20 54 53 43 54 43 49 6E 74 65 72 66 61 63 65 20 73 |

Table 180: Example: sWA TSCTCInterface

Set time server IP address

| Telegram structure: sWN TSCTCSrvAddr (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set time server IP address | String | 12 | All | TSCTCSrvAddr | 54 53 43 54 43 53 72 76 41 64 64 72 |
| IP address data | Set values in hex | Uint_32 | 4 | All | 00 00 00 00h... FF FF FF FFh (decimal values unwieldy) | 00 00 00 00 00... FF FF FF FF |

Table 181: Telegram structure: sWN TSCTCSrvAddr

Example: sWN TSCTCSrvAddr 192.168.0.11

| CoLa A | ASCII | <STX> sWN {SPC} TSCTCSrvAddr {SPC} C0 {SPC} A8 {SPC} 00 {SPC} 0B <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 54 53 43 54 43 53 72 76 41 64 64 72 20 C0 A8 00 0B 03 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 57 4E 20 54 53 43 54 43 53 72 76 41 64 64 72 20 C0 A8 00 0B 3E |

Table 182: Example: sWN TSCTCSrvAddr 192.168.0.11

| Telegram structure: sWA TSCTCSrvAddr |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set time server IP address | String | 12 | All | TSCTCSrvAddr | 54 53 43 54 43 53 72 76 41 64 64 72 |

Table 183: Telegram structure: sWA TSCTCSrvAddr

Example: sWA TSCTCSrvAddr

| CoLa A | ASCII | <STX> sWA {SPC} TSCTCSrvAddr <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 54 53 43 54 43 53 72 76 41 64 64 72 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 57 41 20 54 53 43 54 43 53 72 76 41 64 64 72 20 52 |

Table 184: Example: sWA TSCTCSrvAddr

Set time zone

| Telegram structure: sWN TSCTCtimezone (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set time zone | String | 13 | All | TSCTCtimezone | 54 53 43 54 43 74 69 6D 65 7A 6F 6E 65 |
| Time zone data | Set values in number of hours relative to GMT, hex specially coded | Int_8 | 1 | All | [GMT + &hellip;] -12d... + 12d (00h... 18h) | [GMT + &hellip;] 00... 18 |

Table 185: Telegram structure: sWN TSCTCtimezone

Example: sWN TSCTCtimezone GMT +1 hour

| CoLa A | ASCII | <STX> sWN {SPC} TSCTCtimezone {SPC} +1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 54 53 43 54 43 74 69 6D 65 7A 6F 6E 65 0D 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 4E 20 54 53 43 54 43 74 69 6D 65 7A 6F 6E 65 0D 3F |

Table 186: Example: sWN TSCTCtimezone GMT +1 hour

| Telegram structure: sWA TSCTCtimezone |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set time zone | String | 13 | All | TSCTCtimezone | 54 53 43 54 43 74 69 6D 65 7A 6F 6E 65 |

Table 187: Telegram structure: sWA TSCTCtimezone

Example: sWA TSCTCtimezone

| CoLa A | ASCII | <STX> sWA {SPC} TSCTCtimezone <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 54 53 43 54 43 74 69 6D 65 7A 6F 6E 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 57 41 20 54 53 43 54 43 74 69 6D 65 7A 6F 6E 65 20 3D |

Table 188: Example: sWA TSCTCtimezone

Set update time

| Telegram structure: sWN TSCTCupdatetime (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set update time of synchronization | String | 15 | All | TSCTCupdatetime | 54 53 43 54 43 75 70 64 61 74 65 74 69 6D 65 |
| Update time of synchroni- zation | Set values in seconds | Uint_32 | 4 | All | + 1d... + 3600d (01h... 0e10h) | 00 00 00 00 00... 00 00 0E 10 |

Table 189: Telegram structure: sWN TSCTCupdatetime

Example: sWN TSCTCupdatetime 600 s

| CoLa A | ASCII | <STX> sWN {SPC} TSCTCupdatetime {SPC} +600 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 54 53 43 54 43 75 70 64 61 74 65 74 69 6D 65 20 02 58 03 |
| CoLa B | Binary | 02 02 02 00 00 00 18 73 57 4E 20 54 53 43 54 43 75 70 64 61 74 65 74 69 6D 65 20 00 00 02 58 67 |

Table 190: Example: sWN TSCTCupdatetime 600 s

| Telegram structure: sWA TSCTCupdatetime |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set update time of synchronization | String | 15 | All | TSCTCupdatetime | 54 53 43 54 43 75 70 64 61 74 65 74 69 6D 65 |

Table 191: Telegram structure: sWA TSCTCupdatetime

Example: sWA TSCTCupdatetime

| CoLa A | ASCII | <STX> sWA {SPC} TSCTCupdatetime <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 54 53 43 54 43 75 70 64 61 74 65 74 69 6D 65 |
| CoLa B | Binary | 02 02 02 00 00 14 73 57 41 20 54 53 43 54 43 75 70 64 61 74 65 74 69 6D 65 20 32 |

Table 192: Example: sWA TSCTCupdatetime

Read for maximum offset time

| Telegram structure: sRN TSCTCmaxoffset (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read maximum offset time | String | 14 | All | TSCTCmaxoffset | 54 53 43 54 43 6D 61 78 6F 66 66 73 65 74 |

Table 193: Telegram structure: sRN TSCTCmaxoffset

Example: sRN TSCTCmaxoffset

| CoLa A | ASCII | <STX> sRN {SPC} TSCTCmaxoffset <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 54 53 43 54 43 6D 61 78 6F 66 66 73 65 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 4E 20 54 53 43 54 43 6D 61 78 6F 66 66 73 65 74 65 |

Table 194: Example: sRN TSCTCmaxoffset

| Telegram structure: sRA TSCTCmaxoffset |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read maximum offset time | String | 14 | All | TSCTCmaxoffset | 54 53 43 54 43 6D 61 78 6F 66 66 73 65 74 |
| Max. offset time | [Seconds as float according to IEEE754] | Real | 4 | All | 0h... FFFFFFFFh Min Value ~ -3.403*10^38 s Max Value ~+3.403*10^38 s | 00 00 00 00 00... FF FF FF FF |

Table 195: Telegram structure: sRA TSCTCmaxoffset

Example: sRA TSCTCmaxoffset (18000 s)

| CoLa A | ASCII | <STX> sRA {SPC} TSCTCmaxoffset {SPC} 468CA000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 54 53 43 54 43 6D 61 78 6F 66 66 73 65 74 34 36 38 43 41 30 30 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 17 73 52 41 20 54 53 43 54 43 6D 61 78 6F 66 66 73 65 74 20 46 8C A0 00 20 |

Table 196: Example: sRA TSCTCmaxoffset 18000 s

Read for delay time

| Telegram structure: sRN TSCTCdelay (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read delay time | String | 10 | All | TSCTCdelay | 54 53 43 54 43 64 65 6C 61 79 |

Table 197: Telegram structure: sRN TSCTCdelay

Example: sRN TSCTCdelay

| CoLa A | ASCII | <STX> sRN {SPC} TSCTCdelay <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 54 53 43 54 43 64 65 6C 61 79 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 52 4E 20 54 53 43 54 43 64 65 6C 61 79 69 |

Table 198: Example: sRN TSCTCdelay

| Telegram structure: sRA TSCTCdelay |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read for delay time | String | 10 | All | TSCTCdelay | 54 53 43 54 43 64 65 6C 61 79 |
| Max. offset time | [Seconds as float according to IEEE754] | Real | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

Table 199: Telegram structure: sRA TSCTCdelay

Example: sRA TSCTCdelay (0.003 s)

| CoLa A | ASCII | <STX> sRA {SPC} TSCTCdelay {SPC} 3B435B02 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 54 53 43 54 43 64 65 6C 61 79 20 33 42 34 33 35 42 30 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 52 41 20 54 53 43 54 43 64 65 6C 61 79 20 3B 43 5B 02 67 |

Table 200: Example: sRA TSCTCdelay 0.003 s

Reset maximum offset time

This command resets the maximum offset time, I .e. sets it to zero (0).

| Telegram structure: sMN mResetMaxOff (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Reset maximum offset time | String | 12 | All | mResetMaxOff | 6D 52 65 73 65 74 4D 61 78 4F 66 66 66 |

Table 201: Telegram structure: sMN mResetMaxOff

Example: sMN mResetMaxOff

| CoLa A | ASCII | <STX> sMN {SPC} mResetMaxOff <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 52 65 73 65 74 4D 61 78 4F 66 66 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 10 73 4D 4E 20 6D 52 65 73 65 74 4D 61 78 4F 66 66 73 |

Table 202: Example: sMN mResetMaxOff

| Telegram structure: sAN mResetMaxOff |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Read maximum offset time | String | 12 | All | mResetMaxOff | 6D 52 65 73 65 74 4D 61 78 4F 66 66 66 |

Table 203: Telegram structure: sAN mResetMaxOff

Example: sAN mResetMaxOff

| CoLa A | ASCII | <STX> sAN {SPC} mResetMaxOff <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 52 65 73 65 74 4D 61 78 4F 66 66 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 41 4E 20 6D 52 65 73 65 74 4D 61 78 4F 66 66 20 5F |

Table 204: Example: sAN mResetMaxOff

  - Filter

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.5.1 | Set particle filter | x | x | x | x | x |  |  |  | x | x |  |  |  |
| 4.5.2 | Set mean filter | x | x |  |  |  |  |  |  | x | x |  | x | x |
| 4.5.3 | Set n-pulse to 1-pulse filter (Echo filter) | x |  |  |  |  |  |  |  |  |  |  |  |  |
| 4.5.4 | Set echo filter |  | x |  |  |  |  |  |  | x | x | x |  | x |
| 4.5.5 | Set and read fog filter | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.5.6 | Set sensitivity fog filter (LMS5xx, LMS/MRS1000) |  | x |  |  |  |  |  |  | x | x |  |  |  |
| 4.5.7 | Activate/deactivate "fog filter operating radius active" |  | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.5.8 | Execute the "Teach from 90° angle” button |  | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.5.9 | Setting of the value "Fog filter operating radius active up to” |  | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.5.10 | Enable/disable digital nearfield filter |  |  |  |  |  |  | x |  |  |  |  |  |  |
| 4.5.11 | Set digital nearfield filter sector selection |  |  |  |  |  |  | x | x |  |  |  |  |  |
| 4.5.12 | Activate Median filter |  |  |  |  |  |  |  |  | x | x |  | x | x |
| 4.5.13 | Activate Edge filter |  |  |  |  |  |  |  |  |  |  |  | x |  |
| 4.5.14 | Set Edge filter |  |  |  |  |  |  |  |  |  |  |  | x |  |
| 4.5.15 | Set Cubic Area filter |  |  |  |  |  |  |  |  |  |  |  | x |  |
| 4.5.16 | Activate Gloss compensation filter |  |  |  |  |  |  |  |  |  |  |  | x |  |
| 4.5.17 | Activate long range mode |  |  |  |  |  |  |  |  |  |  | x |  |  |
| 4.5.18 | Read status of long range mode |  |  |  |  |  |  |  |  |  |  | x |  |  |

    - Set particle filter

| Telegram structure: sWN LFPparticle (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set particle filter | String | 11 | All | LFPparticle | 4C 46 50 70 61 72 74 69 63 6C 65 |
| Status code | Code number | Bool_1 | 1 | All | Inactive: 0 Active: 1 | Inactive: 00 Active: 01 |
| Thresh- old 5) | Particle threshold in mm | Uint_16 | 2 | All | +500d (must be taken) (1F4h) | 01F4 (must be taken) |

Table 205: Telegram structure: sWN LFPparticle

Example: sWN LFPparticle

| CoLa A | ASCII | <STX> sWN {SPC} LFPparticle {SPC} 1 {SPC} +500 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 70 61 72 74 69 63 6C 65 20 31 20 2B 35 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 4E 20 4C 46 50 70 61 72 74 69 63 6C 65 01 01 F4 D0 |

Table 206: Example: sWN LFPparticle

| Telegram structure: sWA LFPparticle |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set particle filter | String | 11 | All | LFPparticle | 4C 46 50 70 61 72 74 69 63 6C 65 |

Table 207: Telegram structure: sWA LFPparticle

5) 	Never change the threshold here, it is taken by the device to handle the particles.

Example: sWA LFPparticle

| CoLa A | ASCII | <STX> sWA {SPC} LFPparticle <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 70 61 72 74 69 63 6C 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 57 41 20 4C 46 50 70 61 72 74 69 63 6C 65 20 2B |

Table 208: Example: sWA LFPparticle

    - Set mean filter

| Telegram structure: sWN LFPmeanfilter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set mean filter | String | 13 | All | LFPmeanfilter | 4C 46 50 6D 65 61 6E 66 69 6C 74 65 72 |
| Status code | Code number | Bool_1 | 1 | All | Inactive: 0 Active: 1 | Inactive: 00 Active: 01 |
| Number of scans | Number | Uint_16 | 2 | All | + 2d... + 100d (00 02h... 00 64h) | 00 02... 00 64 |
| Final part | Reserved | Enum_8 | 1 | All | 0 | 00 |

Table 209: Telegram structure: sWN LFPmeanfilter

Example: sWN LFPmeanfilter

| CoLa A | ASCII | <STX> sWN {SPC} LFPmeanfilter {SPC} 1 {SPC} +10 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 6D 65 61 6E 66 69 6C 74 65 72 20 31 20 2B 31 30 20 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 16 73 57 4E 20 4C 46 50 6D 65 61 6E 66 69 6C 74 65 72 20 01 00 64 00 52 |

Table 210: Example: sWN LFPmeanfilter

| Telegram structure: sWA LFPmeanfilter |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set mean filter | String | 13 | All | LFPmeanfilter | 4C 46 50 6D 65 61 6E 66 69 6C 74 65 72 |

Table 211: Telegram structure: sWA LFPmeanfilter

Example: sWA LFPmeanfilter

| CoLa A | ASCII | <STX> sWA {SPC} LFPmeanfilter <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 6D 65 61 6E 66 69 6C 74 65 72 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 57 41 20 4C 46 50 6D 65 61 6E 66 69 6C 74 65 72 |

Table 212: Example: sWA LFPmeanfilter

    - Set n-pulse to 1-pulse filter (Echo filter)

Only LMS1xx, for LMS5xx take the echo filter.

| Telegram structure: sWN LFPnto1filter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | LMS1xx | sWN | 73 57 4E |
| Command | Set n-to-1 filter | String | 13 | LMS1xx | LFPnto1filter | 4C 46 50 6E 74 6F 31 66 69 6C 74 65 72 |
| Status code | Code number | Bool_1 | 1 | LMS1xx | Inactive: 0 Active: 1 | Inactive: 00 Active: 01 |

Table 213: Telegram structure: sWN LFPnto1filter

Example: sWN LFPnto1filter

| CoLa A | ASCII | <STX> sWN {SPC} LFPnto1filter {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 6E 74 6F 31 66 69 6C 74 65 72 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 4E 20 4C 46 50 6E 74 6F 31 66 69 6C 74 65 72 20 01 75 |

Table 214: Example: sWN LFPnto1filter

| Telegram structure: sWA LFPnto1filter |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | LMS1xx | sWA | 73 57 41 |
| Command | Set n-to-1 filter | String | 13 | LMS1xx | LFPnto1filter | 4C 46 50 6E 74 6F 31 66 69 6C 74 65 72 |

Table 215: Telegram structure: sWA LFPnto1filter

Example: sWA LFPnto1filter

| CoLa A | ASCII | <STX> sWA {SPC} LFPnto1filter <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 6E 74 6F 31 66 69 6C 74 65 72 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 57 41 20 4C 46 50 6E 74 6F 31 66 69 6C 74 65 72 7B |

Table 216: Example: sWA LFPnto1filter

    - Set echo filter

Only LMS5xx. For LMS1xx use the n-pulse to 1-pulse filter.

| Telegram structure: sWN FREchoFilter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set echo filter | String | 12 | All | FREchoFilter | 46 52 45 63 68 6F 46 69 |

| Telegram structure: sWN FREchoFilter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  |  | 6C 74 65 72 |
| Status code | Code number | Enum_8 | 1 | All | First echo: 0 All echos: 1 Last echo: 2 | First echo: 00 All echos: 01 Last echo: 02 |

Table 217: Telegram structure: sWN FREchoFilter

Example: sWN FREchoFilter

| CoLa A | ASCII | <STX> sWN {SPC} FREchoFilter {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 46 52 45 63 68 6F 46 69 6C 74 65 72 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 57 4E 20 46 52 45 63 68 6F 46 69 6C 74 65 72 20 01 7E Only available with firmware versions > V1.10. |

Table 218: Example: sWN FREchoFilter

| Telegram structure: sWA FREchoFilter |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set echo filter | String | 12 | All | FREchoFilter | 46 52 45 63 68 6F 46 69 6C 74 65 72 |

Table 219: Telegram structure: sWA FREchoFilter

Example: sWa FREchoFilter

| CoLa A | ASCII | <STX> sWA {SPC} FREchoFilter <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 46 52 45 63 68 6F 46 69 6C 74 65 72 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 57 41 20 46 52 45 63 68 6F 46 69 6C 74 65 72 20 70 Only available with firmware versions > V1.10 LMS5xx. |

Table 220: Example: sWa FREchoFilter

    - Set and read fog filter

Set fog filter (LMS1xx)

| Telegram structure: sWN MSsuppmode (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set fog filter | String | 10 | All | MSsuppmode | 4D 53 73 75 70 70 6D 6F 64 65 |
| Status code | Code number | Bool_1 | 1 | All | Glitch: 0 Fog: 1 | Glitch: 00 Fog: 01 |

Table 221: Telegram structure: sWN MSsuppmode

Example: sWN MSsuppmode

| CoLa A | ASCII | <STX> sWN {SPC} MSsuppmode {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4D 53 73 75 70 6D 6F 64 65 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 57 4E 20 4D 53 73 75 70 70 6D 6F 64 65 20 01 70 |

Table 222: Example: sWN MSsuppmode

| Telegram structure: sWA MSsuppmode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set fog filter | String | 10 | All | MSsuppmode | 4D 53 73 75 70 70 6D 6F 64 65 |

Table 223: Telegram structure: sWA MSsuppmode

Example: sWA MSsuppmode

| CoLa A | ASCII | <STX> sWA {SPC} MSsuppmode <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4D 53 73 75 70 6D 6F 64 65 70 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 57 41 20 4D 53 73 75 70 70 6D 6F 64 65 7E |

Table 224: Example: sWA MSsuppmode

Set fog filter (LMS5xx)

| Telegram structure: sWN CLFogFilterEn (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Enable fog filter | String | 13 | All | CLFogFilterEn | 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E |
| Status code | Enable or disable fog filter | Bool_1 | 1 | All | Disable: 0 Enable: 1 | Disable: 00 Enable: 01 |

Table 225: Telegram structure: sWN CLFogFilterEn

Example: sWN CLFogFilterEn

| CoLa A | ASCII | <STX> sWN {SPC} CLFogFilterEn {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 4E 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 20 01 21 |

Table 226: Example: sWN CLFogFilterEn

| Telegram structure: sWA CLFogFilterEn |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Enable fog filter | String | 13 | All | CLFogFilterEn | 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E |

Table 227: Telegram structure: sWA CLFogFilterEn

Example: sWA CLFogFilterEn

| CoLa A | ASCII | <STX> sWA {SPC} CLFogFilterEn <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 12 73 57 41 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 20 2F |

Table 228: Example: sWA CLFogFilterEn

Read for enabled fog filter (LMS5xx)

| Telegram structure: sRN CLFogFilterEn |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Enabled fog filter | String | 13 | All | CLFogFilterEn | 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E |

Table 229: Telegram structure: sRN CLFogFilterEn

Example: sRN CLFogFilterEn

| CoLa A | ASCII | <STX> sRN {SPC} CLFogFilterEn <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 52 4E 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 05 |

Table 230: Example: sRN CLFogFilterEn

| Telegram structure: sRA CLFogFilterEn |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Enabled fog filter | String | 13 | All | CLFogFilterEn | 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E |
| Status code | Fog filter enabled or disabled | Bool_1 | 1 | All | Disabled: 0 Enabled: 1 | Disabled: 00 Enabled: 01 |

Table 231: Telegram structure: sRA CLFogFilterEn

Example: sRA CLFogFilterEn

| CoLa A | ASCII | <STX> sRA {SPC} CLFogFilterEn {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 20 01 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 13 73 52 41 20 43 4C 46 6F 67 46 69 6C 74 65 72 45 6E 20 01 2B |

Table 232: Example: sRA CLFogFilterEn

    - Set sensitivity fog filter (LMS5xx, LMS/MRS1000)

| Telegram structure: sWN MCSenseLevel (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Sense level | String | 12 | All | MCSenseLevel | 4D 43 53 65 6E 73 65 4C 65 76 65 6C |
| Sensitivity level | Enable or disable fog filter and Sense Level | Uint_8 | 1 | LMS5xx | 1... 6 | 01... 06 |
|  |  |  |  | LMS/ | 0 (Fog Filter off) | 00 |
|  |  |  |  | MRS | 1 (Fog Filter on) | 01 |
|  |  |  |  | 1000 |  |  |

Table 233: Telegram structure: sWN MCSenseLevel

Example: sWN MCSenseLevel

| CoLa A | ASCII | <STX> sWN {SPC} MCSenseLevel {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4D 43 53 65 6E 73 65 4C 65 76 65 6C 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 57 4E 20 4D 43 53 65 6E 73 65 4C 65 76 65 6C 20 01 70 |

Table 234: Example: sWN MCSenseLevel

| Telegram structure: sWA MCSenseLevel |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sense level | String | 12 | All | MCSenseLevel | 4D 43 53 65 6E 73 65 4C 65 76 65 6C |

Table 235: Telegram structure: sWA MCSenseLevel

Example: sWA MCSenseLevel

| CoLa A | ASCII | <STX> sWA {SPC} MCSenseLevel <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4D 43 53 65 6E 73 65 4C 65 76 65 6C 20 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 57 41 20 4D 43 53 65 6E 73 65 4C 65 76 65 6C 20 7E |

Table 236: Example: sWA MCSenseLevel

    - Activate/deactivate "fog filter operating radius active"

| Telegram structure: sWN FogFilterMaxRangeEnable (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Enable max. range restriction for fog filter | String | 23 | All | FogFilterMaxRangeEnable | 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 45 6E 61 62 6C 65 |

| Telegram structure: sWN FogFilterMaxRangeEnable (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Enable/Dis able |  | Bool_1 | 1 | All | Activate : 1 Deactivate: 0 | Activate : 01 Deactivate: 00 |

Table 237: Telegram structure: sWN FogFilterMaxRangeEnable

Example: sWN FogFilterMaxRangeEnable 1

| CoLa A | ASCII | <STX> sWN {SPC} FogFilterMaxRangeEnable {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 45 6E 61 62 6C 65 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1D 73 57 4E 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 45 6E 61 62 6C 65 20 01 2F |

Table 238: Example: sWN FogFilterMaxRangeEnable 1

| Telegram structure: sWA FogFilterMaxRangeEnable (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Enable max. range restriction for fog filter | String | 23 | All | FogFilterMaxRangeEnable | 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 45 6E 61 62 6C 65 |

Table 239: Telegram structure: sWA FogFilterMaxRangeEnable

Example: sWA FogFilterMaxRangeEnable

| CoLa A | ASCII | <STX> sWN SPC FogFilterMaxRangeEnable <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 45 6E 61 62 6C 65 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1C 73 57 41 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 45 6E 61 62 6C 65 20 21 |

Table 240: Example: sWA FogFilterMaxRangeEnable

    - Execute the "Teach from 90° angle” button

| Telegram structure: sMN TeachFogFilterMaxRange (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Teach max. range restriction for fog filter | String | 22 | All | TeachFogFilterMaxRange | 54 65 61 63 68 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 |

Table 241: Telegram structure: sMN TeachFogFilterMaxRange

Example: sMN TeachFogFilterMaxRange

| CoLa A | ASCII | <STX> sMN {SPC} TeachFogFilterMaxRange <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 54 65 61 63 68 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 03 |
| CoLa B | Binary | 02 02 02 00 00 1A 73 4D 4E 20 54 65 61 63 68 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 6E |

Table 242: Example: sMN TeachFogFilterMaxRange

| Telegram structure: sAN TeachFogFilterMaxRange (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Teach max. range restriction for fog filter | String | 22 | All | TeachFogFilterMaxRange | 54 65 61 63 68 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 |

Table 243: Telegram structure: sAN TeachFogFilterMaxRange

Example: sAN TeachFogFilterMaxRange

| CoLa A | ASCII | <STX> sAN {SPC} TeachFogFilterMaxRange <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 54 65 61 63 68 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 03 |

| CoLa B | Binary | 02 02 02 00 00 1B 73 41 4E 20 54 65 61 63 68 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 20 42 |
| --- | --- | --- |

Table 244: Example: sAN TeachFogFilterMaxRange

    - Setting of the value "Fog filter operating radius active up”

| Telegram structure: sWN FogFilterMaxRange (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Max. range restriction for fog filter | String | 17 | All | FogFilterMaxRange | 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 |
| Distance of the radius | [1/1000m] | Real as float accord- ing to IEEE754 | 4 | All | 0.5m : 43FA0000 up 16.0m: 467A0000 | 0.5m: 43 FA 00 00 up 16.0m: 46 7A 00 00 |

Table 245: Telegram structure: sWN FogFilterMaxRange

Example: sWN FogFilterMaxRange 43FA0000 ( 0,5m)

| CoLa A | ASCII | <STX> sWN {SPC} FogFilterMaxRange {SPC} 43FA0000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 20 43 FA 00 00 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1A 73 57 4E 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 20 43 FA 00 00B6 |

Table 246: Example: sWN FogFilterMaxRange 43FA0000

| Telegram structure: sWA FogFilterMaxRange (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Max. range restriction for fog filter | String | 17 | All | FogFilterMaxRange | 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 |

Table 247: Telegram structure: sWA FogFilterMaxRange

Example: sWA FogFilterMaxRange

| CoLa A | ASCII | <STX> sWA {SPC} FogFilterMaxRange <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 16 73 57 41 20 46 6F 67 46 69 6C 74 65 72 4D 61 78 52 61 6E 67 65 00 |

Table 248: Example: sWA FogFilterMaxRange

    - Enable/disable digital nearfield filter

Activates or deactivates the nearfield filter of the LD series. Do not change the setting on LD-LRS XXXX!

| Telegram structure: sWN CLNFDigFilterEn (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Digital nearfield filter | String | 15 | All | CLNFDigFilterEn | 43 4C 4E 46 44 69 67 46 69 6C 74 65 72 45 6E |
| Status code | Code number | Bool_1 | 1 | All | Inactive: 0 Active: 1 | Inactive: 00 Active: 01 |

Table 249: Telegram structure: sWN CLNFDigFilterEn

Example: sWN CLNFDigFilterEn

| CoLa A | ASCII | <STX> sWN {SPC} CLNFDigFilterEn {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 43 4C 4E 46 44 69 67 46 69 6C 74 65 72 45 6E 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 16 73 57 4E 20 43 4C 4E 46 44 69 67 46 69 6C 74 65 72 45 6E 20 01 51 |

Table 250: Example: sWN CLNFDigFilterEn

| Telegram structure: sWA CLNFDigFilterEn |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Digital nearfield filter | String | 15 | All | CLNFDigFilterEn | 43 4C 4E 46 44 69 67 46 69 6C 74 65 72 45 6E |

Table 251: Telegram structure: sWA CLNFDigFilterEn

Example: sWA CLNFDigFilterEn

| CoLa A | ASCII | <STX> sWA {SPC} CLNFDigFilterEn <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 43 4C 4E 46 44 69 67 46 69 6C 74 65 72 45 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 41 20 43 4C 4E 46 44 69 67 46 69 6C 74 65 72 45 6E |

Table 252: Example: sWA CLNFDigFilterEn

    - Set digital nearfield filter sector selection

Do not change the setting on LD-LRS XXXX!

| Telegram structure: sWN CLHWFilterSectEn (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Sector function | String | 16 | All | CLHWFilterSectEn | 43 4C 48 57 46 69 6C 74 65 72 53 65 63 74 45 6E |
| Status code | Active sector vector | Bool_1 | 4 × 1 | All | Active in none of the sectors: 0 0 0 0 | Active in none of the sectors: 00 00 00 00 |
|  |  |  |  |  | Active in all sectors: 1 1 1 1 | Active in all sectors: 01 01 01 01 |

Table 253: Telegram structure: sWN CLHWFilterSectEn

Example: sWN CLHWFilterSectEn

Enable Nearfield Suppression for sector 1, disable for sectors 2, 3 and 4.

| CoLa A | ASCII | <STX> sWN {SPC} CLHWFilterSectEn {SPC} 1 {SPC} 0 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 43 4C 48 57 46 69 6C 74 65 72 53 65 63 74 45 6E 20 31 20 30 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 19 73 57 4E 20 43 4C 48 57 46 69 6C 74 65 72 53 65 63 74 45 6E 20 31 30 30 30 51 |

Table 254: Example: sWN CLHWFilterSectEn 1 0 0 0

| Telegram structure: sWA CLHWFilterSectEn |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sector function | String | 16 | All | CLHWFilterSectEn | 43 4C 48 57 46 69 6C 74 65 72 53 65 63 74 45 6E |

Table 255: Telegram structure: sWA CLHWFilterSectEn

Example: sWA CLHWFilterSectEn

| CoLa A | ASCII | <STX> sWA {SPC} CLHWFilterSectEn <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 43 4C 48 57 46 69 6C 74 65 72 53 65 63 74 45 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 57 41 20 43 4C 48 57 46 69 6C 74 65 72 53 65 63 74 45 6E 20 5F |

Table 256: Example: sWA CLHWFilterSectEn

    - Activate Median filter

Activate a 3x1 Median filter (floating evaluation of 3 measurement points within one scan) for distance values

| Telegram structure: sWN LFPmedianfilter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Sector function | String | 15 | All | LFPmedianfilter | 4C 46 50 6D 65 64 69 61 6E 66 69 6C 74 65 72 |
| Status code | Enable or disable Median filter | Bool_1 | 1 | All | Inactive: 0 Active: 1 | 00 01 |
| Reserved | Always 3 | Uint_16 | 2 | All | 3 | 00 03 |

Table 257: Telegram structure: sWN LFPmedianfilter

Example: sWN LFPmedianfilter

Enable Median filter

| CoL a A | ASCII | <STX> sWN {SPC} LFPmedianfilter {SPC} 1 {SPC} 3 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 6D 65 64 69 61 6E 66 69 6C 74 65 72 20 31 20 33 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 57 4E 20 4C 46 50 6D 65 64 69 61 6E 66 69 6C 74 65 72 20 01 00 03 38 |

Table 258: Example: sWN LFPmedianfilter 1 3

| Telegram structure: sWA LFPmedianfilter |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sector function | String | 15 | All | LFPmedianfilter | 4C 46 50 6D 65 64 69 61 6E 66 69 6C 74 65 72 |

Table 259: Telegram structure: sWA LFPmedianfilter

Example: sWA LFPmedianfilter

| CoLa A | ASCII | <STX> sWA {SPC} LFPmedianfilter <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 6D 65 64 69 61 6E 66 69 6C 74 65 72 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 57 41 20 4C 46 50 6D 65 64 69 61 6E 66 69 6C 74 65 72 20 35 |

Table 260: Example: sWA LFPmedianfilter

    - Activate Edge filter

Activate the Edge filter to eliminate wrong measurement points at object edges.

| Telegram structure: sWN LFPedgefilter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Sector function | String | 13 | All | LFPedgefilter | 4C 46 50 65 64 67 65 66 69 6C 74 65 72 |
| Status code | Activate / deactivate edge filter | Bool_1 | 1 | All | Inactive: 0 Active: 1 | 00 01 |

Table 261: Telegram structure: sWN LFPedgefilter

Example: sWN LFPedgefilter

Enable Edge filter

| CoLa A | ASCII | <STX> sWN {SPC} LFPedgefilter {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 4E 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 20 01 32 |

Table 262: Example: sWN LFPedgefilter 1

| Telegram structure: sWA LFPedgefilter |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sector function | String | 13 | All | LFPedgefilter | 4C 46 50 65 64 67 65 66 69 6C 74 65 72 |

Table 263: Telegram structure: sWA LFPmedianfilter

Example: sWA LFPedgefilter

| CoLa A | ASCII | <STX> sWA {SPC} LFPedgefilter <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 57 41 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 20 3C |

Table 264: Example: sWA LFPedgefilter

    - Set Edge filter

Make the Edge filter work more or less agressively by reducing or increasing the maximum of the allowable distance difference between neighboring (valid) measurement values.

| Telegram structure: sWN LFPedgefilterMaxDist (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command | Write | String | 3 | All | sWN | 73 57 4E |

| Telegram structure: sWN LFPedgefilterMaxDist (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| type |  |  |  |  |  |  |
| Command | Sector function | String | 19 | All | LFPedgefilterMaxDist | 4C 46 50 65 64 67 65 66 69 6C 74 65 72 4D 61 78 44 69 73 74 |
| Distance value | Number | Uint_16 | 2 | All | + 1d... + 6000d (00 01h... 02 58h) | 00 01... 02 58 |

Table 265: Telegram structure: sWN LFPedgefilterMaxDist

Example: sWN LFPedgefilterMaxDist

Set Edge filter

| CoLa A | ASCII | <STX> sWN {SPC} LFPedgefilterMaxDist {SPC} 5 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 4D 61 78 44 69 73 74 20 35 03 |
| CoLa B | Binary | 02 02 02 02 00 00 1B 73 57 4E 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 4D 61 78 44 69 73 74 20 00 05 48 |

Table 266: Example: sWN LFPedgefilterMaxDist +5

| Telegram structure: sWA LFPedgefilterMaxDist |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sector function | String | 19 | All | LFPedgefilterMaxDist | 4C 46 50 65 64 67 65 66 69 6C 74 65 72 4D 61 78 44 69 73 74 |

Table 267: Telegram structure: sWA LFPedgefilterMaxDist

Example: sWA LFPedgefilterMaxDist

| CoLa A | ASCII | <STX> sWA {SPC} LFPedgefilterMaxDist <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 4D 61 78 44 69 73 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 19 73 57 41 20 4C 46 50 65 64 67 65 66 69 6C 74 65 72 4D 61 78 44 69 73 74 20 42 |

Table 268: Example: sWA LFPedgefilterMaxDist

    - Set Cubic Area filter

Define area in relative distance from the device. Only measurement points within this area are valid; all other measurement points are transmitted in the telegram as invald (=0).

| Telegram structure: sWN LFPcubicareafilter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Sector function | String | 18 | All | LFPcubicareafilter | 4C 46 50 63 75 62 69 63 61 72 65 61 66 69 6C 74 65 72 |
| Status code | Enable or disable Cubic Area filter | Bool_1 | 1 | All | Inactive: 0 Active: 1 | 00 01 |
| Distance value | Min. measurement distance in 1/10mm | Uint_32 | 4 | All | 0... 9.1776m: 0&hellip; 91776d (0 &hellip; 16680h) | 00 00 00 00 00... 00 01 66 80 |
| Distance value | Max. measurement distance in 1/10mm | Uint_32 | 4 | All | 0... 9.1776m: 0&hellip; 91776d (0 &hellip; 16680h) | 00 00 00 00 00... 00 01 66 80 |
| Distance value | Negative expansion in 1/10mm | Int_32 | 4 | All | 0... -4.5m: 0&hellip; -45000d Hex calculation: FFFFFFFF h {minus}[value h ]{plus}1 h | FF FF 50 38&hellip; FF FF FF FF |
| Distance value | Positive expansion in 1/10mm | Int_32 | 4 | All | 0... 4.5m: 0&hellip; 45000d (0 &hellip; AFC8h) | 00 00 00 00 00... 00 00 AF C8 |

Table 269: Telegram structure: sWN LFPcubicareafilter

Example: sWN LFPcubicareafilter

Activate Cubic Area filter with area from 1 m&hellip; 2 m and width from -1.5 m&hellip; +1.5 m

| CoLa A | ASCII | <STX> sWN {SPC} LFPcubicareafilter {SPC} 1 {SPC} +10000 {SPC} +20000 {SPC} -15000 {SPC} +15000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 63 75 62 69 63 61 72 65 66 69 6C 74 65 72 20 31 20 2B 31 30 30 30 30 20 2B 32 30 30 30 30 20 2D 31 35 30 30 20 2B 31 35 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 28 73 57 4E 20 4C 46 50 63 75 62 69 63 61 72 65 61 66 69 6C 74 65 72 01 00 00 27 10 00 00 4E 20 FF FF C5 68 00 00 3A 98 48 |

Table 270: Example: sWN LFPcubicareafilter 1 +10000 +20000 -15000 +15000

| Telegram structure: sWA LFPcubicareafilter |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sector function | String | 18 | All | LFPcubicareafilter | 4C 46 50 63 75 62 69 63 61 72 65 61 66 69 6C 74 65 72 |

Table 271: Telegram structure: sWA LFPcubicareafilter

Example: sWA LFPcubicareafilter

| CoLa A | ASCII | <STX> sWA {SPC} LFPcubicareafilter <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 63 75 62 69 63 61 72 65 61 66 69 6C 74 65 72 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 57 41 20 4C 46 50 63 75 62 69 63 61 72 65 61 66 69 6C 74 65 72 20 56 |

Table 272: Example: sWA LFPedgefilterMaxDist

    - Activate Gloss compensation filter

Compensate measurement points that would be invalid due to exceeingly high received signal (direct reflection) with an assumed distance value of the third last valid point.

| Telegram structure: sWN LFPglossfilter (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Sector function | String | 14 | All | LFPglossfilter | 4C 46 50 67 6C 6F 73 73 66 69 6C 74 65 72 |
| Status code | Activate / deactivate edge filter | Bool_1 | 1 | All | Inactive: 0 Active: 1 | 00 01 |

Table 273: Telegram structure: sWN LFPglossfilter

Example: sWN LFPglossfilter

Enable Gloss compensation

| CoLa A | ASCII | <STX> sWN {SPC} LFPglossfilter {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 46 50 67 6C 6F 73 73 66 69 6C 74 65 72 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 57 4E 20 4C 46 50 67 6C 6F 73 73 66 69 6C 74 65 72 20 01 55 |

Table 274: Example: sWN LFPglossfilter 1

| Telegram structure: sWA LFPglossfilter |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sector function | String | 14 | All | LFPglossfilter | 4C 46 50 67 6C 6F 73 73 66 69 6C 74 65 72 |

Table 275: Telegram structure: sWA LFPglossfilter

Example: sWA LFPglossfilter

| CoLa A | ASCII | <STX> sWA {SPC} LFPglossfilter <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 46 50 67 6C 6F 73 73 66 69 6C 74 65 72 |
| CoLa B | Binary | 02 02 02 02 00 00 00 13 73 57 41 20 4C 46 50 67 6C 6F 73 73 66 69 6C 74 65 72 20 5B |

Table 276: Example: sWA LFPglossfilter

    - Activate long range mode

Command extends the maximum scanning distance of the MRS6224 by 1.4 compared to standard range.

| Telegram structure: sWN EnableLongRangeMode (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Enable long range mode | String | 14 | All | EnableLongRangeMode | 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 |
| Status code | Activate / deactivate long range mode | Bool_1 | 1 | All | Inactive: 0 Active: 1 | 00 01 |

Table 277: Telegram structure: sWN EnableLongRangeMode

Example: sWN Enable LongRangeMode

Enable EnableLongRangeMode

| CoLa A | ASCII | <STX> sWN {SPC EnableLongRangeMode {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 19 73 57 4E 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 20 3 1 2C |

Table 278: Example: sWN EnableLongRangeMode 1

| Telegram structure: sWA EnableLongRangeMode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Sector function | String | 14 | All | EnableLongRangeMode | 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 |

Table 279: Telegram structure: sWA EnableLongeRangeModer

Example: sWA EnableLongRangeMode

| CoLa A | ASCII | <STX> sWA {SPC} EnableLongRangeMode <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 |
| CoLa B | Binary | 02 02 02 02 00 00 00 18 73 57 41 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 20 12 |

Table 280: Example: sWA EnableLongeRangeMode

    - Read status of long range mode

| Telegram structure: sRN EnableLongRangeMode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 57 4E |
| Command | Enable long range mode | String | 14 | All | EnableLongRangeMode | 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 |

Table 281: Telegram structure: sWN EnableLongRangeMode

Example: sRN EnableLongRangeMode

Request status Long range mode

| CoLa A | ASCII | <STX> sRN {SPC} EnableLongRangeMode <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 17 73 57 4E 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 38 |

Table 282: Example: sRN EnableLongRangeMode

| Telegram structure: sRA EnableLongRangeMode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Sector function | String | 14 | All | EnableLongRangeMode | 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 |
| Status code | Activate / deactivate long range mode | Bool_1 | 1 | All | Inactive: 0 Active: 1 | 00 01 |

Table 283: Telegram structure: sRA EnableLongRangeMode

Example: sRA EnableLongRangeMode 1

| CoLa A | ASCII | <STX> sRA {SPC} EnableLongRangeMode ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 13 73 52 41 20 45 6E 61 62 6C 65 4C 6F 6E 67 52 61 6E 67 65 4D 6F 64 65 01 16 |

Table 284: Example: sWA EnableLongrangeMode

  - Encoder

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.6.1 | Set increment source | x | x |  |  |  |  |  |  |  |  |  |  | x |
| 4.6.2 | Set encoder settings | x | x |  |  |  |  |  |  |  |  |  | x | x |
| 4.6.3 | Set encoder resolution | x | x |  |  |  |  |  |  |  |  |  |  | x |
| 4.6.4 | Set fixed speed | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.6.5 | Read speed threshold | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.6.6 | Read encoder speed | x | x |  |  |  |  |  |  |  |  |  |  | x |
| 4.6.7 | Reset encoder values |  |  |  |  |  |  |  |  |  |  |  | x |  |

    - Set increment source

| Telegram structure: sWN LICsrc (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set increment source | String | 6 | All | LICsrc | 4C 49 43 73 72 63 |
| Increment source |  | Enum_8 | 1 | All | Fixed speed: 0 Encoder: 1 | Fixed speed: 00 Encoder: 01 |

Table 285: Telegram structure: sWN LICsrc

Example: sWN LICsrc

| CoLa A | ASCII | <STX> sWN {SPC} LICsrc {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 49 43 73 72 63 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 4E 20 4C 49 43 73 72 63 20 01 4F |

Table 286: Example: sWN LICsrc

| Telegram structure: sWA LICsrc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set increment source | String | 6 | All | LICsrc | 4C 49 43 73 72 63 |

Table 287: Telegram structure: sWA LICsrc

Example: sWA LICsrc

| CoLa A | ASCII | <STX> sWA {SPC} LICsrc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 49 43 73 72 63 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0B 73 57 41 20 4C 49 43 73 72 63 41 |

Table 288: Example: sWA LICsrc

    - Set encoder settings

| Telegram structure: sWN LICencset (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Encoder settings | String | 9 | All | LICencset | 4C 49 43 65 6E 63 73 65 74 |
| Encoder setting |  | Enum_8 | 1 | All | Off: 0 Single increment/INC1: 1 | 00 01 |
|  |  |  |  |  | Direction recognition (phase): 2 | 02 |
|  |  |  |  |  | Direction recognition (level): 3 | 03 |
|  |  |  |  | LMS 4000 | (see above +) Fixed increment speed / ticks (1 kHz): 4 | 04 |

Table 289: Telegram structure: sWN LICencset

Example: sWN LICencset

| CoLa A | ASCII | <STX> sWN {SPC} LICencset {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 49 43 65 6E 63 73 65 74 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 57 4E 20 4C 49 43 65 6E 63 73 65 74 20 00 26 |

Table 290: Example: sWN LICencset

| Telegram structure: sWA LICencset |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Encoder settings | String | 9 | All | LICencset | 4C 49 43 65 6E 63 73 65 74 |

Table 291: Telegram structure: sWA LICencset

Example: sWA LICencset

| CoLa A | ASCII | <STX> sWA {SPC} LICencset <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 49 43 65 6E 63 73 65 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 41 20 4C 49 43 65 6E 63 73 65 74 20 29 |

Table 292: Example: sWA LICencset

    - Set encoder resolution

| Telegram structure: sWN LICencres (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set encoder resolution | String | 9 | All | LICencres | 4C 49 43 65 6E 63 72 65 73 |
| Encoder resolution | Resolution value in mm/Inc as float according to IEEE754 | Real | 4 | All | + 0.001d... + 2000d | 3A 83 12 6F... 44 FA 00 00 (see IEEE 754) |

Table 293: Telegram structure: sWN LICencres

Example: sWN LICencres

| CoLa A | ASCII | <STX> sWN {SPC} LICencres {SPC} +1000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 49 43 65 6E 63 72 65 73 20 2B 31 30 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 10 73 57 4E 20 4C 49 43 65 6E 63 72 65 73 20 44 7A 00 00 1E |

Table 294: Example: sWN LICencres

| Telegram structure: sWA LICencres |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set encoder resolution | String | 9 | All | LICencres | 4C 49 43 65 6E 63 72 65 73 |

Table 295: Telegram structure: sWA LICencres

Example: sWA LICencres

| CoLa A | ASCII | <STX> sWA {SPC} LICencres <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 49 43 65 6E 63 72 65 73 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0D 73 57 41 20 4C 49 43 65 6E 63 72 65 73 00 |

Table 296: Example: sWA LICencres

    - Set fixed speed

| Telegram structure: sWN LICFixVel (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set fixed speed | String | 9 | All | LICFixVel | 4C 49 43 46 69 78 56 65 6C |
| Fixed speed | Speed in m/s as float according to IEEE754 | Real | 4 | All | + 0.001d... +10.0d | 3A 83 12 6F... 41 20 00 00 |

Table 297: Telegram structure: sWN LICFixVel

Example: sWN LICFixVel

| CoLa A | ASCII | <STX> sWN {SPC} LICFixVel {SPC} +5 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 49 43 46 69 78 56 65 6C 20 2B 35 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0B 73 57 4E 20 4C 49 43 46 69 78 56 65 6C 20 40 A0 00 00 C4 |

Table 298: Example: sWN LICFixVel

| Telegram structure: sWA LICFixVel |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set fixed speed | String | 9 | All | LICFixVel | 4C 49 43 46 69 78 56 65 6C |

Table 299: Telegram structure: sWA LICFixVel

Example: sWA LICFixVel

| CoLa A | ASCII | <STX> sWA {SPC} LICFixVel <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 49 43 46 69 78 56 65 6C 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 57 41 20 4C 49 43 46 69 78 56 65 6C 0B |

Table 300: Example: sWA LICFixVel

    - Read speed threshold

| Telegram structure: sRN LICSpTh |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read speed threshold | String | 7 | All | LICSpTh | 4C 49 43 53 70 54 68 |

Table 301: Telegram structure: sRN LICSpTh

Example: sRN LICSpTh

| CoLa A | ASCII | <STX> sRN {SPC} LICSpTh <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 49 43 53 70 54 68 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 52 4E 20 4C 49 43 53 70 54 68 16 |

Table 302: Example: sRN LICSpTh

| Telegram structure: sRA LICSpTh |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Speed threshold | String | 7 | All | LICSpTh | 4C 49 43 53 70 54 68 |
| Speed threshold | Speed threshold in % | Uint_8 | 2 | All | + 1d... + 20d (01h &hellip; 14h) | 01... 14 |

Table 303: Telegram structure: sRA LICSpTh

Example: sRA LICSpTh

| CoLa A | ASCII | <STX> sRA {SPC} LICSpTh {SPC} 5 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 49 43 53 70 54 68 35 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 52 41 20 4C 49 43 53 70 54 68 05 3C |

Table 304: Example: sRA LICSpTh

    - Read encoder speed

| Telegram structure: sRN LICencsp |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read encoder speed | String | 8 | All | LICencsp | 4C 49 43 65 6E 63 73 70 |

Table 305: Telegram structure: sRN LICencsp

Example: sRN LICencsp

| CoLa A | ASCII | <STX> sRN {SPC} LICencsp <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 49 43 65 6C 63 73 70 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 52 4E 20 4C 49 43 65 6E 63 73 70 62 |

Table 306: Example: sRN LICencsp

| Telegram structure: sRA LICencsp |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read encoder speed | String | 8 | All | LICencsp | 4C 49 43 65 6E 63 73 70 |
| Encoder speed | [Speed in m/s as float according to IEEE754] | Real | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

Table 307: Telegram structure: sRA LICencsp

Example: sRA LICencsp (0 m/s)

| CoLa A | ASCII | <STX> sRA {SPC} LICencsp {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 49 43 65 6C 63 73 70 30 30 30 30 30 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 52 41 20 4C 49 43 65 6E 63 73 70 00 00 00 00 4D |

Table 308: Example: sRA LICencsp

    - Reset encoder values

| Telegram structure: sMN LIDrstencoderinc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Reset Encoder values | String | 16 | All | LIDrstencoderinc | 4C 49 44 72 73 74 65 6E 63 6F 64 65 72 69 6E 63 |

Table 309: Telegram structure: sMN LIDrstencoderinc

Example: sMN LIDrstencoderinc

| CoLa A | ASCII | <STX> sMN {SPC} LIDrstencoderinc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 4C 49 44 72 73 74 65 6E 63 6F 64 65 72 69 6E 63 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 11 73 4D 4E 20 4C 49 44 72 73 74 65 6E 63 6F 64 65 72 69 6E 63 62 |

Table 310: Example: sMN LIDrstencoderinc

| Telegram structure: sAN LIDrstencoderinc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Read encoder speed | String | 16 | All | LIDrstencoderinc | 4C 49 44 72 73 74 65 6E 63 6F 64 65 72 69 6E 63 |
| Status code | [Speed in m/s as float according to IEEE754] | Enum_8 | 1 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

Table 311: Telegram structure: sAN LIDrstencoderinc

Example: sAN LIDrstencoderinc

| CoLa A | ASCII | <STX> sAN {SPC} LIDrstencoderinc {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 4C 49 44 72 73 74 65 6E 63 6F 64 65 72 69 6E 63 20 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 16 73 41 4E 20 4C 49 44 72 73 74 65 6E 63 6F 64 65 72 69 6E 63 20 00 58 |

Table 312: Example: sAN LIDrstencoderinc

  - Inputs and Outputs

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.7.1 | Read state of the ports |  | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.7.2 | Read Port Configration of all I/OS (LRS4000 : Ports 3-8) |  |  |  |  |  |  |  |  | x | x |  | x | x |
| 4.7.3 | Read state of the inputs | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.7.4 | Read state of the outputs | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.7.5 | Receive outputstate by event | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.7.6 | Set output state | x | x | x | x | x | x | x | x |  |  |  | x | x |
| 4.7.7 | Change output 6/3 function |  | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.7.8 | Change output 1 function |  |  |  |  |  | x | x | x |  |  |  |  |  |
| 4.7.9 | Change output 1 logic state |  |  |  |  | x | x | x | x |  |  |  |  |  |
| 4.7.10 | Change output 2 function |  |  |  |  |  | x | x | x |  |  |  |  |  |
| 4.7.11 | Change output 2 logic state |  |  |  |  | x | x | x | x |  |  |  |  |  |
| 4.7.12 | Set synchronization mode | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.7.13 | Set synchronization phase |  | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.7.14 | Change input 4 function |  | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.7.15 | Set debouncing time for input x | x | x |  | x | x |  |  |  |  |  |  |  |  |
| 4.7.16 | Read status of external sync signal |  | x |  |  |  |  |  |  |  |  |  |  |  |

    - Read state of the ports

LIDportstate has to be available additionally or as successor of the LIDoutputstate telegram.

Valid for all sensors with Ethernet and ports (inputs / outputs).

| Telegram structure: sRN LIDportstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Ask for port configuration | String | 12 | All | LIDportstate | 4C 49 44 70 6F 72 74 73 74 61 74 65 |

Table 313: Telegram structure: sRN LIDportstate

Example: sRN LIDportstate

| CoLa A | ASCII | <STX> sRN {SPC} LIDportstate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 49 44 70 6F 72 74 73 74 61 74 65 |
| CoLa B | Binary | 02 02 02 02 00 00 00 10 73 52 4E 20 4C 49 44 70 6F 72 74 73 74 61 74 65 60 |

Table 314: Example: sRN LIDportstate

| Telegram structure: sRA/sSN LIDportstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA/sSN | 73 52 41/73 53 4E |
| Command | Port state | String | 12 | All | LIDportstate | 4c 49 44 70 6f 72 74 73 74 61 74 65 |
| Status code | Version number | Uint_16 | 2 | All | 0... FFFFh Current version: 1 | 00 01... FF FF |
|  | System counter (time in µs since power up max. 71min then starting from 0 again) | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
| ARRAY which defines the number of internal ports * | 0...n | UINT16 | 1 | All | Hex 00 00-FF FFF 00 = not available 01...n number of ports | 00 00-FF FF |
| State of the ports and count value in hex | Internal port state | Enum_8 | 1 |  | 00 = Output voltage low (Relays: open) 01 = Output voltage high (Relays: closed) |  |
|  |  |  |  |  | 02 = Tri-state |  |
|  |  |  |  |  | 03 = Input voltage high (level) |  |
|  |  |  |  |  | Input voltage from low to high (edge) |  |
|  |  |  |  |  | 04 = Input voltage low (level) |  |

| Telegram structure: sRA/sSN LIDportstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  | Input voltage high to low (edge) |  |
|  | Internal port counter | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | ... .. |  |  |  |  |  |
| ARRAY which defines the number of external or virtual p orts * | 0...n | UINT16 | 1 |  | Hex 00 00-FF FF 00 = not available 01...n number of ports | 00 00-FF FF |
| State of the ports and count value in hex | External port state | Enum_8 | 1 |  | 00 = Output voltage low (Relays: open) 01 = Output voltage high (Relays: closed) 02 = Tri-state 03 = Input voltage high (level) Input voltage from low to high (edge) 04 = Input voltage low (level) Input voltage high to low (edge) |  |
|  | External port counter | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | ... .. |  |  |  |  |  |
| Time | States code | Enum_16 | 1 |  | No time data: 00 00 Time data: 00 01 | No time data: 00 00 Time data: 00 01 |
| Time Block (sensor time from the last change of min. one of the outputs) | Year | Array | 2 |  | E. g. 1970 | E. g. 07 B2 |
|  | Month |  | 1 |  | 1... 12 | 01... 0C |
|  | Day |  | 1 |  | 1... 31 | 01... 1F |
|  | Hour |  | 1 |  | 0... 23 | 00... 17 |
|  | Minute |  | 1 |  | 0... 59 | 00... 3B |
|  | Second |  | 1 |  | 0... 59 | 00... 3B |
|  | Microsecond |  | 4 |  | 0... 999999 | 00 00 00 00 00... 00 0F 42 3F |

Table 315: Telegram structure: sRA/sSN LIDportstate

Inputs/outputs: If the device has separate inputs and outputs (instead of general purpose ports) the ARRAY shall start with inputs followed by the outputs.

Virtual ports are ports that can be used to expand the number of ports but are not physically available. They just show up in the corresponding ethernet telegrams (like LIDportstate).

Tri-State: Port is neither input nor output; the port is set inactive in SOPAS Example with 3 internal and 4 external ports:

Example: sRA LIDportstate

| CoLa A | ASCII | <STX> sRA {SPC} LIDportstate {SPC} 1 {SPC} 41F84EC5 {SPC} 3 {SPC} 1 {SPC} 20 {SPC} 1 {SPC} 20 {SPC} 3 {SP C} 20 {SPC} 4 {SPC} 1 {SPC} 20 {SPC} 1{ SPC} 20 {SPC} 0 {SPC} 20 {SPC} 03 {SPC} 20 {SPC} 1 {SPC} 7D9 {SPC} 2 {S PC} 12 {SPC} C {SPC} 29 {SPC} E {SPC} 975E0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 49 44 70 6F 72 74 73 74 61 74 65 31 20 34 31 46 38 34 45 43 35 20 33 20 31 20 32 30 20 31 20 32 30 20 33 20 32 30 20 34 20 31 20 32 30 20 31 20 32 30 20 30 20 32 30 20 30 33 20 32 30 20 31 20 37 44 39 20 32 20 31 32 20 43 20 32 39 20 45 20 39 37 35 45 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 5B 73 52 41 20 4C 49 44 70 6F 72 74 73 74 61 74 65 20 01 00 41 F8 4E C5 00 03 00 01 00 20 00 04 00 01 00 20 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 01 00 07 D9 00 02 00 12 00 0C 00 29 00 0E 00 09 75 E0 21 |

Table 316: Example: sRA LIDportstate

    - Read Port Configration of all I/OS (LRS4000 : Ports 3-8)

| Telegram structure: sRN PortConfiguration |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Ask for port configuration | String | 12 | All | PortConfiguration | 50 6F 72 74 43 6F 6E 66 69 67 75 72 61 74 69 6F 6E |

Table 317: Telegram structure: sRN PortConfiguration

Example: sRN PortConfiguration

| CoLa A | ASCII | <STX> sRN {SPC} PortConfiguration <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 50 6F 72 74 43 6F 6E 66 69 67 75 72 61 74 69 6F 6E 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 15 73 52 4E 20 50 6F 72 74 43 6F 6E 66 69 67 75 72 61 74 69 6F 6E 26 |

Table 318: Example: sRN PortConfiguration

| Telegram structure: sRA PortConfiguration |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type |  | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | Configuration of all I/OS | String | 12 | All | PortConfiguration | 50 6F 72 74 43 6F 6E 66 69 67 75 72 61 74 69 6F 6E |
| Start of loop, number of loops = amount of all current and future Inputs and Outputs of device family |  |  |  |  |  |  |  |
| Port Type |  | Input or Output | Enum_8 | 1 | All | Input: 0 Ouput: 1 | 00 01 |
| Port Name |  | Amount of characters of the following port name | Uint_16 | 2 | All | 0h... 20h | 00 00... 00 20 |
|  |  | Port name | String | 16 | All | [Port name] | [Port name] |
| Input Settings | Logic | Logic of the input | Bool_1 | 1 | All | Active High: 0 Active low: 1 | 00 01 |
|  | Debounc ing | Select debouncing time in ms | Uint_16 | 2 | All | 0h... 27 10 (max. 10,000ms) | 00 00... 27 10 |
|  | Sensitivit y | Status change at Edge or Level | Enum_8 | 1 | All | Edge: 0 Level: 1 | 00 01 |
|  | Reserved | Reserved value 1 | Uint_16 | 2 | All | 0h | 00 00 |
|  | Reserved | Reserved value 2 | Uint_16 | 2 | All | 0h | 00 00 |
| Output Settings | Logic | Logic of the output | Bool_1 | 1 | All | Active High: 0 Active low: 1 | 00 01 |
|  | Output Mode | PNP, NPN or Push- Pull | Enum_8 | 1 | All | PNP: 0 NPN: 1 Push-Pull: 2 | 00 01 02 |
|  | Restart type | Restart behavior of output after event: immediatly or after specific time | Enum_8 | 1 | All | Immediatly: 0 Time: 1 | 00 01 |
|  | Restart time | [Only with restart type = Time], time in ms | Uint_32 | 4 | All | 14h... 927C0 (20 ms &hellip; 600,000 ms) | 00 00 00 14... 00 09 27 C0 |
|  | Combina tion | Combining multiple Events and/or Inputs | Enum_8 | 1 | All | AND: 0 OR: 1 XOR: 2 | 00 01 02 |
|  |  |  |  | 1 | LMS 4000 | Always: 0 | 00 |
|  | Reserved | Reserved value 3 | Uint_16 | 2 | All | 0h | 00 00 |
|  | Reserved | Reserved value 4 | Uint_16 | 2 | All | 0h | 00 00 |
|  | Sources | Amount (n) | Uint_16 | 2 | All | 0h... FFFFh | 00 00... FF FF |

| Telegram structure: sRA PortConfiguration |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  | combined sources |  |  | LMS 4000 | 0... 1 | 00 00... 00 01 |
|  |  | Start of source loop, number of loops = amount of combined sources |  |  |  |  |  |
|  |  | Source name | String | 4 | All | [Source] DRDY = Device State SCxx = Sopas Command (xx = Number of output port) SROT = Indexsignal SCLK = SyncOutByClock Input = Port Name = INxx (INxx = Number of input port) | [Source] |
|  |  | Source Inverted or not | Bool_1 | 1 | All | Not inverted: 0 Inverted: 1 | 00 01 |
|  |  | Reserved value 5 | Uint_8 | 1 | All | 0h | 00 |
|  |  | Reserved value 6 | Uint_8 | 1 | All | 0h | 00 |
|  |  | Stop of source loop |  |  |  |  |  |
| Reserved |  | Reserved value 7 | Uint_16 | 2 | All | 0h | 00 00 |
| Reserved |  | Reserved value 8 | Uint_16 | 2 | All | 0h | 00 00 |
| Reserved |  | Reserved value 9 | Uint_16 | 2 | All | 0h | 00 00 |
| Reserved |  | Reserved value 10 | Uint_16 | 2 | All | 0h | 00 00 |
| Stop of loop |  |  |  |  |  |  |  |

    - Read state of the inputs

Use "sEN LIDinputstate 1” to receive a telegram each time an input signal (e.g. by trigger) changes. Compare with chapter 4.7.5 Receive outputstate by event.

| Telegram structure: sRN LIDinputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Output state | String | 14 | All | LIDinputstate | 4C 49 44 69 6E 70 75 74 73 74 61 74 65 |

Table 319: Telegram structure: sRN LIDinputstate

Example: sRN LIDinputstate

| CoLa A | ASCII | <STX> sRN {SPC} LIDinputstate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 49 44 69 6E 70 75 74 73 74 61 74 65 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 52 4E 20 4C 49 44 69 6E 70 75 74 73 74 61 74 65 0F |

Table 320: Example: sRN LIDinputstate

| Telegram structure: sRA LIDinputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Output state | String | 14 | All | LIDinputstate | 4C 49 44 69 6E 70 75 74 73 74 61 74 65 |
| Status | Version number | Uint_16 | 2 | All | 0... FFFFh | 00 00... FF FF |

| Telegram structure: sRA LIDinputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| code | System counter (time in µs since power up max. 71min then starting from 0 again) | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
| State of the inputs 1&hellip; n | Amount of inputs (n) depending of device family | Enum_8 | 1 | All | Not active: 0 Active: 1 Input not used: 2 | Not active: 00 Active: 01 Input not used: 02 |
| Time | States code | Uint_16 | 2 | All | No time data: 0 Time data: 1 | No time data: 00 00 Time data: 00 01 |
| Time Block | Year | Array | 2 | All | E. g. 1970 | E. g. 07 B2 |
|  | Month |  | 1 |  | 1... 12 | 01... 0C |
| (sensor- | Day |  | 1 |  | 1... 31 | 01... 1F |
| time from |  |  |  |  |  |  |
|  | Hour |  | 1 |  | 0... 23 | 00... 17 |
| the last |  |  |  |  |  |  |
| change |  |  |  |  |  |  |
|  | Minute |  | 1 |  | 0... 59 | 00... 3B |
| min. one |  |  |  |  |  |  |
|  | Second |  | 1 |  | 0... 59 | 00... 3B |
| the |  |  |  |  |  |  |
| outputs) |  |  |  |  |  |  |
|  | Microsecond |  | 4 |  | 0... 999999 | 00 00 00 00 00... 00 0F 42 3F |

Table 321: Telegram structure: sRA LIDinputstate

Example: sRA LIDinputstate In1 active, In2 active, In3 inactive, In4 not used, In5 not used, In6 inactive, In7 inactive, In8 not used, time: 2022-04-25 14:10 46.946sec

| CoLa A | ASCII | <STX> sRA {SPC} LIDinputstate {SPC} 1 {SPC} 98E71BD7 {SPC} 1 {SPC} 1 {SPC} 0 {SPC} 2 {SPC} 2 {SPC} 0 {SPC} 0 {SPC} 2 {S PC} 1 {SPC} 7E6 {SPC} 4 {SPC} 19 {SPC} E {SPC} A {SPC} 2E {SPC} E6F50 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 49 44 69 6E 70 75 74 73 74 61 74 65 31 20 39 38 45 37 31 42 44 37 20 31 20 31 20 30 20 32 20 32 20 30 20 30 20 32 20 31 20 37 45 36 20 34 20 31 39 20 45 20 41 20 32 45 20 45 36 46 35 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 2D 73 52 41 20 4C 49 44 69 6E 70 75 74 73 74 61 74 65 20 00 01 98 E7 1B D7 01 01 00 02 02 00 00 02 00 01 07 E6 04 19 0E 0A 2E 00 0E 6F 50 |

Table 322: Example: sRA LIDinputstate

    - Read state of the outputs

Generally status of all outputs, for LRS4000 Out 3-8

| Telegram structure: sRN LIDoutputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Output state | String | 14 | All | LIDoutputstate | 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 |

Table 323: Telegram structure: sRN LIDoutputstate

Example: sRN LIDoutputstate

| CoLa A | ASCII | <STX> sRN {SPC} LIDoutputstate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 4E 20 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 66 |

Table 324: Example: sRN LIDoutputstate

| Telegram structure: sRA LIDoutputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Find complete telegram structure of the answer in section 4.7.5 „ Receive outputstate by event "on page 175. |  |  |  |  |  |  |

Table 325: Telegram structure: sRA LIDoutputstate

    - Receive outputstate by event

Output telegram is sent every time an output state changes.

| Telegram structure: sEN LIDoutputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Event | String | 3 | All | sEN | 73 45 4E |
| Command | Output state | String | 14 | All | LIDoutputstate | 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 |
|  | Start/stop | Enum_8 | 1 | All | Start: 1 Stop: 0 | Start: 01 Stop: 00 |

Table 326: Telegram structure: sEN LIDoutputstate

Example: sEN LIDoutputstate

| CoLa A | ASCII | <STX> sEN {SPC} LIDoutputstate {SPC}1<ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 4E 20 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 45 4E 20 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 01 50 |

Table 327: Example: sEN LIDoutputstate

| Telegram structure: sRA/sSN LIDoutputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA/sSN | 73 52 41/73 53 4E |
| Command | Output state | String | 14 | All | LIDoutputstate | 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 |
| Status code | Version number | Uint_16 | 2 | All | 0... FFFFh | 00 00... FF FF |
|  | System counter (time in µs since power up max. 71min then starting from 0 again) | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
| State of the outputs and count value in hex | Out1 state | Enum_8 | 1 | All | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |

| Telegram structure: sRA/sSN LIDoutputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | Out1 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out2 state | Enum_8 | 1 | TiM24x TiM7xx LMS1xx LMS5xx LD-OEM15xx LD-LRS36xx MRS1000 LMS4000 MRS6000 | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |
|  | Out2 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out3 state | Enum_8 | 1 | TiM7xx LMS1xx LMS5xx LD-OEM15xx LD-LRS36xx MRS1000 LMS4000 MRS6000 LRS4000 | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |
|  | Out3 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out4 state | Enum_8 | 1 | TiM7xx LMS5xx LD-OEM15xx LD-LRS36xx MRS1000 LMS4000 MRS6000 LRS4000 | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |
|  | Out4 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out5 state | Enum_8 | 1 | LMS5xx LD-OEM15xx LD-LRS36xx MRS1000 LMS4000 MRS6000 LRS4000 | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |
|  | Out5 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out6 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |
|  | Out6 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out7 state | Enum_8 | 1 | MRS1000 LMS4000 MRS6000 LRS4000 | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |
|  | Out7 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out8 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 00 Active: 01 Output not used: 02 |
|  | Out8 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Out9 state | Enum_8 | 1 | LMS4000 | Always not used: 2 | Output not used: 02 |
|  | Out9 count | Uint_32 | 4 |  | 0 | 00 00 00 00 00 |
|  | Ext.Out1 state | Enum_8 | 1 | LMS1xx | Not active: 0 | Not active: 0 |

| Telegram structure: sRA/sSN LIDoutputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  | LMS5xx | Active: 1 Output not used: 2 | Active: 1 Output not used: 2 |
|  | Ext.Out1 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Ext.Out2 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 0 Active: 1 Output not used: 2 |
|  | Ext.Out2 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Ext.Out3 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 0 Active: 1 Output not used: 2 |
|  | Ext.Out3 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Ext.Out4 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 0 Active: 1 Output not used: 2 |
|  | Ext.Out4 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Ext.Out5 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 0 Active: 1 Output not used: 2 |
|  | Ext.Out5 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Ext.Out6 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 0 Active: 1 Output not used: 2 |
|  | Ext.Out6 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Ext.Out7 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 0 Active: 1 Output not used: 2 |
|  | Ext.Out7 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
|  | Ext.Out8 state | Enum_8 | 1 |  | Not active: 0 Active: 1 Output not used: 2 | Not active: 0 Active: 1 Output not used: 2 |
|  | Ext.Out8 count | Uint_32 | 4 |  | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
| Time | States code | Uint_16 | 2 | All | No time data: 0 Time data: 1 | No time data: 00 00 Time data: 00 01 |

| Telegram structure: sRA/sSN LIDoutputstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Time Block (sensor- time from the last change of min. one of the outputs) | Year | Array | 2 | All | E. g. 1970 | E. g. 07 B2 |
|  | Month |  | 1 |  | 1... 12 | 01... 0C |
|  | Day |  | 1 |  | 1... 31 | 01... 1F |
|  | Hour |  | 1 |  | 0... 23 | 00... 17 |
|  | Minute |  | 1 |  | 0... 59 | 00... 3B |
|  | Second |  | 1 |  | 0... 59 | 00... 3B |
|  | Microsecond |  | 4 |  | 0... 999999 | 00 00 00 00 00... 00 0F 42 3F |

Table 328: Telegram structure: sRA/sSN LIDoutputstate

Example: sRA LIDoutputstate Out1 active Count 5, Out2 active Count 5, Out3 active Count 5, all other Outputs not used, time: 2009-02-18 12:41 14.62sec

| CoLa A | ASCII | <STX> sRA {SPC} LIDoutputstate {SPC} 1 {SPC} 41F84EC5 {SPC} 1 {SPC} 5 {SPC} 1 {SPC} 5 {SPC} 1 {SPC} 5 {SPC} 2 {SPC} 0 {S PC} 2 {SPC} 0 {SPC} 2 {SPC} 0 {SPC} 2 {SPC} 0 {SPC} 2 {SPC} 0 {SPC} 2 {SPC} 0 {SPC} 2 {SPC} 0 {SPC} 2 {SPC} 0 {SPC} 1 {SPC} 7D9 {SPC} 2 {SPC} 12 {SPC} C {SPC} 29 {SPC} E {SPC} 975E0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 31 20 41 F8 4E C5 20 31 20 35 20 31 20 35 20 31 20 35 20 32 20 30 20 32 20 30 20 32 20 30 20 32 20 30 20 32 20 30 20 32 20 30 20 32 20 30 20 32 20 30 20 31 20 37 44 39 20 32 20 31 32 20 43 20 32 39 20 45 20 39 37 35 45 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 5D 73 52 41 20 4C 49 44 6F 75 74 70 75 74 73 74 61 74 65 20 00 01 41 F8 4E C5 |
|  |  | 01 00 00 00 05 01 00 00 05 01 00 00 05 02 00 00 00 00 02 00 00 00 02 00 00 00 |
|  |  | 00 00 00 02 00 00 00 00 02 00 00 00 00 02 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 01 07 D9 02 12 0C 29 0E |
|  |  | 00 09 75 E0 06 |

Table 329: Example: sRA LIDoutputstate

    - Set output state

NOTE

Output source needs to be set to "SOPAS command” and the port configured as Output (in case of I /O).

| Telegram structure: sMN mDOSetOutput |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set output state | String | 12 | All | mDOSetOutput | 6D 44 4F 53 65 74 4F 75 74 70 75 74 |
| Output number |  | Uint_8 | 1 | LMS1xx | 1... 3 | 01... 03 |
|  |  |  |  | LMS12x | 1... 2 | 01... 02 |
|  |  |  |  | LMS 4000 | 1... 4 | 01... 04 |
|  |  |  |  | LMS5xx | 1... 6 | 01... 06 |
|  |  |  |  | TiMxxx | 1 | 01 |
|  |  |  |  | TiM24x | 2 | 02 |
|  |  |  |  | LD-OEM | 1... 4 | 01... 04 |
|  |  |  |  | 1500 |  |  |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 3600 |  |  |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 3610 |  |  |
|  |  |  |  | LD-OEM | 01 | 01 |
|  |  |  |  | 1501 |  |  |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 3601 |  |  |
|  |  |  |  | LD-LRS |  |  |
|  |  |  |  | 3611 |  |  |
|  |  |  |  | LRS4000 | 03... 08 | 03&hellip;008 |
| Output state |  | Enum_8 | 1 | All | Inactive: 0 Active: 1 | Inactive: 00 Active: 01 |

Table 330: Telegram structure: sMN mDOSetOutput

Example: sMN mDOSetOutput

| CoLa A | ASCII | <STX> sMN {SPC} mDOSetOutput {SPC} 1 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 44 4F 53 65 74 4F 75 74 70 75 74 20 31 20 31 03 |

| CoLa B | Binary | 02 02 02 02 00 00 00 13 73 4D 4E 20 6D 44 4F 53 65 74 4F 75 74 70 75 74 20 01 01 69 |
| --- | --- | --- |

Table 331: Example: sMN mDOSetOutput

| Telegram structure: sAN mDOSetOutput |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Set output state | String | 12 | All | mDOSetOutput | 6D 44 4F 53 65 74 4F 75 74 70 75 74 |
| Status Code | Status code | Bool_1 | 1 | All | Error: 0 Success: 1 | Error: 00 Success: 01 |

Table 332: Telegram structure: sAN mDOSetOutput

Example: sAN mDOSetOutput

| CoLa A | ASCII | <STX> sAN {SPC} mDOSetOutput {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 44 4F 53 65 74 4F 75 74 70 75 74 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 41 4E 20 6D 44 4F 53 65 74 4F 75 74 70 75 74 20 01 67 |

Table 333: Example: sAN mDOSetOutput

    - Change output 6/3 function

| Telegram structure: sWN DO6Fnc/sWN DO3Fnc (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Output function | String | 6 | LMS5xx PRO | DO6Fnc | 44 4F 36 46 6E 63 |
|  |  |  |  | LMS5xx Lite | DO3Fnc | 44 4F 33 46 6E 63 |
| Output state |  | Enum_8 | 1 | All | No Function: 0 SOPAS command: 1 | Not available |
|  |  |  |  |  | Device Ready: 2 |  |
|  |  |  |  |  | Application: 3 |  |
|  |  |  |  |  | Applic./Device Ready: 4 |  |
|  |  |  |  |  | Dev.ready/Contamination: 5 |  |
|  |  |  |  |  | Contamination: 6 |  |
|  |  |  |  |  | Master Synchronisation: 7 |  |

Table 334: Telegram structure PRO: sWN DO6Fnc/Lite: sWN DO3Fnc

Example: sWN DO6Fnc  Set Out6 to Master Synchronisation

| CoLa A | ASCII | <STX> sWN {SPC} DO6Fnc {SPC} 7 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 4F 36 46 6E 63 20 37 03 |
| CoLa B | Binary | Unavailable with current firmware. |

Table 335: Example: sWN DO6Fnc  Out6 to master sync

| Telegram structure: sWA DO6Fnc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Output function | String | 6 | LMS5xx PRO | DO6Fnc | 44 4F 36 46 6E 63 |
|  |  |  |  | LMS5xx Lite | DO3Fnc | 44 4F 33 46 6E 63 |

Table 336: Telegram structure: PRO: sWN DO6Fnc/Lite: sWN DO3Fnc

Example: sWA DO6Fnc

| CoLa A | ASCII | <STX> sWA {SPC} DO6Fnc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 44 4F 36 46 6E 63 03 |
| CoLa B | Binary | Not available with firmware V1.10 |

Table 337: Example: sWA DO6Fnc

    - Change output 1 function

| Telegram structure: sWN DO1Fnc (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Output function | String | 6 | All | DO1Fnc | 44 4F 31 46 6E 63 |
| Output 1 function | Selected function | Enum_8 | 1 | All | No function: 0 Command: 1 | No function: 00 Command: 01 |
|  |  |  |  |  | Device ready: 2 | Device ready: 02 |
|  |  |  |  |  | Application dev. ready: 3 | Application dev. ready: 03 |
|  |  |  |  |  | Sync pulse: 4 | Sync pulse: 04 |
|  |  |  |  |  | Sync index: 5 | Sync index: 05 |

Table 338: Telegram structure: sWN DO1Fnc

Example: sWN DO1Fnc  Set Out1 to Device Ready

| CoLa A | ASCII | <STX> sWN {SPC} DO1Fnc {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 4F 31 46 6E 63 20 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 4E 20 44 4F 31 46 6E 63 20 02 19 |

Table 339: Example: sWN DO1Fnc  Out1 to device ready

| Telegram structure: sWA DO1Fnc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Output function | String | 6 | All | DO1Fnc | 44 4F 31 46 6E 63 |

Table 340: Telegram structure: sWA DO1Fnc

Example: sWA DO1Fnc

| CoLa A | ASCII | <STX> sWA {SPC} DO1Fnc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 44 4F 31 46 6E 63 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0A 73 57 41 20 44 4F 31 46 6E 63 34 |

Table 341: Example: sWA DO1Fnc

Functions:

No function: 0

Command: 1

Device ready (for field application): 2 Application dev. ready: 3

Sync pulse (10 ms puls when timer register is read "sRN STlms”): 4

Sync index: 5

The output signal depends on the scanner head position (high (+24 V): 0° &hellip; 179°/low (0 V): 180° &hellip; 360°).

    - Change output 1 logic state

| Telegram structure: sWN DO1Logic (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Output function | String | 8 | All | DO1Logic | 44 4F 31 4C 6F 67 69 63 |
| Output 1 logic state | State of the output | Enum_8 | 1 | All | Active_High: 0 Active_Low: 1 | Active_High: 00 Active_Low: 01 |

Table 342: Telegram structure: sWN DO1Logic

Example: sWN DO1Logic  Active_High

| CoLa A | ASCII | <STX> sWN {SPC} DO1Logic {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 4F 31 4C 6F 67 69 63 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 4E 20 44 4F 31 4C 6F 67 69 63 20 01 1F |

Table 343: Example: sWN DO1Logic  Active_Low

| Telegram structure: sWA DO1Logic |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Output logic | String | 8 | All | DO1Logic | 44 4F 31 4C 6F 67 69 63 |

Table 344: Telegram structure: sWA DO1Logic

Example: sWA DO1Logic

| CoLa A | ASCII | <STX> sWA {SPC} DO1Logic <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 44 4F 31 4C 6F 67 69 63 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 41 20 44 4F 31 4C 6F 67 69 63 31 |

Table 345: Example: sWA DO1Logic

    - Change output 2 function

| Telegram structure: sWN DO2Fnc (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Output function | String | 6 | All | DO2Fnc | 44 4F 32 46 6E 63 |
| Output 2 function | Code number | Enum_8 | 1 | All | No function: 0 Command: 1 | No function: 00 Command: 01 |
|  |  |  |  |  | Device ready: 2 | Device ready: 02 |
|  |  |  |  |  | Application dev. ready: 3 | Application dev. ready: 03 |

Table 346: Telegram structure: sWN DO2Fnc

Example: sWN DO2Fnc  Out2 to device ready

| CoLa A | ASCII | <STX> sWN {SPC} DO2Fnc {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 4F 32 46 6E 63 20 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 4E 20 44 4F 32 46 6E 63 20 02 1A |

Table 347: Example: sWN DO2Fnc  Out2 to device ready

| Telegram structure: sWA DO2Fnc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Output function | String | 6 | All | DO2Fnc | 44 4F 32 46 6E 63 |

Table 348: Telegram structure: sWA DO2Fnc

Example: sWA DO2Fnc

| CoLa A | ASCII | <STX> sWA {SPC} DO2Fnc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 44 4F 32 46 6E 63 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 57 41 20 44 4F 32 46 6E 63 37 |

Table 349: Example: sWA DO2Fnc

    - Change output 2 logic state

Identical for TiM7xx Outputs 3 and 4 with DO3Logic and DO4Logic.

| Telegram structure: sWN DO2Logic (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Output function | String | 8 | All | DO2Logic | 44 4F 32 4C 6F 67 69 63 |
| Output 2 logic state | State of the output | Enum_8 | 1 | All | Active_High: 0 Active_Low: 1 | Active_High: 00 Active_Low: 01 |

Table 350: Telegram structure: sWN DO2Logic

Example: sWN DO2Logic  Active_High

| CoLa A | ASCII | <STX> sWN {SPC} DO2Logic {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 4F 32 4C 6F 67 69 63 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 4E 20 44 4F 32 4C 6F 67 69 63 20 00 1C |

Table 351: Example: sWN DO2Logic  Active_High

| Telegram structure: sWA DO2Logic |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Output logic | String | 8 | All | DO2Logic | 44 4F 32 4C 6F 67 69 63 |

Table 352: Telegram structure: sWA DO2Logic

Example: sWA DO2Logic

| CoLa A | ASCII | <STX> sWA {SPC} DO2Logic <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 44 4F 32 4C 6F 67 69 63 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 41 20 44 4F 32 4C 6F 67 69 63 32 |

Table 353: Example: sWA DO2Logic

    - Set synchronization mode

| Telegram structure: sWN SYMode (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set sync mode | String | 6 | All | SYMode | 53 59 4D 6F 64 65 |
| Sync mode data | Synchronization mode data | Bool_1 | 1 | All | No sync = 0 Sync by wire = 1 Sync by CAN = 2 | Not possible |

Table 354: Telegram structure: sWN SYMode

Example: sWN SYMode

| CoLa A | ASCII | <STX> sWN {SPC} SYMode {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 53 59 4D 6F 64 65 20 31 03 |
| CoLa B | Binary | Not possible |

Table 355: Example: sWN SYMode

| Telegram structure: sWA SYMode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set sync mode | String | 6 | All | SYMode | 53 59 4D 6F 64 65 |

Table 356: Telegram structure: sWA SYMode

Example: sWA SYMode

| CoLa A | ASCII | <STX> sWA {SPC} SYMode <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 53 59 4D 6F 64 65 03 |
| CoLa B | Binary | Not possible |

Table 357: Example: sWA SYMode

    - Set synchronization phase

| Telegram structure: sWN SYPhase (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | Not possible |
| Command | Set sync phase | String | 7 | All | SYPhase | Not possible |
| Sync phase data | Synchronization phase data | Int_16 | 2 | All | -180d... + 180d (FF4Ch &hellip; 00 B4h) | Not possible |

Table 358: Telegram structure: sWN SYPhase

Example: sWN SYPhase +90

| CoLa A | ASCII | <STX> sWN {SPC} SYPhase {SPC} +90 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 53 59 50 68 61 73 65 20 2B 39 30 03 |
| CoLa B | Binary | Not possible |

Table 359: Example: sWN SYPhase +90

| Telegram structure: sWA SYPhase |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | Not possible |
| Command | Set sync phase | String | 7 | All | SYPhase | Not possible |

Table 360: Telegram structure: sWA SYPhase

Example: sWA SYPhase

| CoLa A | ASCII | <STX> sWA {SPC} SYPhase <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 53 59 50 68 61 73 65 03 |
| CoLa B | Binary | Not possible |

Table 361: Example: sWA SYPhase

    - Change input 4 function

| Telegram structure: sWN DO3And4Fnc (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Input function | String | 10 | All | DO3And4Fnc | 44 4F 33 41 6E 64 34 46 6E 63 |
| Input state | Code number | Enum_8 | 1 | All | No function: 0 Encoder: 1 Slave sync: 2 Digital input: 3 |  |

Table 362: Telegram structure: sWN DO3And4Fnc

Example: sWN In4  In3+4 to slave sync

| CoLa A | ASCII | <STX> sWN {SPC} DO3And4Fnc {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 4F 33 41 6E 64 34 46 6E 63 20 02 03 |
| CoLa B | Binary | Not available with firmware V1.10 |

Table 363: Example: sWN In4  In3+4 to slave sync

| Telegram structure: sWA DO3And4Fnc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Input function | String | 10 | All | DO3And4Fnc | 44 4F 33 41 6E 64 34 46 6E 63 |

Table 364: Telegram structure: sWA DO3And4Fnc

Example: sWA DO3And4Fnc

| CoLa A | ASCII | <STX> sWA {SPC} DO3And4Fnc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 44 4F 33 41 6E 64 34 46 6E 63 03 |
| CoLa B | Binary | Not available with firmware V1.10 |

Table 365: Example: sWA DO3And4Fnc

    - Set debouncing time for input x

The telegram applies for the inputs 1 to 4 (DIxDebTim, x = 1 &hellip; 4). The following tables show the data for input 3.

| Telegram structure: sWN DI3DebTim (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set debouncing time for input 3 | String | 9 | All | DI3DebTim | 44 49 33 44 65 62 54 69 6D |
| Debounc- ing time data | [ms] | Uint_16 | 2 | All | 0d... + 10000d (00h... 2710h) | 00 00... 27 10 |

Table 366: Telegram structure: sWN DI3DebTim

Example: sWN DI3DebTim

| CoLa A | ASCII | <STX> sWN {SPC} DI3DebTim {SPC} +10 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 49 33 44 65 62 54 69 6D 20 2B 31 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 57 4E 20 44 49 33 44 65 62 54 69 6D 20 00 0A 77 |

Table 367: Example: sWN DI3DebTim

| Telegram structure: sWA DI3DebTim |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set debouncing time for input 3 | String | 9 | All | DI3DebTim | 44 49 33 44 65 62 54 69 6D |

Table 368: Telegram structure: sWA DI3DebTim

Example: sWA DI3DebTim

| CoLa A | ASCII | <STX> sWA {SPC} DI3DebTim <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 44 49 33 44 65 62 54 69 6D 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 41 20 44 49 33 44 65 62 54 69 6D 20 48 |

Table 369: Example: sWA DI3DebTim

    - Read status of external sync signal

| Telegram structure: sRN SYextmon |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Status of external sync signal | String | 8 | All | SYextmon | 53 59 65 78 74 6D 6F 6E |

Table 370: Telegram structure: sRN SYextmon

Example: sRN SYextmon

| CoLa A | ASCII | <STX> sRN {SPC} SYextmon <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 53 59 65 78 74 6D 6F 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 52 4E 20 53 59 65 78 74 6D 6F 6E 40 |

Table 371: Example: sRN SYextmon

| Telegram structure: sRA SYextmon |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Status of external sync signal | String | 8 | All | SYextmon | 53 59 65 78 74 6D 6F 6E |
| Sync status data | Synchronization status data | Uint_8 | 1 | All | None: 1 Too slow: 2 | None: 01 Too slow: 02 |
|  |  |  |  |  | Good: 4 | Good: 04 |
|  |  |  |  |  | Too fast: 8 | Too fast: 08 |
| Signal frequency | [Frequency in Hz as float according to IEEE754] | Real | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

Table 372: Telegram structure: sRA SYextmon

Example: sRA SYextmon (49.9 Hz)

| CoLa A | ASCII | <STX> sRA {SPC} SYextmon {SPC} 4 {SPC} 4247BD87 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 53 59 65 78 74 6D 6F 6E 20 34 20 34 32 34 37 42 44 38 37 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 12 73 52 41 20 53 59 65 78 74 6D 6F 6E 20 04 42 47 BD 87 54 |

Table 373: Example: sRA SYextmon

  - Status

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.8.1 | Read contamination status of the LMS | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.8.2 | Read firmware version | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.8.3 | Read the device state | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.8.4 | Status commands for LD-XXX and NAV310 |  |  |  |  |  | x | x | x |  |  |  |  |  |
| 4.8.5 | Read device order number | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.8.6 | Read device type | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.8.7 | Read operating hours | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.8.8 | Read power on counter | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.8.9 | Read temperature | x | x |  | x | x |  |  |  |  |  |  | x | x |
| 4.8.10 | Set device name | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.8.11 | Read device name | x | x | x | x | x |  |  |  | x | x | x | x | x |
| 4.8.12 | Read angle compensation sine |  |  |  |  |  | x |  |  |  |  |  |  |  |
| 4.8.13 | Reset output counter | x | x | x | x | x |  |  |  |  |  | x | x | x |
| 4.8.14 | Read heating state | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.8.15 | FieldEvalLoggingList by Event |  |  |  |  |  |  |  |  | x | x |  |  |  |
| 4.8.16 | FieldEvalLoggingList on Request |  |  |  |  |  |  |  |  | x | x |  |  |  |

    - Read contamination status of the LMS

| Telegram structure: sRN LCMstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Status of LMS | String | 8 | All | LCMstate | 4C 43 4D 73 74 61 74 65 |

Table 374: Telegram structure: sRN LCMstate

Example: sRN LCMstate

| CoLa A | ASCII | <STX> sRN {SPC} LCMstate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 43 4D 73 74 61 74 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 52 4E 20 4C 43 4D 73 74 61 74 65 7A |

Table 375: Example: sRN LCMstate

| Telegram structure: sRA LCMstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Status of LMS | String | 8 | All | LCMstate | 4C 43 4D 73 74 61 74 65 |
| Status code |  | Enum_8 | 1 | All | No contamination: 0 Contamination warning: 1 | No contamination: 00 Contamination warning: 01 |
|  |  |  |  |  | Contamination error: 2 | Contamination error: 02 |
|  |  |  |  |  | Contamination measurement functionality defective: 3 | Contamination measurement functionality defective: 03 |

Table 376: Telegram structure: sRA LCMstate

Example for LMS100: sRA LCMstate

| CoLa A | ASCII | <STX> sRA {SPC} LCMstate {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 43 4D 73 74 61 74 65 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 52 41 20 4C 43 4D 73 74 61 74 65 00 55 |

Table 377: Example for LMS100: sRA LCMstate

    - Read firmware version

| Telegram structure: sRN DeviceIdent |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read firmware version | String | 11 | All | DeviceIdent | 44 65 76 69 63 65 49 64 65 6E 74 |

Table 378: Telegram structure: sRN DeviceIdent

Example: sRN DeviceIdent

| CoLa A | ASCII | <STX> sRN {SPC} DeviceIdent <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 44 65 76 69 63 65 49 64 65 6E 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 52 4E 20 44 65 76 69 63 65 49 64 65 6E 74 25 |

Table 379: Example: sRN DeviceIdent

| Telegram structure: sRA DeviceIdent |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 11 | All | DeviceIdent | 44 65 76 69 63 65 49 64 65 6E 74 |
| Value | Length of firmware designation | Enum_16 | 2 | All | 0... 22h | 0... 22h |
| Value | Firmware designation for device family | String |  | All | (See example) | (See example) |
| Value | Length of firmware version | Enum_16 | 2 | All | 0... 22h | 0... 22h |
| Value | Firmware version | String |  | All | (See example) | (See example) |

Table 380: Telegram structure: sRA DeviceIdent

Example: sRA DeviceIdent

| CoLa A | ASCII | <STX> sRA {SPC} DeviceIdent {SPC} 10 {SPC} LMS10x_FieldEval {SPC} 10 {SPC} V1.36-21.10.2010 <ETX> |
| --- | --- | --- |
|  | Hex | Always ASCII answer |
| CoLa B | Binary | 02 02 02 00 00 00 34 73 52 41 20 44 65 76 69 63 65 49 64 65 6E 74 20 00 10 4C 4D 53 31 30 78 5F 46 69 65 6C 64 45 76 61 6C 00 10 56 31 2E 33 36 2D 32 31 2E 31 30 2E 32 30 31 30 62 |

Table 381: Example: sRA DeviceIdent

    - Read the device state

This telegram reads the general device state.

Remark: The status of the measurement function of LMS1 and LMS5 can be read separately with the telegram STlms

| Telegram structure: sRN SCdevicestate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 13 | All | SCdevicestate | 53 43 64 65 76 69 63 65 73 74 61 74 65 |

Table 382: Telegram structure: sRN SCdevicestate

Example: sRN SCdevicestate

| CoLa A | ASCII | <STX> sRN {SPC} SCdevicestate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 53 43 64 65 76 69 63 65 73 74 61 74 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 52 4E 20 53 43 64 65 76 69 63 65 73 74 61 74 65 30 |

Table 383: Example: sRN SCdevicestate

| Telegram structure: sRA SCdevicestate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 13 | All | SCdevicestate | 53 43 64 65 76 69 63 65 73 74 61 74 65 |
| Status code | Code number | Enum_8 | 1 | All | Busy / logged-in: 0 Ready: 1 | 00 01 |
|  |  |  |  |  | Error: 2 | 02 |
|  |  |  |  | LMS5xx | Busy / logged-in: 0 | 00 |
|  |  |  |  | LMS1xx | Ready: 1 | 01 |
|  |  |  |  |  | Error: 2 | 02 |
|  |  |  |  |  | Standby: 3 | 03 |

Table 384: Telegram structure: sRA SCdevicestate

Example: sRA SCdevicestate

| CoLa A | ASCII | <STX> sRA {SPC} SCdevicestate {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 53 43 64 65 76 69 63 65 73 74 61 74 65 20 00 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 52 41 20 53 43 64 65 76 69 63 65 73 74 61 74 65 20 00 1F |

Table 385: Example: sRA SCdevicestate

    - Status commands for LD-XXX and NAV310

The following status commands will be explained in the subsequent sections:

      - LMCmeasstate: Status of the internal Statemachine

      - SCdevicestate: 	Status of the Sensors ( actual measurement status)

      - EMCustomerInfo: Additional error information

      - LDMSenStat: 	Status of the state machine of the measurment core, Motor status

How status commands for for LD-XXX and NAV310 work together:

      - If LMCmeasstate changes to "Idle” or an other status, although the measurement status "Measure2D” is expected, there is an error during the measurement (or during start up of the measurement).

      - SCdevicestate is always "Ready", if the measurement is active.

If "Busy” will be indicated the unit is not measuring ( e.g, IDLE). If there is any failure "Error” will be indicated. (However LMCmeasstate could indicate "Measure2D”, if the failure occurs during the measurement, because it is only an indication of the status of the State machine ).

      - In case of a failure EMCustomerInfo can provide an information about the error.

In case of an motor failure there are following condition visible:

        - Motor blocked during operation  DEVICE_FAILURE

        - Motor blocked during spin up  CHECK_PARAMETER

      - It is also possible to read LDMSenStat (and to register as an event). This value equals the Sensorstatus of the NAV310/LD­XXX. A status "B1” of the measurement core means "Motor error and Idle”).

      - During the measurement it is possible to monitor a deviation of the target rotation frequency. (If the device detects rotation values that are too slow, it will terminate the measurement.)

In case of an failure this value will not always be updated, therefore it is necessary to monitor LMCmeasstate and SCdevicestate in parallel.

NOTE

      - In case of an failure ( Scanner does not change to MEASURE2D or switches back to IDLE ), it is necessary to send the command LMCstopmeas ( even if the Status is indicated as IDLE)

      - If at EMCustomerInfo the message CHECK_PARAMETER is indicated, a reset is only possible by a power cycle of the scanner.

Ask for Device Measurement State

| Telegram structure: sRN LMCmeasstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Ask for measurement state | String | 12 | All | LMCmeasstate | 4C 4D 43 6D 65 61 73 73 74 61 74 65 |

Table 386: Telegram structure: sRN LMCmeasstate

Example: sRN LMCmeasstate

| CoLa A | ASCII | <STX> sRN {SPC} LMCmeasstate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 4D 43 6D 65 61 73 73 74 61 74 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 52 4E 20 4C 4D 43 6D 65 61 73 73 74 61 74 65 30 |

Table 387: Example: sRN LMCmeasstate

| Telegram structure: sRA LMCmeasstate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Report | String | 12 | All | LMCmeasstate | 4C 4D 43 6D 65 61 73 73 |
|  | measurement state |  |  |  |  | 74 61 74 65 |
| Status code | Current measurement state | Enum_16 | 2 | All | Idle: 3 Ready 2D: 6 | Idle: 0003 Ready 2D: 0006 |
|  |  |  |  |  | Measure 2D: 7 | Measure 2D: 0007 |
|  |  |  |  |  | Other state codes may show up during booting, firmware update or transition between states. |  |

Table 388: Telegram structure: sRA LMCmeasstate

Example: sRA LMCmeasstate is Measure 2D

| CoLa A | ASCII | <STX> sRA {SPC} LMCmeasstate {SPC} 7 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 4D 43 6D 65 61 73 73 74 61 74 65 00 07 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 52 41 20 4C 4D 43 6D 65 61 73 73 74 61 74 65 00 07 1F |

Table 389: Example: sRA LMCmeasstate is Measure 2D

Ask for customer info of sensor

This telegram will provide additional error information.

| Telegram structure: sRN EMCustomerInfo |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Ask for customer info | String | 14 | All | EMCustomerInfo | 45 4D 43 75 73 74 6F 6D 65 72 49 6E 66 6F |

Table 390: Telegram structure: sRN EMCustomerInfo

Example: sRN EMCustomerInfo

| CoLa A | ASCII | <STX> sRN {SPC} EMCustomerInfo <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 45 4D 43 75 73 74 6F 6D 65 72 49 6E 66 6F 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 4E 20 45 4D 43 75 73 74 6F 6D 65 72 49 6E 66 6F 4D |

Table 391: Example: sRN EMCustomerInfo

| Telegram structure: sRA EMCustomerInfo |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Report customer info | String | 14 | All | EMCustomerInfo | 45 4D 43 75 73 74 6F 6D 65 72 49 6E 66 6F |
| Status code | Customer info | Enum_16 | 2 | All | 0: DEVICE_ OK 1: DEFECTIVE_DEVICE | 0000: DEVICE_ OK 0001: DEFECTIVE_DEVICE |
|  |  |  |  |  | 2: DEVICE_TEMP_FAILURE 3: DEVICE_FAILURE | 0002: DEVICE_TEMP_FAILURE |
|  |  |  |  |  | 4: DEVICE_NOT_READY | 0003: DEVICE_FAILURE |
|  |  |  |  |  | 5: CHECK_PARAMETER | 0004: DEVICE_NOT_READY 0005: CHECK_PARAMETER |
|  |  |  |  |  | DEFECTIVE_DEVICE: Please return device to SICK |  |
|  |  |  |  |  | DEVICE_TEMP_FAILURED: Device failure. Please check temperature. |  |
|  |  |  |  |  | DEVICE_FAILURE : Please switch off for 20 seconds and power up again. |  |
|  |  |  |  |  | DEVICE_NOT_READY: Please wait. |  |
|  |  |  |  |  | CHECK_PARAMETER: Warning - please check parametrization. |  |

Table 392: Telegram structure: sRA EMCustomerInfo

Example: sRA EMCustomerInfo = Device OK

| CoLa A | ASCII | <STX> sRA {SPC} EMCustomerInfo {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 45 4D 43 75 73 74 6F 6D 65 72 49 6E 66 6F 20 00 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 52 4E 20 45 4D 43 75 73 74 6F 6D 65 72 49 6E 66 6F 20 00 6D |

Table 393: Example: sRA EMCustomerInfo = Device OK

Ask for Sensorstatus

This telegram provides status information of the State Machine of measurement core and the Motor Status

| Telegram structure: sRN LDMSenStat |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Ask for state | String | 10 | All | LDMSenStat | 4C 44 4D 53 65 6E 53 74 61 74 |

Table 394: Telegram structure: sRN LDMSenStat

Example: sRN LDMSenStat

| CoLa A | ASCII | <STX> sRN {SPC} LDMSenStat <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 44 4D 53 65 6E 53 74 61 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 52 4E 20 4C 44 4D 53 65 6E 53 74 61 74 60 |

Table 395: Example: sRN LDMSenStat

| Telegram structure: sRA LDMSenStat |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Report state | String | 10 | All | LDMSenStat | 4C 44 4D 53 65 6E 53 74 61 74 |

| Telegram structure: sRA LDMSenStat |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Status code | Current state regarding&hellip; | Uint_32 | 4 | All | Idle: 3 Ready 2D: 6 | Idle: 0003 Ready 2D: 0006 |
|  |  |  |  |  | Measure 2D: 7 | Measure 2D: 0007 |
|  |  |  |  |  | Other state codes may show up during booting, firmware update or transition between states. |  |
|  | Working mode |  | Bit 0&hellip;3 |  | Idle: 1 | Idle: 1 |
|  |  |  |  |  | Rotate: 2 | Rotate: 2 |
|  |  |  |  |  | Measure: 3 | Measure: 3 |
|  |  |  |  |  | Error: 4 | Error: 4 |
|  |  |  |  |  | (Other bits: reserved) | (Other bits: reserved) |
|  | Motor mode |  | Bit 4&hellip;7 |  | Motor OK: 0 | Motor OK: 0 |
|  |  |  |  |  | Motor spin to low: 4 | Motor spin to low: 4 |
|  |  |  |  |  | Motor spin to high: 9 | Motor spin to high: 9 |
|  |  |  |  |  | Motor stops or coder error: B | Motor stop or coder error: B |
|  |  |  |  |  | (Other bits: reserved) | (Other bits: reserved) |
|  | (Reserved) |  | Bit 8&hellip;31 |  | (Reserved) | (Reserved) |

Table 396: Telegram structure: sRA LDMSenStat

Example: sRA LDMSenStat Device in Idle mode

| CoLa A | ASCII | <STX> sRA {SPC} LDMSenStat {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 44 4D 53 65 6E 53 74 61 74 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 13 73 52 41 20 4C 44 4D 53 65 6E 53 74 61 74 20 00 00 00 01 4E |

Table 397: Example: sRA LDMSenStat Device is in Idle mode

    - Read device order number

This telegram reads the device order number.

| Telegram structure: sRN DIornr |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read state | String | 6 | All | DIornr | 44 49 6F 72 6E 72 |

Table 398: Telegram structure: sRN DIornr

Example: sRN DIornr

| CoLa A | ASCII | <STX> sRN {SPC} DIornr <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 44 49 6F 72 6E 72 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 52 4E 20 44 49 6F 72 6E 72 43 |

Table 399: Example: sRN DIornr

| Telegram structure: sRA DIornr |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 6 | All | DIornr | 44 49 6F 72 6E 72 |
| Order number | Order number in 7 digits | String | 7 | All | 0000000... 9999999 | 00 00 00 00 00 00 00 00... FF FF FF FF FF FF FF |

Table 400: Telegram structure: sRA DIornr

Example: sRA DIornr 1047782 (Order Number for LMS511-20100)

| CoLa A | ASCII | <STX> sRA {SPC} DIornr {SPC} 1047782 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 44 49 6F 72 6E 72 20 31 30 34 37 37 38 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 41 20 44 49 6F 72 6E 72 31 30 34 37 37 38 32 53 |

Table 401: Example for LMS511-20100: sRA DIornr

Example: sRA DIornr 1067299 (Order Number for TiM561-2050101)

| CoLa A | ASCII | <STX> sRA {SPC} DIornr {SPC} 1071419 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 44 49 6F 72 6E 72 20 31 30 37 31 34 31 39 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 41 20 44 49 6F 72 6E 72 31 30 37 31 34 31 39 57 |

Table 402: Example for TiM561-2050101: sRA DIornr

    - Read device type

This telegram asks for the device type.

| Telegram structure: sRN DItype |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Ask state | String | 6 | All | DItype | 44 49 74 79 70 65 |

Table 403: Telegram structure: sRN DItype

Example: sRN DItype

| CoLa A | ASCII | <STX> sRN {SPC} DItype <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 44 49 74 79 70 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 52 4E 20 44 49 74 79 70 65 5A |

Table 404: Example: sRN DItype

| Telegram structure: sRA DItype |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Ask state | String | 6 | All | DItype | 44 49 74 79 70 65 |
| Length of type key | Number of digits of the following type code length | Uint_8 | 1 | All | 0d... 255d (0h &hellip; FF) | 00... FF |
| Device type | Type code of the device | String | (var.) | All | (Device type) | (Device type) |

Table 405: Telegram structure: sRA DItype

Example for LMS511-20100

| CoLa A | ASCII | <STX> sRA {SPC} DItype {SPC} C {SPC} LMS511-20100 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 44 49 74 79 70 65 20 43 20 4C 4D 53 35 31 31 2D 32 30 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 18 73 52 41 20 44 49 74 79 70 65 20 0C 4C 4D 53 35 31 31 2D 32 30 31 30 30 00 |

Table 406: Example for LMS511-20100: sRA DItype

Example for TiM561-2050101

| CoLa A | ASCII | <STX> sRA {SPC} DItype {SPC} E {SPC} TIM561-2050101 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 44 49 74 79 70 65 20 45 20 54 49 4D 35 36 31 2D 32 30 35 30 31 30 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 1A 73 52 41 20 44 49 74 79 70 65 20 0E 54 49 4D 35 36 31 2D 32 30 35 30 31 30 31 03 |

Table 407: Example for TiM561-2050101: sRA DItype

    - Read operating hours

| Telegram structure: sRN ODoprh |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read operating hours | String | 6 | All | ODoprh | 4F 44 6F 70 72 68 |

Table 408: Telegram structure: sRN ODoprh

Example: sRN ODoprh

| CoLa A | ASCII | <STX> sRN {SPC} ODoprh <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4F 44 6F 70 72 68 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 52 4E 20 4F 44 6F 70 72 68 41 |

Table 409: Example: sRN ODoprh

| Telegram structure: sRA ODoprh |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read operating hours | String | 6 | All | ODoprh | 4F 44 6F 70 72 68 |
| Value | Operating hours in 1/10 h | Uint_32 | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

Table 410: Telegram structure: sRA ODoprh

Example: sRA ODoprh

| CoLa A | ASCII | <STX> sRA {SPC} ODoprh {SPC} 2DC8B <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4F 44 6F 70 72 68 20 32 44 43 38 42 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 52 41 20 4F 44 6F 70 72 68 20 00 02 DC 8B 36 |

Table 411: Example: sRA ODoprh

Calculation of the value: 2DC8B (hex)  187531 (dez) × 1/10 h = 18753.1 h

    - Read power on counter

| Telegram structure: sRN ODpwrc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read power on counter | String | 6 | All | ODpwrc | 4F 44 70 77 72 63 |

Table 412: Telegram structure: sRN ODpwrc

Example: sRN ODpwrc

| CoLa A | ASCII | <STX> sRN {SPC} ODpwrc <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4F 44 70 77 72 63 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 52 4E 20 4F 44 70 77 72 63 52 |

Table 413: Example: sRN ODpwrc

| Telegram structure: sRA ODpwrc |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read power on counter | String | 6 | All | ODpwrc | 4F 44 70 77 72 63 |
| Value | Power on counter | Uint_32 | 4 | All | 0h... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

Table 414: Telegram structure: sRA ODpwrc

Example: sRA ODpwrc

| CoLa A | ASCII | <STX> sRA {SPC} ODpwrc {SPC} 752D <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4F 44 70 77 72 63 20 752D 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 52 41 20 4F 44 70 77 72 63 20 00 00 75 2D 36 |

Table 415: Example: sRA ODpwrc

    - Read temperature

With this command the internal temperature of the device can be identified. Please note that it does not give an indication of the current ambient temperature.

| Telegram structure: sRN OPcurtmpdev |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read temperature of the device | String | 11 | All | OPcurtmpdev | 4F 50 63 75 72 74 6D 70 64 65 76 |

Table 416: Telegram structure: sRN OPcurtmpdev

Example: sRN OPcurtmpdev

| CoLa A | ASCII | <STX> sRN {SPC} OPcurtmpdev <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4F 50 63 75 72 74 6D 70 64 65 76 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 52 4E 20 4F 50 63 75 72 74 6D 70 64 65 76 2A |

Table 417: Example: sRN OPcurtmpdev

| Telegram structure: sRA OPcurtmpdev |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read temperature of the device | String | 11 | All | OPcurtmpdev | 4F 50 63 75 72 74 6D 70 64 65 76 |
| Tempera- ture data | [°C as float accord- ing to IEEE754] | Real | 4 | All | C2480000h... 42C80000h (-50 c... +100 c) | C2 48 00 00... 42C8 00 00 00 |

Table 418: Telegram structure: sRA OPcurtmpdev

Example: sRA OPcurtmpdev (35°C)

| CoLa A | ASCII | <STX> sRA {SPC} OPcurtmpdev {SPC} 420C0000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4F 50 63 75 72 74 6D 70 64 65 76 34 32 30 43 30 30 30 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 14 73 52 41 20 4F 50 63 75 72 74 6D 70 64 65 76 20 42 0C 00 00 4B |

Table 419: Example: sRA OPcurtmpdev

    - Set device name

| Telegram structure: sWN LocationName (Maintenance) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set device name | String | 12 | All | LocationName | 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 |
| Value | Array of characters of the following device name | Uint_16 | 2 | All | 0d... + 16d (0h &hellip; 10h) | 00 00... 00 10 |
| Value | Device name | String | 16 | All | [Device name] | [Device name] |

Table 420: Telegram structure: sWN LocationName

Example: sWN LocationName +13 OutdoorDevice

| CoLa A | ASCII | <STX> sWN {SPC} LocationName {SPC} +13 {SPC} OutdoorDevice <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 20 2B 31 33 20 4F 75 74 64 6F 6F 72 44 65 76 69 63 65 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 20 73 57 4E 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 20 00 0D 4F 75 74 64 6F 6F 72 44 65 76 69 63 65 1D |

Table 421: Example: sWN LocationName +13 OutdoorDevice

| Telegram structure: sWA LocationName |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set device name | String | 12 | All | LocationName | 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 |

Table 422: Telegram structure: sWA LocationName

Example: sWA LocationName

| CoLa A | ASCII | <STX> sWA {SPC} LocationName <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 20 73 57 41 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 74 |

Table 423: Example: sWA LocationName

    - Read device name

| Telegram structure: sRN LocationName |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read device name | String | 12 | All | LocationName | 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 |

Table 424: Telegram structure: sRN LocationName

Example: sRN LocationName

| CoLa A | ASCII | <STX> sRN {SPC} LocationName <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 52 4E 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 55 |

Table 425: Example: sRN LocationName

| Telegram structure: sRA LocationName |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read device name | String | 12 | All | LocationName | 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 |
| Value | Array of characters of the following device name | Uint_16 | 2 | All | 0d... + 16d (0h &hellip; 10h) | 00 00... 00 10 |
| Value | Device name | String | 16 | All | [Device name] | [Device name] |

Table 426: Telegram structure: sRA LocationName

Example: sRA LocationName

| CoLa A | ASCII | <STX> sRA {SPC} LocationName {SPC} D {SPC} OutdoorDevice <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 44 20 4F 75 74 64 6F 6F 72 44 65 76 69 63 65 03 |
| CoL a B | Binary | 02 02 02 02 00 00 00 17 73 52 41 20 4C 6F 63 61 74 69 6F 6E 4E 61 6D 65 00 0D 4F 75 74 64 6F 6F 72 44 65 76 69 63 65 20 |

Table 427: Example: sRA LocationName

    - Read angle compensation sine

| Telegram structure: sRN MCAngleCompSin |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read angle compensation sine | String | 14 | All | MCAngleCompSin | 4D 43 41 6E 67 6C 65 43 6F 6D 70 53 69 6E |

Table 428: Telegram structure: sRN MCAngleCompSin

Example: sRN MCAngleCompSin

| CoLa A | ASCII | <STX> sRN {SPC} MCAngleCompSin <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4D 43 41 6E 67 6C 65 43 6F 6D 70 53 69 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 4E 20 4D 43 41 6E 67 6C 65 43 6F 6D 70 53 69 6E 65 |

Table 429: Example: sRN MCAngleCompSin

| Telegram structure: sRA MCAngleCompSin |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Angle compensation sine | String | 14 | All | MCAngleCompSin | 4D 43 41 6E 67 6C 65 43 6F 6D 70 53 69 6E |
| Amplitude | [1/10000 °] | Int_16 | 2 | All | -10000d... + 10000d (D8F0h &hellip; 2710h) | D8 F0... 27 10 |
| Phase | [1/10000 °] | Int_32 | 4 | All | -3600000d... + 3600000d (FFC91180h &hellip; 36 EE80h) | FF C9 11 80&hellip; 00 36 EE 80 |
| Offset | [1/10000 °] | Int_16 | 2 | All | -10000d... + 10000d (D8F0h &hellip; 2710h) | D8 F0... 27 10 |

Table 430: Telegram structure: sRA MCAngleCompSin

Example: sRA MCAngleCompSin

| CoLa A | ASCII | <STX> sRA {SPC} MCAngleCompSin {SPC} 0 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4D 43 41 6E 67 6C 65 43 6F 6D 70 53 69 6E 20 30 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 18 73 52 41 20 4D 43 41 6E 67 6C 65 43 6F 6D 70 53 69 6E 20 00 00 00 00 00 00 00 00 4A |

Table 431: Example: sRA MCAngleCompSin

The values of the angular compensation could be retrieved form the memory of the NAV310 to improve the angular measurement accuracy.

The applied formula is:

AngleComp = AngleRaw + (AngleCompAmp * sin(AngleRaw - AngleCompPhase) + AngleCompOffset

Example (C coded):

angleRaw: Raw angle as float in degrees (0.000 &hellip; 359999) angleComp:Compensated angle as float in degrees (0.000 &hellip; 359999) AngleCompAmp

AngleCompPhase

AngleCompOffset: Compensation parameters as int in 1/1000 degrees float compensateAngle(float angleRaw)

{

float angleComp;

angleRaw += ((float) AngleCompOffset)/1000.0; angleRaw += (((float) AngleCompAmp)/1000.0) *

sin((DEGTORAD * (angle - ((float) AngleCompPhase)/1000.0))); return angleComp;

}

    - Reset output counter

| Telegram structure: sMN LIDrstoutpcnt (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Reset output counter | String | 13 | All | LIDrstoutpcnt | 4C 49 44 72 73 74 6F 75 74 70 63 6E 74 |

Table 432: Telegram structure: sMN LIDrstoutpcnt

Example: sMN LIDrstoutpcnt

| CoLa A | ASCII | <STX> sMN {SPC} LIDrstoutpcnt <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 4C 49 44 72 73 74 6F 75 74 70 63 6E 74 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 11 73 4D 4E 20 4C 49 44 72 73 74 6F 75 74 70 63 6E 74 03 |

Table 433: Example: sMN LIDrstoutpcnt

| Telegram structure: sAN LIDrstoutpcnt |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Reset state | String | 13 | All | LIDrstoutpcnt | 4C 49 44 72 73 74 6F 75 74 70 63 6E 74 |
| Status code | Code number | Bool_1 | 1 | All | Success: 0 Error: 1 | Success: 00 Error: 01 |

Table 434: Telegram structure: sAN LIDrstoutpcnt

Example: sAN LIDrstoutpcnt

| CoLa A | ASCII | <STX> sAN {SPC} LIDrstoutpcnt {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 4C 49 44 72 73 74 6F 75 74 70 63 6E 74 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 41 4E 20 4C 49 44 72 73 74 6F 75 74 70 63 6E 74 20 00 2F |

Table 435: Example: sAN LIDrstoutpcnt

    - Read heating state

Note

It is not allowed to request this telegram in a faster cycle than 10 ms!

| Telegram structure: sRN OPheatstateext (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 14 | All | OPheatstateext | 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 |

Table 436: Telegram structure: sRN OPheatstateext

Example: sRN OPheatstateext

| CoLa A | ASCII | <STX> sRN {SPC} OPheatstateext <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 52 4E 20 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 56 |

Table 437: Example: sRN OPheatstateext

| Telegram structure: sRA OPheatstateext (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 14 | All | OPheatstateext | 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 |

| Telegram structure: sRA OPheatstateext (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Heating state | Heating state | Enum_8 | 1 | All | Heating active : 0d Electrical current for heating too low: 1d | Heating active : 00 Electrical current for heating too low: 01 |
|  |  |  |  |  | Heating inactive: 2d | Heating inactive: 02 |
|  |  |  |  |  | Power supply for heating not connected: 3d | Power supply for heating not connected: 03 |

Table 438: Telegram structure: sRA OPheatstateext

Example: sRA OPheatstateext 2

| CoLa A | ASCII | <STX> sRA {SPC} OPheatstateext {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 14 73 52 41 20 4F 50 68 65 61 74 73 74 61 74 65 65 78 74 20 02 7B |

Table 439: Example: sRA OPheatstateext 2

    - FieldEvalLoggingList by Event

Start/Stop of the Event.

| Telegram structure: sEN FieldEvalLoggingList |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Event | String | 3 | All | sEN | 73 45 4E |
| Command |  | String | 20 | All | FieldEvalLoggingList | 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 |
|  |  | Bool_1 | 1 | All | Start: 1 Stop: 0 | Start: 01 Stop: 00 |

Table 440: Telegram structure: sEN FieldEvalLoggingList

Example: sEN FieldEvalLoggingList 1

Explanation

| CoLa A | ASCII | <STX> sEN {SPC} FieldEvalLoggingList {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 4E 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 1A 73 45 4E 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 20 0 1 64 |

Table 441: Example sEN FieldEvalLoggingList

Acknowledge for Start/Stop of the Event

| Telegram structure: sEA FieldEvalLoggingList |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sEA | 73 45 41 |
| Command |  | String | 20 | All | FieldEvalLoggingList | 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 |
|  |  | Bool_1 | 1 | All | Start: 1 Stop: 0 | Start: 01 Stop: 00 |

Table 442: Telegram structure: sEA FieldEvalLoggingList

Example: sEA FieldEvalLoggingList 1

Explanation

| CoLa A | ASCII | <STX> sEA {SPC} FieldEvalLoggingList {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 41 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 1A 73 45 41 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 20 0 1 6B |

Table 443: Example: sEA FieldEvalLoggingList 1

This is a readonly telegram

| Telegram |  |  | structure: | sSN FieldEvalLogging List |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type |  | Answer | String | 3 | All | sSN | 73 53 4E |
| Command |  |  | String | 20 | All | FieldEvalLogging List | 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 |
| Array size |  | 0 .. 100 | Uint_16 | 2 | All | 0h .. 00 64h | 00 00... 00 64 |
| Index of Array |  | 0 .. 65535 | Uint_16 | 2 | All | 0h..FFFFh | 00 00..FF FF |
|  | Year | 0 .. 2105 | Uint_16 | 2 | All | 0h .. 839h | 00 00 .. 0839 |
|  | Month | 1 .. 12 | Uint_8 | 1 | All | 1h..Ch | 01 .. 0C |
|  | Day | 1 .. 31 | Uint_8 | 1 | All | 1h .. 1Fh | 01 .. 1F |
|  | Hour | 0 .. 23 | Uint_8 | 1 | All | 0h .. 17h | 00 .. 17 |
|  | Minute | 0 .. 59 | Uint_8 | 1 | All | 0h .. 3Bh | 00 .. 3B |
|  | second | 0 .. 59 | Uint_8 | 1 | All | 0h .. 3Bh | 00 .. 3B |
|  | microsecond | 0 .. 999 | Uint_32 | 4 | All | 0h .. 3E7 | 00 .. 03E7 |
|  | EvaluationNumb er | 1 .. 255 | Uint_8 | 1 | All | 0h..FFh | 00..FF |
|  | X Coordinate | X-Position of first violation in world coordinates [mm] | Int_32 | 4 | All |  |  |
|  |  | -2147483648 |  |  |  |  |  |
|  |  | .. 2147483647 |  |  |  |  |  |
|  | Y Coordinate | Y-Position of first violation in world coordinates [mm] | Int_32 | 4 | All |  |  |
|  |  | -2147483648 |  |  |  |  |  |
|  |  | .. 2147483647 |  |  |  |  |  |
|  | Z Coordinate | Z-Position of first violation in world coordinates [mm] | Int_32 | 4 | All |  |  |
|  |  | -2147483648 |  |  |  |  |  |
|  |  | .. 2147483647 |  |  |  |  |  |
|  | RadialDistance | 0 .. 4294967295 | Uint_32 | 4 | All | 0h&hellip;FF FF FF FFh | 00 00 00 00..FF FF FF FF |
|  | HorizontalAngle | -1800000 .. 1800000 | Int_32 | 4 | All |  |  |
|  | VerticalAngle | -9000 .. 9000 | Int_16 | 2 | All |  |  |
|  | ObjectSize in mm | 0 .. 4294967295 | Uint_32 | 4 | All | 0h&hellip;FF FF FF FFh | 00 00 00 00..FF FF FF FF |
|  | Reserved 1 | 0 .. 4294967295 | Uint_32 | 4 | All | 0h&hellip;FF FF FF FFh | 00 00 00 00..FF FF FF |

| Telegram structure: sSN FieldEvalLogging List |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  |  |  |  |  |  | FF |
|  | Reserverd 2 | 0 .. 4294967295 | Uint_32 | 4 | All | 0h&hellip;FF FF FF FFh | 00 00 00 00..FF FF FF FF |

Table 444: Telegram structure: sSN Field EvalLogging List

Example: sSN Field EvalLogging List

Explanation

| CoLa A | ASCII | <STX >sSN {SPC} FieldEvalLoggingList {SPC} 64 {SPC} 1 {SPC} 7E4 {SPC} C {SPC} 16 {SPC} 9 {SPC} 3A {SPC} 20 {SPC} 13 {S PC} 1 {SPC} 140 {SPC} FFFFFEEC {SPC} 0 {SPC} 1A7 {SPC} FFF9B4AC {SPC} 0 {SPC} E2 {SPC} 0 {SPC} 0 {SPC} 2&hellip;&hellip; <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 53 4E 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 20 36 34 20 31 20 37 45 34 20 43 20 31 36 20 39 20 33 41 20 32 30 20 31 33 20 31 20 31 34 30 20 46 46 46 46 46 45 45 43 20 30 20 31 41 37 20 46 46 46 39 42 34 41 43 20 30 20 45 32 20 30 20 30 20 32 .... 03 |
| CoLa B | Binary | 02 02 02 00 00 12 13 73 53 4E 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 20 00 64 00 01 07 E4 0C 16 09 3A 20 00 13 01 00 00 01 40 FF FF FE EC 00 00 00 00 00 00 01 A7 FF F9 B4 AC 00 00 00 00 00 E2 00 00 00 00 00 00 00 00 &hellip;&hellip; 00 00 7A |

Table 445: Example: sSN Field EvalLogging List

    - FieldEvalLoggingList on Request

Request the actual Field Eval logging List

| Telegram structure: sRN FieldEvalLoggingList |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Event | String | 3 | All | sRN | 73 52 4E |
| Command |  | String | 20 | All | FieldEvalLoggingList | 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 |

Table 446: Telegram structure: sRN FieldEvalLoggingList

Example: sRN FieldEvalLoggingList

Explanation

| CoLa A | ASCII | <STX> sRN {SPC}FieldEvalLoggingList<ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 18 73 52 4E 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 52 |

Table 447: Example sEN FieldEvalLoggingList 1

| Telegram structure: sRA FieldEvalLoggingList |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command |  | String | 20 | All | FieldEvalLoggingList | 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 |
| Please check the list in SN Field EvalLogging List. Same content |  |  |  |  |  |  |

Table 448: Telegram structure: sRA FieldEvalLoggingList

Example: sRA FieldEvalLoggingList

Explanation

| CoLa A | ASCII | <STX> sRA {SPC} FieldEvalLoggingList {SPC} 64 {SPC} 1 {SPC} 7E4 {SPC} C {SPC} 16 {SPC} 9 {SPC} 3A {SPC} 20 {SPC} 13 {SP C} 1 {SPC} 140 {SPC} FFFFFEEC {SPC} 0 {SPC} 1A7 {SPC} FFF9B4AC {SPC} 0 {SPC} E2 {SPC} 0 {SPC} 0 {SPC} 2&hellip;&hellip;. <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 36 34 20 31 20 37 45 34 20 43 20 31 36 20 39 20 33 41 20 32 30 20 31 33 20 31 20 31 34 30 20 46 46 46 46 46 45 45 43 20 30 20 31 41 37 20 46 46 46 39 42 34 41 43 20 30 20 45 32 20 30 20 30 20 32 20 30 20 20 32... .. 03 |
| CoLa B | Binary | 02 02 02 00 00 12 13 73 52 41 20 46 69 65 6C 64 45 76 61 6C 4C 6F 67 67 69 6E 67 4C 69 73 74 20 00 64 00 01 07 E4 0C 16 09 3A 20 00 13 01 00 00 01 40 FF FF FE EC 00 00 00 00 00 00 01 A7 FF F9 B4 AC 00 00 00 00 00 E2 00 00 00 00 00 00 00 00 00 02 &hellip;&hellip; F5 |

Table 449: Example: sEA FieldEvalLoggingList 1

  - Interfaces

Telegram validity overview

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.9.1 | Set IP address | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.9.2 | Read IP address | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.9.3 | Set Ethernet gateway | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.9.4 | Read Ethernet gateway | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.9.5 | Set IP mask | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.9.6 | Read IP mask | x | x | x | x | x | x | x | x | x | x | x | x | x |
| 4.9.7 | Set baud rate for host interface | x | x |  |  |  | x |  | x |  |  |  |  |  |
| 4.9.8 | Read baud rate of host interface | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.9.9 | Set interface type | x | x |  | x | x |  |  |  |  |  |  |  |  |
| 4.9.10 | Read interface type | x | x |  | x | x |  |  |  |  |  |  |  |  |
| 4.9.11 | Set Host port number | x | x |  | x | x | x | x | x | x | x | x | x | x |
| 4.9.12 | Set Host port Command Language (CoLa dialect) | x | x |  | x | x | x | x | x | x | x | x | x | x |
| 4.9.13 | Set Motor synchronization |  |  |  |  |  |  |  |  |  |  |  | x |  |
| 4.9.14 | Set Phase Shift of Slave for motor synchronization |  |  |  |  |  |  |  |  |  |  |  | x |  |
| 4.9.15 | Enable/Disable Front Panel | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.9.16 | Set function front panel | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.9.17 | Set front LEDs | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.9.18 | Set function of LED1 |  |  |  |  |  | x | x |  |  |  |  | x |  |
| 4.9.19 | Set function of LED2 |  |  |  |  |  | x | x |  |  |  |  |  |  |
| 4.9.20 | Switch on/off LED1 or LED2 |  |  |  |  |  | x | x |  |  |  |  |  |  |
| 4.9.21 | Set 7-segment display to specific symbol or number | x | x |  |  |  |  |  |  |  |  |  |  |  |

    - Set IP address

IMPORTANT

 	Save permanently to set values. Changes will be active after rebooting the device.

 		Settings must correspond with network in which scanner is used. Else device cannot be found any more.

| Telegram structure: sWN EIIpAddr (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set IP address | String | 8 | All | EIIpAddr | 45 49 49 50 41 64 64 72 |
| IP address | Set values in hex | Uint_32 | 4 | All | 00 00 00 00h... FF FF FF FFh (decimal values unwieldy) | 00 00 00 00 00... FF FF FF FF |

Table 450: Telegram structure: sWN EIIpAddr

Example: sWN EIIpAddr 192.168.0.2

| C o La A | ASCII | <STX> sWN {SPC} EIIpAddr {SPC} C0 {SPC} A8 {SPC} 0 {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 49 70 41 64 64 72 20 43 30 20 41 38 20 30 20 32 03 |
| C oLa B | Binary | 02 02 02 00 00 00 11 73 57 4E 20 45 49 49 70 41 64 64 72 20 C0 A8 00 02 05 |

Table 451: Example: sWN EIIpAddr 192.168.0.2

| Telegram structure: sWA EIIpAddr |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set IP address | String | 8 | All | EIIpAddr | 45 49 49 50 41 64 64 72 |

Table 452: Telegram structure: sWA EIIpAddr

Example: sWA EIIpAddr

| CoLa A | ASCII | <STX> sWA {SPC} EIIpAddr <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 49 70 41 64 64 72 03 |
| CoLa B | Binary | 02 02 02 02 00 00 0D 73 57 41 20 45 49 49 70 41 64 64 72 63 |

Table 453: Example: sWA EIIpAddr

    - Read IP address

| Telegram structure: sRN EIIpAddr |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read IP address | String | 8 | All | EIIpAddr | 45 49 49 50 41 64 64 72 |

Table 454: Telegram structure: sRN EIIpAddr

Example: srN EIIpAddr

| CoLa A | ASCII | <STX> sRN {SPC} EIIpAddr <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 49 70 41 64 64 72 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 52 4E 20 45 49 49 70 41 64 64 72 49 |

Table 455: Example: srN EIIpAddr

| Telegram structure: sRA EIIpAddr |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Set IP address | String | 8 | All | EIIpAddr | 45 49 49 50 41 64 64 72 |
| IP address | Set values in hex | Uint_32 | 4 | All | 00 00 00 00h... FF FF FF FFh (decimal values unwieldy) | 00 00 00 00 00... FF FF FF FF |

Table 456: Telegram structure: sRA EIIpAddr

Example: sRA EIIpAddr 192.168.0.2

| CoLa A | ASCII | <STX> sRA {SPC} EIIpAddr {SPC} C0 {SPC} A8 {SPC} 00 {SPC} 02 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 49 70 41 64 64 72 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 52 41 20 45 49 49 70 41 64 64 72 20 C0 A8 00 02 0C |

Table 457: Example: sRA EIIpAddr 192.168.0.2

    - Set Ethernet gateway

Change Ethernet gateway IP address (TCP/IP)

IMPORTANT

 	Save permanently to set values. Changes will be active after rebooting the device.

 	Settings must correspond with network in which scanner is used. Else device cannot be found any more.

| Telegram structure: sWN EIgate (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set gateway adress | String | 6 | All | EIgate | 45 49 67 61 74 65 |
| Gateway address | Set values in hex | Uint_32 | 4 | All | 00 00 00 00h... FF FF FF FFh (decimal values unwieldy) | 00 00 00 00 00... FF FF FF FF |

Table 458: Telegram structure: sWN Elgate

Example: sWN EIgate 192.168.0.1

| CoLa A | ASCII | <STX> sWN {SPC} EIgate {SPC} C0 {SPC} A8 {SPC} 00 {SPC} 01 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 67 61 74 65 20 43 30 20 41 38 20 30 30 20 30 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 57 4E 20 45 49 67 61 74 65 20 C0 A8 00 01 5A |

Table 459: Example: sWN EIgate 192.168.0.1

| Telegram structure: sWA EIgate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set gateway address | String | 6 | All | EIgate | 45 49 67 61 74 65 |

Table 460: Telegram structure: sWA Elgate

Example: sWA EIgate

| CoLa A | ASCII | <STX> sWA {SPC} EIgate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 67 61 74 65 03 |
| CoLa B | Binary | 02 02 02 02 00 00 0A 73 57 41 20 45 49 67 61 74 65 5E |

Table 461: Example: sWA EIgate

    - Read Ethernet gateway

Read for the Ethernet gateway (TCP/IP)

| Telegram structure: sRN Elgate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read gateway address | String | 6 | All | EIgate | 45 49 67 61 74 65 |

Table 462: Telegram structure: sRN Elgate

Example: sRN EIgate

| CoLa A | ASCII | <STX> sRN {SPC} EIgate <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 45 49 67 61 74 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 52 4E 20 45 49 67 61 74 65 54 |

Table 463: Example: sRN EIgate

| Telegram structure: sRA EIgate |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read gateway address | String | 6 | All | EIgate | 45 49 67 61 74 65 |
| Gateway address | Values in hex | Uint_32 | 4 | All | 00 00 00 00h... FF FF FF FFh (decimal values unwieldy) | 00 00 00 00 00... FF FF FF FF |

Table 464: Telegram structure: sRA Elgate

Example: sRA EIgate 192.168.0.1

| CoLa A | ASCII | <STX> sRA {SPC} EIgate {SPC} C0 {SPC} A8 {SPC} 00 {SPC} 01 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 45 49 67 61 74 65 20 C0 A8 00 01 03 |
| CoLa B | Binary | 02 02 02 02 00 00 0F 73 52 41 20 45 49 67 61 74 65 20 C0 A8 00 01 12 |

Table 465: Example: sRA EIgate 192.168.0.1

    - Set IP mask

IMPORTANT

 	Save permanently to set values. Changes will be active after rebooting the device.

 	Settings must correspond with network in which scanner is used. Else device cannot be found any more.

| Telegram structure: sWN EImask (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set IP mask | String | 6 | All | EImask | 45 49 6D 61 73 6B |
| IP mask | Set values in hex | Uint_32 | 4 | All | 00 00 00 00h... FF FF FF FFh (decimal values unwieldy) | 00 00 00 00 00... FF FF FF FF |

Table 466: Telegram structure: sWN EImask

Example: sWN EImask 255.255.254.0

| CoLa A | ASCII | <STX> sWN {SPC} EImask {SPC} FF {SPC} FF {SPC} FE {SPC} 00 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 6D 61 73 6B 20 46 46 20 46 46 20 46 45 20 30 30 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 57 4E 20 45 49 6D 61 73 6B 20 FF FF FE 00 8C |

Table 467: Example: sWN EImask 255.255.254.0

| Telegram structure: sWA EImask |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set IP mask | String | 6 | All | EImask | 45 49 6D 61 73 6B |

Table 468: Telegram structure: sWA EImask

Example: sWA EImask

| CoLa A | ASCII | <STX> sWA {SPC} EImask <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 6D 61 73 6B 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 57 41 20 45 49 6D 61 73 6B 63 |

Table 469: Example: sWA EImask

    - Read IP mask

| Telegram structure: sRN EImask |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read IP mask | String | 6 | All | EImask | 45 49 6D 61 73 6B |

Table 470: Telegram structure: sRN Elmask

Example: sRN EImask

| C oLa A | ASCII | <STX> sRN {SPC} EImask <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 45 49 6D 61 73 6B 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0A 73 52 4E 20 45 49 6D 61 73 6B 57 |

Table 471: Example: sRN EImask

| Telegram structure: sRA EImask |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read IP mask | String | 6 | All | EImask | 45 49 6D 61 73 6B |
| IP mask | Values in hex | Uint_32 | 4 | All | 00 00 00 00h... FF FF FF FFh (decimal values unwieldy) | 00 00 00 00 00... FF FF FF FF |

Table 472: Telegram structure: sRA Elmask

Example: sRA EImask 255.255.254.0

| CoLa A | ASCII | <STX> sRA {SPC} EImask {SPC} FF {SPC} FF {SPC} FE {SPC} 00 <ETX> <STX> sRN {SPC} EImask <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 45 49 6D 61 73 6B 20 45 49 6D 61 73 6B 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 52 41 20 45 49 6D 61 73 6B 20 FF FF FE 00 86 |

Table 473: Example: sRA EImask 255.255.254.0

    - Set baud rate for host interface

| Telegram structure: sWN SIHstBaud (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set baud rate for host interface | String | 9 | All | SIHstBaud | 53 49 48 73 74 42 61 75 64 |
| Baud rate data | Baud rate data for host interface | Enum_8 | 1 | All | 9600: + 5d (05h) 19200: + 6d (06h) | 9600: 05 19200: 06 |
|  |  |  |  |  | 38400: + 7d (07h) | 38400: 07 |
|  |  |  |  |  | 57600: + 8d (08h) | 57600: 08 |
|  |  |  |  |  | 115200: + 9d (09h) | 115200: 09 |
|  |  |  |  | LMS1xx, | 250000: + 10d (0Ah) | 250000: 0A |
|  |  |  |  | LMS5xx | 500000: + 11d (0Bh) | 500000: 0B |

Table 474: Telegram structure: sWN SIHstBaud

Example: sWN SIHstBaud

| CoLa A | ASCII | <STX> sWN {SPC} SIHstBaud {SPC} +8 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 53 49 48 73 74 42 61 75 64 20 08 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 57 4E 20 53 49 48 73 74 42 61 75 64 20 08 05 |

Table 475: Example: sWN SIHstBaud

| Telegram structure: sWA SIHstBaud |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set baud rate for host interface | String | 9 | All | SIHstBaud | 53 49 48 73 74 42 61 75 64 |

Table 476: Telegram structure: sWA SIHstBaud

Example: sWA SIHstBaud

| CoLa A | ASCII | <STX> sWA {SPC} SIHstBaud <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 53 49 48 73 74 42 61 75 64 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 41 20 53 49 48 73 74 42 61 75 64 20 02 |

Table 477: Example: sWA SIHstBaud

    - Read baud rate of host interface

| Telegram structure: sRN SIHstBaud |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read baud rate of host interface | String | 9 | All | SIHstBaud | 53 49 48 73 74 42 61 75 64 |

Table 478: Telegram structure: sRN SIHstBaud

Example: sRN SIHstBaud

| CoLa A | ASCII | <STX> sRN {SPC} SIHstBaud <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 53 49 48 73 74 42 61 75 64 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 52 4E 20 53 49 48 73 74 42 61 75 64 28 |

Table 479: Example: sRN SIHstBaud

| Telegram structure: sRA SIHstBaud |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read baud rate of host interface | String | 9 | All | SIHstBaud | 53 49 48 73 74 42 61 75 64 |
| Baud rate data | Baud rate data of host interface | Enum_8 | 1 | All | 9600: 5d (05h) 19200: 6d (06h) | 9600: 05 19200: 06 |
|  |  |  |  |  | 38400: 7d (07h) | 38400: 07 |
|  |  |  |  |  | 57600: 8d (08h) | 57600: 08 |
|  |  |  |  |  | 115200: 9d (09h) | 115200: 09 |
|  |  |  |  | LMS1xx, | 250000: 10d (0Ah) | 250000: 0A |
|  |  |  |  | LMS5xx | 500000: 11d (0Bh) | 500000: 0B |

Table 480: Telegram structure: sRA SIHstBaud

Example: sRA SIHstBaud

| CoLa A | ASCII | <STX> sRA {SPC} SIHstBaud {SPC} 8 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 53 49 48 73 74 42 61 75 64 20 08 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 52 41 20 53 49 48 73 74 42 61 75 64 20 08 0F |

Table 481: Example: sRA SIHstBaud

    - Set interface type

| Telegram structure: sWN SIHstHw (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set hardware settings for host interface | String | 7 | All | SIHstHw | 53 49 48 73 74 48 77 |
| Interface type data | Hardware settings data for host interface | Enum_8 | 1 | All | TX_RS232: 0 TX_RS485_2WIRE: 1 TX_RS422_485_4WIRE: 2 | TX_RS232: 00 TX_RS485_2WIRE: 01 TX_RS422_485_4WIRE: 02 |

Table 482: Telegram structure: sWN SIHstHw

Example: sWN SIHstHw

| CoLa A | ASCII | <STX> sWN {SPC} SIHstHw {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 53 49 48 73 74 48 77 20 00 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 57 4E 20 53 49 48 73 74 48 77 20 00 00 |

Table 483: Example: sWN SIHstHw

| Telegram structure: sWA SIHstHw |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set hardware settings for host interface | String | 7 | All | SIHstHw | 53 49 48 73 74 48 77 |

Table 484: Telegram structure: sWA SIHstHw

Example: sWA SIHstHw

| CoLa A | ASCII | <STX> sWA {SPC} SIHstHw <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 53 49 48 73 74 48 77 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 41 20 53 49 48 73 74 48 77 20 0F |

Table 485: Example: sWA SIHstHw

    - Read interface type

| Telegram structure: sRN SIHstHw |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Read hardware settings of host interface | String | 7 | All | SIHstHw | 53 49 48 73 74 48 77 |

Table 486: Telegram structure: sRN SIHstHw

Example: sRN SIHstHw

| CoLa A | ASCII | <STX> sRN {SPC} SIHstHw <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 53 49 48 73 74 48 77 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0B 73 52 4E 20 53 49 48 73 74 48 77 25 |

Table 487: Example: sRN SIHstHw

| Telegram structure: sRA SIHstHw |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Read hardware settings of host interface | String | 7 | All | SIHstHw | 53 49 48 73 74 48 77 |
| Interface type data | Hardware settings data of host interface | Enum_8 | 1 | All | TX_RS232: 0 TX_RS485_2WIRE: 1 TX_RS422_485_4WIRE: 2 | TX_RS232: 00 TX_RS485_2WIRE: 01 TX_RS422_485_4WIRE: 02 |

Table 488: Telegram structure: sRA SIHstHw

Example: sRA SIHstHw

| CoLa A | ASCII | <STX> sRA {SPC} SIHstHw {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 53 49 48 73 74 48 77 20 00 03 |
| CoLa B | Binary | 02 02 02 00 00 0D 73 52 41 20 53 49 48 73 74 48 77 20 00 0A |

Table 489: Example: sRA SIHstHw

    - Set Host port number

Note

Command for UDP port = EIUDPPort (45 49 55 44 50 50 6F 72 74)

| Telegram structure: sWN EIHstPort (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sWN | 73 57 4E |
| Command | Read hardware settings of host interface | String | 7 | All | EIHstPort | 45 49 48 73 74 50 6F 72 74 |
| Port number | Host port number (Default = 2112) | Uint_16 | 2 | All | + 0d... +65535d (0h &hellip; FF FFh) | 00 00... FF FF |

Table 490: Telegram structure: sWN EIHstPort

Example: sWN EIHstPort +2110 (Host Port 2110)

| CoLa A | ASCII | <STX> sWN {SPC} EIHstPort {SPC} +2110 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 48 73 74 50 6F 72 74 20 08 3E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 57 4E 20 45 49 48 73 74 50 6F 72 74 20 08 3E 26 |

Table 491: Example: sWN EIHstPort +2110

Example: sWN EIUDPPort +2214 (UDP Port 2214)

| CoLa A | ASCII | <STX> sWN {SPC} EIUDPPort {SPC} +2214 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 55 44 50 50 6F 72 74 20 08 A6 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 57 4E 20 45 49 55 44 50 50 6F 72 74 20 08 A6 B0 |

Table 492: Example: sWN EIUDPPort +2214

| Telegram structure: sWA EIHstPort |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Read hardware settings of host interface | String | 7 | All | EIHstPort | 45 49 48 73 74 50 6F 72 74 |

Table 493: Telegram structure: sWA EIHstPort

Example: sWA EIHstPort (Host Port)

| CoLa A | ASCII | <STX> sWA {SPC} EIHstPort <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 48 73 74 50 6F 72 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 41 20 45 49 48 73 74 50 6F 72 74 20 1F |

Table 494: Example: sWA EIHstPort

Example: sWA EIUDPPort (UDP Port)

| CoLa A | ASCII | <STX> sWA {SPC} EIUDPPort <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 55 44 50 50 6F 72 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 41 20 45 49 55 44 50 50 6F 72 74 20 11 |

Table 495: Example: sWA EIUDPPort

    - Set Host port Command Language (CoLa dialect)

Note

It is not allowed to use this telegram in a faster cycle than 10 ms!

After switching the CoLa dialect by this telegram, you have to store the changes permanentely and reboot the sensor to activate the chosen CoLA dialect.

Binary CRC32 is available since the firmware version V1.80.0 Command for UDP port = EIUDPCola (45 49 55 44 50 43 6F 6C 61)

| Telegram structure: sWN EIHstCola (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sWN | 73 57 4E |
| Command | Read hardware settings of host interface | String | 7 | All | EIHstCola | 45 49 48 73 74 43 6F 6C 61 |
| Command language | Host port Command language | Enum_8 | 1 | All | CoLa ASCII: 0 CoLa binary: 1 | 00 01 |
|  |  |  |  | LMS1xx, LMS5xx LRS4000 | CoLa ASCII: 0 CoLa binary: 1 Binary CRC32: 2 | 00 01 02 |
|  |  |  |  | LRS4000 | CoLa ASCII: 0 CoLa binary: 1 | 00 01 |

Table 496: Telegram structure: sWN EIHstCola

Example: sWN EIHstCola 1 (Host Port CoLa binary)

| CoLa A | ASCII | <STX> sWN {SPC} EIHstCola {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 48 73 74 43 6F 6C 61 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 57 4E 20 45 49 48 73 74 43 6F 6C 61 20 01 09 |

Table 497: Example: sWN EIHstCola 1

Example: sWN EIUDPCola 0 (UDP Port CoLa ASCII)

| CoLa A | ASCII | <STX> sWN {SPC} EIUDPCola {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 45 49 55 44 50 43 6F 6C 61 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 57 4E 20 45 49 55 44 50 43 6F 6C 61 20 00 06 |

Table 498: Example: sWN EIUDPCola 0

| Telegram structure: sWA EIHstCola |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Read hardware settings of host interface | String | 7 | All | EIHstCola | 45 49 48 73 74 43 6F 6C 61 |

Table 499: Telegram structure: sWA EIHstCola

Example: sWA EIHstCola (Host Port)

| CoLa A | ASCII | <STX> sWA {SPC} EIHstCola <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 48 73 74 43 6F 6C 61 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 41 20 45 49 48 73 74 43 6F 6C 61 20 07 |

Table 500: Example: sWA EIHstCola

Example: sWA EIUDPCola (UDP Port

| CoLa A | ASCII | <STX> sWA {SPC} EIUDPCola <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 45 49 55 44 50 43 6F 6C 61 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0E 73 57 41 20 45 49 55 44 50 43 6F 6C 61 20 09 |

Table 501: Example: sWA EIUDPCola

    - Set Motor synchronization

To increase the width of the field for the intended application or to avoid shadowing effects caused by geometrical properties, mulitple sensors (one Master, several Slaves) can be assembled side by side. For further infomration see operating instructions.

| Telegram structure: sWN SYtype (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set motor sycronization | String | 6 | All | SYtype | 53 59 74 79 70 65 |
| Function | Code number | Enum_8 | 1 | All | No function: 0 | 00 |
|  |  |  |  |  | Master: 1 | 01 |
|  |  |  |  |  | Slave: 2 | 02 |

Table 502: Telegram structure: sWN SYtype

Example: sWN SYtype

| CoLa A | ASCII | <STX> sWN {SPC} SYtype {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 53 59 74 79 70 65 20 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 4E 20 53 59 74 79 70 65 20 02 7A |

Table 503: Example: sWN SYtype 2

| Telegram structure: sWA SYtype |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set motor sycronization | String | 11 | All | SYtype | 53 59 74 79 70 65 |

Table 504: Telegram structure: sWA SYtype

Example: sWA SYtype

| CoLa A | ASCII | <STX> sWA {SPC} SYtype <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 53 59 74 79 70 65 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0B 73 57 41 20 53 59 74 79 70 65 20 77 |

Table 505: Example: sWA SYtype

    - Set Phase Shift of Slave for motor synchronization

The phase shift should correspond with the angle of rotation of each slave device ralative to the master

| Telegram structure: sWN SYphas (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set phase shift of slave device | String | 6 | All | SYphas | 53 59 70 68 61 73 |
| Function | Phase value in 1/10000° | Uint_32 | 4 | All | 0... 80: 0... 800000d (0... C3500h) | 00 00 00 00 00... 00 0C 35 00 |

Table 506: Telegram structure: sWN SYphas

Example: sWN SYphas (Slave phase shift = 10°)

| C o La A | ASCII | <STX> sWN {SPC} SYphas {SPC} +100000 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 53 59 70 68 61 73 20 2B 31 30 30 30 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0F 73 57 4E 20 53 59 70 68 61 73 00 01 86 A0 4D |

Table 507: Example: sWN SYphas +100000

| Telegram structure: sWA SYphas |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set phase shift of slave device | String | 6 | All | SYphas | 53 59 70 68 61 73 |

Table 508: Telegram structure: sWA SYphas

Example: sWA SYphas

| CoLa A | ASCII | <STX> sWA {SPC} SYphas <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 53 59 70 68 61 73 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0B 73 57 41 20 53 59 70 68 61 73 20 65 |

Table 509: Example: sWA SYphas

    - Enable/Disable Front Panel

The following tellegram "LMLfpen” corresponds to the SOPAS GUI check box "Display and serial aux interface active” in LMS1xx and "Display active” in LMS5xx.

| Telegram structure: sWN LMLfpen (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set front panel display and serial aux interface | String | 7 | All | LMLfpen | 4C 4D 4C 66 70 65 6E |
| Status | Enable/Disable(Activ e/inactive) | Bool_1 | 1 | All | Enable(Active): 1 Disable(Inactive): 0 | Enable(Active): 01 Disable(Inactive): 00 |

Table 510: Telegram structure: sWN LMLfpen

Example: sWN LMLfpen

| CoLa A | ASCII | <STX> sWN {SPC} LMLfpen {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 4C 66 70 65 6E 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 57 4E 20 4C 4D 4C 66 70 65 6E 20 01 3B |

Table 511: Example: sWN LMLfpen 1 (Enable Front Panel)

| Telegram structure: sWA LMLfpen |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Set front panel display and serial aux interface | String | 8 | All | LMLfpen | 4C 4D 4C 66 70 65 6E |

Table 512: Telegram structure: sWA LMLfpen

Example: sWA LMLfpen

| CoLa A | ASCII | <STX> sWA {SPC} LMLfpFen <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 4D 4C 66 70 65 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 41 20 4C 4D 4C 66 70 65 6E 20 35 |

Table 513: Example: sWA LMLfpen

For LMS1xx Security sensors (LMS12x, LMS13x, LMS14x), telegram "LLMfpmode” corresponds to the selection of "Mode” between "State dependent” (the mode is dependent on the state of the selection of "Arm”/”Disarmed”/”Walk Test”) or "Permanent”(the always active mode).

| Telegram structure: sWN LLMfpmode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | LMS1xx Security | sWN | 73 57 4E |
| Command | Set front panel display and serial aux interface | String | 7 | LMS1xx Security | LLMfpmode | 4C 4C 4D 66 70 6D 6F 64 65 |
| Status | Enable/Disable(Activ e/inactive) | Bool_1 | 1 | LMS1xx Security | Permanent(Active): 0 State dependent(Inactive): 1 | Permanent(Active): 00 State dependent(Inactive): 01 |

Table 514: Telegram structure: sWN LLMfpmode

Example: sWN LLMfpmode

| CoLa A | ASCII | <STX> sWN {SPC} LLMfpmode {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4C 4D 66 70 6D 6F 64 65 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0D 73 57 4E 20 4C 4C 4D 66 70 6D 6F 64 65 20 00 12 |

Table 515: Example: sWN LLMfpmode 0 (set to "Permanent”)

| Telegram structure: sWA LLMfpmode |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | LMS1xx Security | sWA | 73 57 41 |
| Command | Set front panel display and serial aux interface | String | 8 | LMS1xx Security | LLMfpmode | 4C 4C 4D 66 70 6D 6F 64 65 |

Table 516: Telegram structure: sWA LLMfpmode

Example: sWA LLMfpmode

| CoLa A | ASCII | <STX> sWA {SPC} LLMfpmode <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 4C 4D 66 70 6D 6F 64 65 |
| CoLa B | Binary | 02 02 02 00 00 00 0C 73 57 41 20 4C 4D 4C 66 70 65 6E 20 3D |

Table 517: Example: sWA LLMfpmode

    - Set function front panel

| Telegram structure: sWN LMLfpFcn (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set function of the front panel | String | 8 | All | LMLfpFcn | 4C 4D 4C 66 70 46 63 6E |
| Reserved | Reserved | Bool_1 | 1 | All | 1 | 01 |
| LED func- tion Q1/Q2 | Code number | Enum_8 | 1 | All | No function: 0 Application: 1 Command: 2 | No function: 00 Application: 01 Command: 02 |
| LED func- tion OK /Stop | Code number | Enum_8 | 1 | All | Application: 0 Command: 1 | Application: 00 Command: 01 |
| Display function | Code number | Enum_8 | 1 | All | No function: 0 Application: 1 Command: 2 | No function: 00 Application: 01 Command: 02 |

Table 518: Telegram structure: sWN LMLfpFcn

Example: sWN LMLfpFcn

| CoLa A | ASCII | <STX> sWN {SPC} LMLfpFcn {SPC} 1 {SPC} 1 {SPC} 0 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 4C 4D 4C 66 70 46 63 6E 20 31 20 31 20 30 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 11 73 57 4E 20 4 C 4D 4C 66 70 46 63 6E 20 01 01 00 01 7B |

Table 519: Example: sWN LMLfpFcn

| Telegram structure: sWA LMLfpFcn |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Front panel function | String | 8 | All | LMLfpFcn | 4C 4D 4C 66 70 46 63 6E |

Table 520: Telegram structure: sWA LMLfpFcn

Example: sWA LMLfpFcn

| CoLa A | ASCII | <STX> sWA {SPC} LMLfpFcn <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 4C 4D 4C 66 70 46 63 6E 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0D 73 57 41 20 4C 4D 4C 66 70 46 63 6E 75 |

Table 521: Example: sWA LMLfpFcn

    - Set front LEDs

To use this command, it is necessary to set the function of the LED to "Command” (use sWN LMLfpFcn), otherwise this command will have no influence to the LEDs.

OK and Stop LED can only alternate, if one is switched on, the other will turn automatically off.

| Telegram structure: sMN mLMLSetLed |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set front LED | String | 10 | All | mLMLSetLed | 6D 4C 4D 4C 53 65 74 4C 65 64 |
| LED | LED to turn on/off | Int_8 | 1 | All | Stop: 1 OK: 2 Q1: 3 Q2: 4 | Stop: 01 OK: 02 Q1: 03 Q2: 04 |
| Status | On or Off | Int_8 | 1 | All | On: 1 Off: 0 | On: 01 Off: 00 |

Table 522: Telegram structure: sMN mLMLSetLed

Example: sMN mLMLSetLed 1 1 (Stop LED)

| CoLa A | ASCII | <STX> sMN {SPC} mLMLSetLed {SPC} 1 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 4C 53 65 74 4C 65 64 20 31 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 12 73 4D 4E 20 6D 4C 4D 4C 53 65 74 4C 65 64 20 01 20 01 7F |

Table 523: Example: sMN mLMLSetLed 1 1 (Stop LED)

| Telegram structure: sAN mLMLSetLed |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | Front LED | String | 10 | All | mLMLSetLed | 6D 4C 4D 4C 53 65 74 4C 65 64 |
| Status code | Code number | Bool_1 | 1 | All | Error: 0 Success: 1 | Error: 00 Success: 01 |

Table 524: Telegram structure: sAN mLMLSetLed

Example: sAN mLMLSetLed

| CoLa A | ASCII | <STX> sAN {SPC} mLMLSetLed {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 4C 4D 4C 53 65 74 4C 65 64 20 00 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 41 4E 20 6D 4C 4D 4C 53 65 74 4C 65 64 20 00 53 |

Table 525: Example: sAN mLMLSetLed

    - Set function of LED1

With this command the operation of LED1 can be defined.

NAV310 / LE-OEM15xx: Either it has no function, it flashes when output Q1 or application is active or it can be turned on and off by another telegram command (sMN mHMISetLed). LMS4000: Either it has no function, or it flashes when either In/Out1, In/Out2 or Out4 is acktive (selectable).

| Telegram structure: sWN HMIfpFcn_Y1 (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set function of the front panel LED1 | String | 11 | All | HMIfpFcn_Y1 | 48 4D 49 66 70 46 63 6E 5F 59 31 |
| LED1 func- tion Q1 | Code number | Enum_8 | 1 | NAV310, LD- OEM15xx | No function: 0 Application: 1 Command: 2 | No function: 00 Application: 01 Command: 02 |
|  |  |  |  | LMS 4000 | No function: 0 In/Out1: 1 In/Out2: 2 Out4: 4 | 00 01 02 04 |

Table 526: Telegram structure: sWN HMIfpFcn_Y1

Example: sWN HMIfpFcn_Y1 +2 = Command / In/Out2

| CoLa A | ASCII | <STX> sWN {SPC} HMIfpFcn_Y1 {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 48 4D 49 66 70 46 63 6E 5F 59 31 20 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 57 4E 20 48 4D 49 66 70 46 63 6E 5F 59 31 20 02 4E |

Table 527: Example: sWN HMIfpFcn_Y1 2

| Telegram structure: sWA HMIfpFcn_Y1 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | LED1 function | String | 11 | All | HMIfpFcn_Y1 | 48 4D 49 66 70 46 63 6E 5F 59 31 |

Table 528: Telegram structure: sWA HMIfpFcn_Y1

Example: sWA HMIfpFcn_Y1

| CoLa A | ASCII | <STX> sWA {SPC} HMIfpFcn_Y1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 48 4D 49 66 70 46 63 6E 5F 59 31 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 57 41 20 48 4D 49 66 70 46 63 6E 5F 59 31 20 43 |

Table 529: Example: sWA HMIfpFcn_Y1

    - Set function of LED2

With this command the operation of LED2 can be defined. Either it has no function (00), it flashes when output Q2 or application is active (01) or it can be turned on and off (02) by another telegram command (sMN mHMISetLed).

| Telegram structure: sWN HMIfpFcn_Y2 (Authorized client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Set function of the front panel LED2 | String | 11 | All | HMIfpFcn_Y2 | 48 4D 49 66 70 46 63 6E 5F 59 32 |
| LED2 func- tion Q2 | Code number | Enum_8 | 1 | All | No function: 0 Application: 1 Command: 2 | No function: 00 Application: 01 Command: 02 |

Table 530: Telegram structure: sWN HMIfpFcn_Y2

Example: sWN HMIfpFcn_Y2 = Command

| CoLa A | ASCII | <STX> sWN {SPC} HMIfpFcn_Y2 {SPC} 2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 48 4D 49 66 70 46 63 6E 5F 59 32 32 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 57 4E 20 48 4D 49 66 70 46 63 6E 5F 59 32 20 02 7D |

Table 531: Example: sWN HMIfpFcn_Y2 = Command

| Telegram structure: sWA HMIfpFcn_Y2 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | LED2 function | String | 11 | All | HMIfpFcn_Y2 | 48 4D 49 66 70 46 63 6E 5F 59 32 |

Table 532: Telegram structure: sWA HMIfpFcn_Y2

Example: sWA HMIfpFcn_Y2

| CoLa A | ASCII | <STX> sWA {SPC} HMIfpFcn_Y2 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 48 4D 49 66 70 46 63 6E 5F 59 32 |
| CoLa B | Binary | 02 02 02 02 00 00 00 0F 73 57 41 20 48 4D 49 66 70 46 63 6E 5F 59 32 60 |

Table 533: Example: sWA HMIfpFcn_Y2

    - Switch on/off LED1 or LED2

With this command the LEDs can be switched on and off (e.g. to locate the sensor or test the conncection). As a prerequisite, the operation of LED1 and LED2 must have been set to the right function (sWN HMIfpFcn_).

| Telegram structure: sMN mHMISetLed |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Method | String | 3 | All | sMN | 73 4D 4E |
| Command | Set function of the front panel | String | 10 | All | mHMISetLed | 6D 48 4D 49 53 65 74 4C 65 64 |
| LED num- ber 1/2 | LED number | Uint_8 | 1 | All | LED 1: 3 LED 2: 4 | LED 1: 03 LED 2: 04 |
| LED func- tion off/on | Code number | Uint_8 | 1 | All | Off: 0 On: 1 | Off: 00 On: 01 |

Table 534: Telegram structure: sMN mHMISetLed

Example: sMN mHMISetLed 1 = On

| CoLa A | ASCII | <STX> sMN {SPC} mHMISetLed {SPC} 3 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 48 4D 49 53 65 74 4C 65 64 20 33 20 31 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 12 73 4D 4E 20 6D 48 4D 49 53 65 74 4C 65 64 20 03 20 01 7C |

Table 535: Example: sMN mHMISetLed 1 = On

| Telegram structure: sAN mHMISetLed |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sAN | 73 41 4E |
| Command | LED status | String | 10 | All | mHMISetLed | 6D 48 4D 49 53 65 74 4C 65 64 |
| Result | Code number | Bool_1 | 1 | All | No success: 0 Success: 1 | No success: 00 Success: 01 |

Table 536: Telegram structure: sAN mHMISetLed

Example: sAN mHMISetLed 01

| CoLa A | ASCII | <STX> sAN {SPC} mHMISetLed {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 48 4D 49 53 65 74 4C 65 64 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 10 73 41 4E 20 6D 48 4D 49 53 65 74 4C 65 64 20 01 53 |

Table 537: Example: sAN mHMISetLed 01

    - Set 7-segment display to specific symbol or number

Precondition

It is mandatory to define that the 7-segment display should react to SOPAS commands. This option needs to be activated via sWN LMLfpFcn or using the configuration software SOPAS ET. Choose "SOPAS command” in the drop-down list for the 7-segment display (as shown in the figure below). Afterwards the segments of the display can be set via telegram.

The display is consisting of the segments A- P:

The segments A- G are operated using the bits 0 to 6, the segment P using bit 7:

| Segment | p | g | f | e | d | c | B | a |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Related bit | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 0 |

Example: Showing the number "7” on the display:

| Segment | p | g | f | e | d | c | B | a |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| On/Off for showing the required symbol | off | off | off | off | off | on | on | on |
| Binary | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 |

Transfer binary into ASCII: 00000111 bin ≙ 07 ASCII

| Telegram structure: sMN mLMLSetDisp (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Request (SOPAS method by name) | String | 3 | All | sMN | 73 4D 4E |
| Command | Set 7-segment display | String | 11 | All | mLMLSetDisp | 6D 4C 4D 4C 53 65 74 44 69 73 70 |
| Display | 7-segment display in the display | Uint_8 | 1 | All | Display off: 0 ... | 00 ... |
|  | the LMS. |  |  |  | Display shows 2: 5Bh | 5B |
|  | segments A- G are |  |  |  | Display shows 7: 7 | 07 |
|  | operated using |  |  |  | ... | ... |
|  | bits 0 to 6, |  |  |  | Display fully on (8.): FFh | FF |
|  | point P using bit 7. |  |  |  |  |  |

Table 538: Telegram structure: sMN mLMLSetDisp

Example: sMN mLMLSetDisp 07 (Showing the number "7” on the display)

| CoLa A | ASCII | <STX> sMN {SPC} mLMLSetDisp {SPC} 07 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 4D 4C 53 65 74 44 69 73 70 20 30 37 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 4D 4E 20 6D 4C 4D 4C 53 65 74 44 69 73 70 20 07 3B |

Table 539: Example: sMN mLMLSetDisp 07 (Showing the number "7” on the display)

| Telegram structure: sAN mLMLSetDisp |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer (SOPAS answer) | String | 3 | All | sAN | 73 41 4E |
| Command | Set 7-segment display | String | 11 | All | mLMLSetDisp | 6D 4C 4D 4C 53 65 74 44 69 73 70 |
| ErrorCode | The command has been accepted if the error code 1 is returned. | Enum8 | 1 | All | - error - no error | - error - no error |

Table 540: Telegram structure: sAN mLMLSetDisp

Example: sAN mLMLSetDisp 1

| CoLa A | ASCII | <STX> sAN {SPC} mLMLSetDisp {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 4C 4D 4C 53 65 74 44 69 73 70 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 11 73 41 4E 20 6D 4C 4D 4C 53 65 74 44 69 73 70 20 01 31 |

Table 541: Example: sAN mLMLSetDisp 1

  - Application

|  |  | LMS1xx | LMS5xx | TiM2xx | TiM5xx | TIM7xx | NAV310 | LD-OEM15xx | LD-LRS36xx | MRS1000 | LMS1000 | MRS6000 | LMS4000 | LRS4000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4.10.1 | Request status change of monitoring fields on event | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.10.2 | Individual request of monitoring fields to their status changes - ECR xy | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.10.3 | Request SOPAS field data structure | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.10.4 | Request perpendicular distance once | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.10.5 | Request perpendicular distance continiously on event | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.10.6 | Request latest field infringement info | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.10.7 | Request field infringement info continiously on event | x | x |  |  |  |  |  |  |  |  |  |  |  |
| 4.10.8 | Fieldset selection method |  |  |  |  | x |  |  |  |  |  |  |  |  |
| 4.10.9 | Read active fieldset |  |  |  |  | x |  |  |  |  |  |  |  |  |
| 4.10.10 | Write active fieldset |  |  |  |  | x |  |  |  |  |  |  |  |  |

    - Request status change of monitoring fields on event

Precondition

Necessary sensor setup:

- Setup detection fields

- Setup evaluation cases and assign outputs to the evaluation cases Advantage of this telegram:

Status of the evaluation case (1=field free, 2=detecting or 3=infringement) is transmitted if there is a status change of the monitoring field within the application.

For example if the object size in the field exceeds the set object size parameter but the duration of the object inside the field is below the set time parameter. Then the status change from "field free” to "detecting” will be given out via telegram.

Necessary procedure after each power up of the sensor:

- Establish Ethernet connection to LMS

- Activate Ethernet output

| Telegram structure: sEN ECRChangeArr (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sEN | 73 45 4E |
| Command | Data telegram | String | 12 | All | ECRChangeArr | 45 43 52 43 68 61 6E 67 65 41 72 72 |
| Reporting | Start/stop | Enum_8 | 1 | LM1xx LMS5XX | Stop: 0 Start: 1 | 00 01 |

Table 542: Telegram structure: sEN ECRChangeArr

Example: sEN ECRChangeArr 1

| CoLa A | ASCII | <STX> sEN {SPC} ECRChangeArr {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 4E 20 45 43 52 43 68 61 6E 67 65 41 72 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 45 4E 20 45 43 52 43 68 61 6E 67 65 41 72 72 20 01 4A |

Table 543: Example: sEN ECRChangeArr 1

| Telegram structure: sEA ECRChangeArr (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sEA | 73 45 41 |
| Command | Data telegram | String | 12 | All | ECRChangeArr | 45 43 52 43 68 61 6E 67 65 41 72 72 |
| Reporting | Start/stop | Enum_8 | 1 | LM1xx LMS5XX | Stop: 0 Start: 1 | 00 01 |

Table 544: Telegram structure: sEA ECRChangeArr

Example: sEA ECRChangeArr 1

| CoLa A | ASCII | <STX> sEA {SPC} ECRChangeArr {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 41 20 45 43 52 43 68 61 6E 67 65 41 72 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 12 73 45 41 20 45 43 52 43 68 61 6E 67 65 41 72 72 20 01 45 |

Table 545: Example: sEA ECRChangeArr 1

NOTE

The answer to the telegram will be followed by data that is sent on event.

The sensor only sends the following answer if there is a status change of the evaluation case within the application.

| Telegram structure: sSN ECRChangeArr (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sSN | 73 53 4E |
| Command | Data telegram | String | 12 | All | ECRChangeArr | 45 43 52 43 68 61 6E 67 65 41 72 72 |
| System counter | Time in &mu;s since power up max. | Uint_32 | 4 | LM1xx LMS5XX | 0... FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |

| Telegram structure: sSN ECRChangeArr (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | 71min then starting from 0 again |  |  |  |  |  |
| Array | 1-10 | UInt_16 | 2 | LM1xx LMS5XX | 0h... Ah (0d... 10d) | 00 01... 00 0A |
| EVC number | 1-10 | UInt_8 | 1 | LM1xx LMS5XX | 0h... Ah (0d... 10d) | 00 01... 00 0A |
| Object detection |  | Enum_8 | 1 | LM1xx LMS5XX | 1= field free 2= detecting 3= infrigement | LMS1xx: 01 = field free 02 = detecting 03 = infrigement LMS5xx: 31= field free 32= detecting 33= infrigement |
| Year |  | Uint_16 | 2 | LM1xx LMS5XX | 0000h... 270Fh (0d... 9999d) | 00 00... 27 0F |
| Month | 1 to 12 | Uint_8 | 1 | LM1xx LMS5XX | 00h... 0Ch (0d... 12d) | 00... 0C |
| Day | Day of month 1 to 31 | Uint_8 | 1 | LM1xx LMS5XX | 00h... 1Fh (0d... 31d) | 00... 1F |
| Hour | 0 to 23 | Uint_8 | 1 | LM1xx LMS5XX | 00h... 17h (0d... 23d) | 00... 17 |
| Minute | 0 to 59 | Uint_8 | 1 | LM1xx LMS5XX | 00h... 3Bh 0d... 59d | 00... 3B |
| Second | 0 to 59 | Uint_8 | 1 | LM1xx LMS5XX | 00h... 3Bh (0d... 59d) | 00... 3B |
| µSecond | 0 to 9999999 | Uint_32 | 4 | LM1xx LMS5XX | 00000000h&hellip; 000F423Fh (0d... 999999d) | 00 00 00 00 00... 00 0F 42 3F |

Table 546: Telegram structure: sSN ECRChangeArr

Example: sSN ECRChangeArr 1

| CoLa A | ASCII | <STX> sSN {SPC} ECRChangeArr {SPC} 97F8C2E3 {SPC} 1 {SPC} 1 {SPC} 1 {SPC} 7B2 {SPC} 1 {SPC} 1 {SPC} 0 {SPC} 2A {SPC } 1D {SPC} 63DA8 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 53 4E 20 45 43 52 43 68 61 6E 67 65 41 72 72 20 39 37 46 38 43 32 45 33 20 31 20 31 20 31 20 37 42 32 20 31 20 31 20 30 20 32 41 20 31 44 20 36 33 44 41 38 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 24 73 53 4E 20 45 43 52 43 68 61 6E 67 65 41 72 72 20 CF 09 10 99 00 01 01 02 07B2 01 01 02 09 1C 00 03 47 D8 2E |

Table 547: Example: sSN ECRChangeArr 1

    - Individual request of monitoring fields to their status changes - ECR xy

NOTE

The telegram "ECRxy” is available with the LMS12x, LMS13x, LMS14x from firmware version V1.32 The telegram "ECRxy” available with the LMS531 from firmware version V2.10

Precondition:

Necessary sensor setup: 1.Setup detection fields 2.Setup evaluation cases.

- Setup output assignments to the evaluation cases

Necessary proceeding after each Laser Scanner power up:

- Establish Ethernet connection to LMS

- If the user wants to register an event, it is necessary to send a telegram to register it each time after establishing connection

Read/Poll the status of an evaluation case

It is possible to send (poll repeatedly) this telegram to Ethernet port 2111 or Ethernet port 2112 to check the evaluation status.

LMS 531

| Telegram structure: sRN ECRxy (no required User Level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Data telegram | String | 12 | LMS1xx | ECRxy: | ECR01: 45 43 52 30 31 |
|  |  |  |  |  | (concrete number should be | ECR02: 45 43 52 30 32 |
|  |  |  |  |  | used for xy as shown below) | ... |
|  |  |  |  |  | ECR01(evaluation1) ECR02(evaluation2) | ECR 10: 45 43 52 31 30 |
|  |  |  |  |  | ECR10(evaluation10) |  |

Table 548: Telegram structure: sRN ECRxy

Example: sRN ECR01

| CoLa A | ASCII | <STX> sRN {SPC} ECR01 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 45 43 52 30 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 09 73 52 4E 20 45 43 52 30 31 1A |

Table 549: Example: sRN ECR01

The response telegram is shown in Table 7: Telegram structure: sRA/sSN ECRxy

Register an event to get the status update of an evaluation

It is also possible to register an event by "sEN ECRxy” for receiving the change of the an evaluation. So, when there a change happens in the field of this evaluation, an update telegram will be sent out automatically.

The registration telegram format is: sEN ECRxy 0/1 (shown below)

LMS 531

| Telegram structure: sEN ECRxy (no required User Level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sEN | 73 45 4E |
| Command | Data telegram | String | 12 | LMS1xx | ECRxy: | ECR01: 45 43 52 30 31 |
|  |  |  |  |  | (concrete number should be | ECR02: 45 43 52 30 32 |
|  |  |  |  |  | used for xy as shown below) | ... |
|  |  |  |  |  | ECR01(evaluation1) ECR02(evaluation2) | ECR 10: 45 43 52 31 30 |
|  |  |  |  |  | ECR10(evaluation10) |  |
| Reporting | Start/stop | Enum_8 | 1 | LMS1XX | Stop: 0 Start: 1 | Stop: 00 Start: 01 |

Table 550: Telegram structure: sEN ECRxy

Example: sEN ECR01 1

| CoLa A | ASCII | <STX> sEN {SPC} ECR01 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 4E 20 45 43 52 30 31 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0B 73 45 4E 20 45 43 52 30 31 20 01 2C |

Table 551: Example: sEN ECR01 1

LMS 531

| Telegram structure: sEA ECRxy (no required User Level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sEA | 73 45 41 |
| Command | Data telegram | String | 12 | LMS1xx | ECRxy: | ECR01: 45 43 52 30 31 |
|  |  |  |  |  | (concrete number should be | ECR02: 45 43 52 30 32 |
|  |  |  |  |  | used for xy as shown below) | ... |
|  |  |  |  |  | ECR01(evaluation1) ECR02(evaluation2) | ECR 10: 45 43 52 31 30 |
|  |  |  |  |  | ECR10(evaluation10) |  |
| Reporting | Start/stop | Enum_8 | 1 | LMS1XX | Stop: 0 Start: 1 | Stop: 00 Start: 01 |

Table 552: Telegram structure: sEA ECRxy

Example: sEA ECR01 1

| CoLa A | ASCII | <STX> sEA {SPC} ECR01 {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 41 20 45 43 52 30 31 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 0B 73 45 41 20 45 43 52 30 31 20 01 23 |

Table 553: Example: sEA ECR01 1

Notes.

  - The event registration is not a parameter and cannot be saved permanently.

Therefor it is necessary to send the telegram to register an event each time after establishing a connection to the device (especially, when the Ethernet port is in server mode).

  - Ethernet port 2112 has a special function and can be set as client mode (as shown in Figure 1). Then, the corresponding telegram can be selected for the expected evaluation (for example evaluation 1, as shown below). There is no further need to manually register an event by sending a separate telegram. When there is a change in the registered field evaluation, LMS12x, LMS13x, LMS14x will automatically connect to the specified server (such as 192.168.0.1, as shown below) to send an update telegram.If the settings are saved permanently, this is the same situation even after the sensor power cycle.

Figure 1: Ethernet host port setting (LMS12x,LMS13x, LMS14x)

The response telegram to "sRN ECRxy” and the event status update telegram for "sEA ECRxy 1” follows the format in table 7 as shown below.

LMS 531

| Telegram structure: sRA/sSN ECRxy (no required User Level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRA: for sRN ECRxy sSN: for sEN ECRxy | 73 52 41 73 53 4E |
| Command | Data telegram | String | 12 | LMS1xx | ECRxy: | ECR01: 45 43 52 30 31 |
|  |  |  |  |  | (concrete number should be | ECR02: 45 43 52 30 32 |
|  |  |  |  |  | used for xy as shown below) | ... |
|  |  |  |  |  | ECR01(evaluation1) ECR02(evaluation2) | ECR 10: 45 43 52 31 30 |
|  |  |  |  |  | ECR10(evaluation10) |  |
| Reporting | Object detection | Enum_8 | 1 | LMS1XX | 0 =DON'T CARE | 00=DON'T CARE |
|  |  |  |  |  | 1 =FIELD FREE | 01=FIELD FREE |
|  |  |  |  |  | 2 =DETECTING | 02=DETECTING |
|  |  |  |  |  | 3 =INFRINGEMENT | 03=INFRIGEMENT |
| Timestamp | Transmisssion time stamp of the current telegram (unit:µs) | Uint_32 | 4 | LMS1XX | 00000000h&hellip;. FFFFFFFFh | 00 00 00 00 00... FF FF FF FF |
| Year |  | Uint_16 | 2 | LMS1XX | 0000h... 270Fh (0d... 9999d) | 00 00... 27 0F |
| Month | 1 to 12 | Uint_8 | 1 | LMS1XX | 00h... 0Ch (0d... 12d) | 00... 0C |
| Day | Day of month 1 to 31 | Uint_8 | 1 | LMS1XX | 00h... 1Fh (0d... 31d) | 00... 1F |
| Hour | 0 to 23 | Uint_8 | 1 | LMS1XX | 00h... 17h (0d... 23d) | 00... 17 |
| Minute | 0 to 59 | Uint_8 | 1 | LMS1XX | 00h... 3Bh 0d... 59d | 00... 3B |
| Second | 0 to 59 | Uint_8 | 1 | LMS1XX | 00h... 3Bh (0d... 59d) | 00... 3B |
| µSecond | 0 to 9999999 | Uint_32 | 4 | LMS1XX | 00000000h&hellip; 000F423Fh (0d... 999999d) | 00 00 00 00 00... 00 0F 42 3F |

Table 554: Telegram structure: sRA/sSN ECRxy

Example: sRA ECR01

| CoLa A | ASCII | <STX> sRA {SPC} ECR01 {SPC} 1 {SPC} 1B14A215 7B2 1 1 0 7 22 4CE78 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 45 43 52 30 31 20 31 20 31 42 31 34 41 32 31 35 20 37 42 32 20 31 20 31 20 30 20 37 20 32 32 20 34 43 45 37 38 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 29 73 52 41 20 45 43 52 30 31 20 01 1B 14 A2 15 07 B2 01 01 00 07 22 04 CE 78AE |

Table 555: Example: sRA ECR01

    - Request SOPAS field data structure

Note

The SOPAS telegram "mLFEgetField” requests a field number as parameter and returns the corresponding SOPAS field data structure. If the field number is number is not configured, the answer telegram will be filled with 0.

Example - Request the field data structure of an evaluation field (field number: 1) that has been parameterized in the engineering tool SOPAS.

The answer telegram will include information regarding the field data structure of the requested evaluation field (type: segmented, number of field points etc.)

Please note: The sensor will switch to the state "Stop measurement” during read out. After the read out you have to switch the sensor back to "Run measurement ”.

| Telegram structure: sMN mLFEgetField (required user level: Authorized Client) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Request | String | 3 | LMS1xx LMS5XX | sMN | 73 4D 4E |
| Command | Only one telegram | String | 14 | LMS1xx LMS5XX | mLFEgetField | 6D 4C 46 45 67 65 74 46 69 65 6C 64 |
| Command |  | UInt_32 | 4 | LMS1xx LMS5XX | 1... 10 | 00 00 00 01... 00 00 01 01 |

Table 556: Telegram structure: sMN mLFEgetField

Example: sMN mLFEgetField 1

| CoLa A | ASCII | <STX> sMN {SPC} mLFEgetField {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 4D 4E 20 6D 4C 46 45 67 65 74 46 69 65 6C 64 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 4D 4E 20 6D 4C 46 45 67 65 74 46 69 65 6C 64 20 00 00 00 01 67 |

Table 557: Example: sMN mLFEgetField 1

| Telegram structure: sAN mLFEgetField |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Read |  |  | String | 3 | All | sAN | 73 41 4E |
| Data telegram |  |  | String | 12 | All | mLFEgetField | 6D 4C 46 45 67 65 74 46 69 65 6C 64 |
| Field header | Distance Scale Factor | Scale factor or factor of the measurement values (for the LMS5xx this depends on angular resolution) | Real as float accordin g to IEEE754 | 4 | All | Factor × 1: 3F800000h Factor × 2: 40000000h | 3F 80 00 00 40 00 00 00 00 |
|  | Distance Scale Offset | Sets starting point of measurement | Real as float accordin g to IEEE754 | 4 | All | No offset: 0 | 00 00 00 00 00 |
|  | Angle Scale Factor | Angle resolution | UInt_32 | 4 | LMS1xx | 50Hz: 1388h | 00 00 13 88 |
|  |  |  |  |  | LMS5xx | 25Hz: 9C4h 35Hz: DACh 50Hz: 1388h 75Hz: 1D4Ch 100Hz:2710h | 00 00 09 C4 00 00 0D AC 00 00 13 88 00 00 1D 4C 00 00 27 10 |
|  | Angle Scale Offset | Starting angle | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  | Field type |  | Enum_8 | 1 | All | 0 = Radial 1 = Rectangle 2 = Segmented 3 = Dynamic | 00 01 02 03 |
|  | Field number |  | UInt_8 | 1 | All | 0 .. Ah | 00... 0A |

| Segmented field | Segment ed field configure d |  | UInt_16 | 2 | All | 0 = no data for segmented field available, I .e. field type is not segmented field 1 = data for segmented field available | 00 00... 00 01 |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Only shown if field type is configure d | Number of field points | UInt_16 | 2 | All | 0 ... 571: 0h... 23Bh | 00 00... 02 3B |
|  |  | Angle index n | UInt_16 | 2 | LMS1xx | 0 ... 1081: 0h... 0439h | 00 00... 04 39 |
|  |  |  |  |  | LMS5xx | 0... 1140: 0h... 0474h | 00 00... 04 74 |
|  |  | Start distance | UInt_16 | 2 | All | 0 .. 65535: 0h... FFFFh | 00 00... FF FF |
|  |  | End distance | UInt_16 | 2 | All | 0 .. 65535: 0h... FFFFh | 00 00... FF FF |
|  |  | Angle index n+1 | ... | ... | ... | ... | ... |
|  |  | Start distance | ... | ... | ... | ... | ... |
|  |  | End distance | ... | ... | ... | ... | ... |
|  |  | ... | ... | ... | ... | ... | ... |
| Rectangular field | Rectangu lar field configure d |  | UInt_16 | 2 | All | 0 = no data for rectangular field available, I .e. field type is not rectangular field 1 = data for rectangular field available | 00 00... 00 01 |
|  | Only shown if field type is configure d | Angle of reference point | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  | Distance of reference point | UInt_16 | 2 | All | 0 .. 65535: 0h... FFFFh | 00 00... FF FF |
|  |  | Rotation angle | Int_32 | 4 | All | -1800000d .. 1800000d (FFE488C0h... 001 B7740h) | FF E4 88 C0&hellip; 00 1B 77 40 |
|  |  | Width | UInt_32 | 4 | All | 0... 4294967295d (0 &hellip; FFFFFFFFh) | 00 00 00 00 00... FF FF FF FF |
|  |  | Length | UInt_32 | 4 | All | 0... 4294967295d (0 &hellip; FFFFFFFFh) | 00 00 00 00 00... FF FF FF FF |
| Radial field |  | Radial field configured | UInt_16 | 2 | All | 0 (radial fields are not available) | 00 00 |
| Dynamic field | Dynamic field configure d |  | UInt_16 | 2 | All | 0 = no data for dynamic field available, I .e. field type is not dynamic field 1 = data for dynamic field available | 00 00... 00 01 |
|  | Only shown if field type is configure d | Angle of reference point | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  |  | Distance of reference point | UInt_16 | 2 | All | 0 .. 65535: 0h... FFFFh | 00 00... FF FF |
|  |  | Rotation angle | Int_32 | 4 | All | -1800000d .. 1800000d (FFE488C0h... 001 B7740h) | FF E4 88 C0&hellip; 00 1B 77 40 |

|  |  |  | UInt_32 | 4 | All | 0... 4294967295d (0 &hellip; FFFFFFFFh) | 00 00 00 00 00... FF FF FF FF |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | UInt_32 | 4 | All | 0... 4294967295d (0 &hellip; FFFFFFFFh) | 00 00 00 00 00... FF FF FF FF |
|  |  | Maximum speed | Int_16 | 2 | All | 0... 30000d (0... 7530h) | 00 00... 75 30 |
|  |  | Maximum length | UInt_32 | 4 | All | 0 .. 4294967295 | 00 00 00 00 00... 7F FF FF FF |
| Version number |  |  | UInt_16 | 2 | All | 0 .. 65535 | 00 00... FF FF |
| Length of field name |  |  | UInt_16 | 2 | All | 0 .. 32 | 00 00 .. 00 20 |
| Field name |  |  | String | 0... 32 | All | FIELD1 | 46 49 45 4c 44 31 |
| Length of comment |  | If no comment, also no length with a string type... | UInt_16 | 2 | All | 0... 128d 0... 80h) | 00 00... 00 80 |
| Comment |  |  | String | 0... 128 | All |  |  |

Table 558: Example: sAN mLFEgetField

Example: sAN mLFEgetField

| CoLa A | ASCII | <STX> sAN {SPC} mLFEgetField {SPC} 40000000 {SPC} 00000000 {SPC} 1388 {SPC} FFFF3CB0 {SPC} 2 {SPC} 1 {SPC} 1 {SPC} 4 {SPC} 12 {SPC} FFFF {SPC} F7 {SPC} 72 {SPC} FFFF {SPC} 18B {SPC} 12D {SPC} FFFF {SPC} 233 {SPC} 16D {SPC} FFF F {SPC} 1DC {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 1 {SPC} 6 {SPC} FIELD1 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 41 4E 20 6D 4C 46 45 67 65 74 46 69 65 6C 64 20 34 30 30 30 30 30 30 30 20 30 30 30 30 30 30 30 30 20 31 33 38 38 20 46 46 46 46 33 43 42 30 20 32 20 31 20 31 20 34 20 31 32 20 46 46 46 46 20 46 37 20 37 32 20 46 46 46 46 20 31 38 42 20 31 32 44 20 46 46 46 46 20 32 33 33 20 31 36 44 20 46 46 46 46 20 31 44 43 20 30 20 30 20 30 20 31 20 36 20 46 49 45 4C 44 31 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 51 73 41 4E 20 6D 4C 46 45 67 65 74 46 69 65 6C 64 20 40 00 00 00 00 00 00 00 00 00 13 88 FF FF 3C B0 02 01 00 01 00 04 00 12 FF FF 00 F7 00 72 FF FF 01 8B 01 2D FF FF 02 33 01 6D FF FF 01 DC 00 00 00 00 00 00 00 01 00 06 46 49 45 4C 44 31 00 00 FE |

Table 559: Example: sAN mLFEgetField

    - Request perpendicular distance once

Precondition

The evaluation stategy "Perpendicular distance” has to be activated in the engineering tool SOPAS. After you get the perpendicular distance displayed in SOPAS as shown below, you are able to read out the telegram.

Set the EVC to evaluation strategy "Perpendicular distance” and choose the field which you like to read out. Also keep in mind, that the field evalutation will only be activated by choosing an output.

Fehler!Es ist nicht möglich, durch die Bearbeitung von Feldfunktionen Objekte zu erstellen.

| Telegram structure: sRN LFEperpdistresult (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Only one telegram | String | 14 | All | LFEperpdistresult | 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 |

Table 560: Telegram structure: sRN LFEperpdistresult

Example: sRN LFEperpdistresult

| CoLa A | ASCII | <STX> sRN {SPC} LFEperpdistresult <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 52 4E 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 14 |

Table 561: Example: sRN LFEperpdistresult

| Telegram structure: sRA LFEperpdistresult (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRA | 73 52 41 |
| Command | Data telegram | String | 14 | All | LFEperpdistresult | 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 |
| Array | 1-10 | UInt_16 | 2 | All | 0... 10d (0 &hellip; Ah) | 00 00... 00 0A |
| EVC number | 1-10 | UInt_8 | 1 | All | 0... 10d (0 &hellip; Ah) | 00... 0A |
| Perpendicu lar distance | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |
| Reserved | default 0 | UInt_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | UInt_32 | 4 | All | 0 | 00 00 00 00 00 |
| X-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
| Y-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
| Reserved | default 0 | Int_32 | 4 | LMS1xx LMS5XX | 0 | 00 00 00 00 00 |
| Radial distance | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |
| Radial distance | LMS1xx: -45 °... 225 ° | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |

| Telegram structure: sRA LFEperpdistresult (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| angle | LMS5xx: -5 °... 185 ° |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reference line angle | LMS1xx: -45... 225 LMS5xx: -5 °... 185 ° | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |

Table 562: Telegram structure: sRA LFEperpdistresult

Example: sRA LFEperpdistresult

| CoLa A | ASCII | <STX> sRA {SPC} LFEperpdistresult {SPC} 1 {SPC} 2 {SPC} 21A {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 21A {SPC} 0 {SPC} 21A {SPC} D BBA0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 20 31 20 32 20 32 31 41 20 30 20 30 20 30 20 32 31 41 20 30 20 32 31 41 20 44 42 42 41 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 61 73 52 41 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 00 01 20 02 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 38 |

Table 563: Example: sRA LFEperpdistresult

    - Request perpendicular distance continiously on event

Precondition

The evaluation stategy "Perpendicular distance” has to be activated in the engineering tool SOPAS. After you get the perpendicular distance displayed in SOPAS as shown below, you are able to read out the telegram.

Set the EVC to evaluation strategy "Perpendicular distance” and choose the field which you like to read out. Also keep in mind, that the field evalutation will only be activated by choosing an output.

Fehler!Es ist nicht möglich, durch die Bearbeitung von Feldfunktionen Objekte zu erstellen.

| Telegram structure: sEN LFEperpdistresult (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sEN | 73 45 4E |
| Command | Only one telegram | String | 14 | All | LFEperpdistresult | 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 |
| Reporting | Start/stop | Enum_8 | 1 | LMS1xx LMS5XX | Stop: 0 Start: 1 | Stop: 00 Start: 01 |

Table 564: Telegram structure: sEN LFEperpdistresult

Example: sEN LFEperpdistresult 1

| CoLa A | ASCII | <STX> sEN {SPC} LFEperpdistresult {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 4E 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 45 4E 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 20 01 22 |

Table 565: Example: sEN LFEperpdistresult 1

| Telegram structure: sEA LFEperpdistresult (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sEA | 73 45 41 |
| Command | Only one telegram | String | 14 | All | LFEperpdistresult | 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 |
| Reporting | Start/stop | Enum_8 | 1 | LMS1xx LMS5XX | Stop: 0 Start: 1 | Stop: 00 Start: 01 |

Table 566: Telegram structure: sEA LFEperpdistresult

Example: sEA LFEperpdistresult 1

| CoLa A | ASCII | <STX> sEA {SPC} LFEperpdistresult {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 41 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 17 73 45 41 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 20 01 2D |

Table 567: Example: sEA LFEperpdistresult 1

NOTE

The answer to the telegram will be followed by data that is sent on event.

The sensor only sends the following answer if there are perpendicular distance values calculated within the application.

| Telegram structure: sSN LFEperpdistresult (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sSN | 73 53 4E |
| Command | Data telegram | String | 14 | All | LFEperpdistresult | 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 |
| Array | 1-10 | UInt_16 | 2 | All | 0... 10d (0 &hellip; Ah) | 00 00... 00 0A |
| EVC number | 1-10 | UInt_8 | 1 | All | 0... 10d (0 &hellip; Ah) | 00... 0A |
| Perpendicu lar distance | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |
| Reserved | default 0 | UInt_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | UInt_32 | 4 | All | 0 | 00 00 00 00 00 |
| X-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
| Y-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Radial distance | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |
| Radial distance angle | LMS1xx: -45 °... 225 ° LMS5xx: | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d | FF FF 3C B0&hellip; 00 1C 3A 90 |

| Telegram structure: sSN LFEperpdistresult (no required user level) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | -5 °... 185 ° |  |  |  | (FFFF3CB0h... 1 C3A90h) |  |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reference line angle | LMS1xx: -45... 225 LMS5xx: -5 °... 185 ° | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |
| Reserved | default 0 | Int_32 | 4 | All | 0 | 00 00 00 00 00 |

Table 568: Telegram structure: sSN LFEperpdistresult

Example: sSN LFEperpdistresult

| CoLa A | ASCII | <STX> sSN {SPC} LFEperpdistresult {SPC} 1 {SPC} 2 {SPC} B0 {SPC} 0 {SPC} 0 {SPC} FFFFFFE8 {SPC} B0 {SPC} 0 {SPC} B2 {S PC} EF420 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 {SPC} 0 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 53 4E 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 20 31 20 32 20 42 30 20 30 20 30 20 46 46 46 46 46 46 45 38 20 42 30 20 30 20 42 32 20 45 46 34 32 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 20 30 03 |
| CoLa B | Binary | 02 02 02 00 00 00 61 73 53 4E 20 4C 46 45 70 65 72 70 64 69 73 74 72 65 73 75 6C 74 20 00 01 01 00 00 01 5E 00 00 00 00 00 00 00 00 00 00 00 4A 00 00 01 5E 00 00 00 00 00 00 00 01 66 00 0B E6 E0 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 00 15 |

Table 569: Example: sSN LFEperpdistresult

    - Request latest field infringement info

Note

The command is used to request entries from the "field infringement logging” via telegram. Using this command, only the latest field infringement is given out. The answer telegram includes information regarding date and time of the infringement as well as the associated EVC and infringement position.

Please note: Only activated EVCs will be considered. An EVC is only valid, when an output is defined in the EVC. Using this telegram, an EVC without a defined output is not taken into considaration.

| Telegram structure: sRN LFEinfringementinfo |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Only one telegram | String | 19 | LMS1xx | LFEinfringementinfo | 4C 46 45 69 6E 66 72 69 |
|  |  |  |  | LMS5XX |  | 6E 67 65 6D 65 6E 74 69 6E 66 6F |

Table 570: Telegram structure: sRN LFEinfringementinfo

Example: sRN LFEinfringementinfo

| CoLa A | ASCII | <STX> sRN {SPC} LFEinfringementinfo <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 03 |
| CoLa B | Binary | 02 02 02 02 00 00 00 73 52 4E 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F A |

Table 571: Example: sEN LFEinfringementinfo

NOTE

The answer telegram refers to the latest recorded field infringement. Therefore the device outputs the latest field infringement entry from the logging table.

| Telegram structure: sRA LFEinfringementinfo |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type |  | Read | String | 3 | All | sRA | 73 52 41 |
| Command |  | Data telegram | String | 19 | All | LFEinfringementinfo | 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F |
| Time Info | Counter | continious counter of infrigements | Uint_16 | 2 | All | 0... 9999d (0... 270Fh) | 00 00... 27 0F |
|  | Year |  | Uint_16 | 2 | All | 0... 9999d (0... 270Fh) | 00 00... 27 0F |
|  | Month | 1 to 12 | Uint_8 | 1 | All | 0... 12d (0 &hellip; Ch) | 00... 0C |
|  | Day | Day of month 1 to 31 | Uint_8 | 1 | All | 0... 31d (0... 1Fh) | 00... 1F |
|  | Hour | 0 to 23 | Uint_8 | 1 | All | 0... 23d (0... 17h) | 00... 17 |
|  | Minute | 0 to 59 | Uint_8 | 1 | All | 0... 59d (0... 3Bh) | 00... 3B |
|  | Second | 0 to 59 | Uint_8 | 1 | All | 0... 59d (0... 3Bh) | 00... 3B |
|  | µSecond | 0 to 999999 | Uint_32 | 4 | All | 0... 999999d (0 &hellip; 000 F423Fh) | 00 00 00 00 00... 00 0F 42 3F |
| Infrigment-Info | EVC number | 1-10 | UInt_8 | 1 | All | 0... 10d (0 &hellip; Ah) | 00... 0A |
|  | X-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
|  | Y-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
|  | Distance [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |

| Telegram structure: sRA LFEinfringementinfo |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  |  | 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |
|  | Angle [°] | LMS1xx: -45... 225 LMS5xx: -5 °... 185 ° | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  | Angle of Layer [°] | LMS1xx: -45... 225 LMS5xx: -5 °... 185 ° | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  | Object Size [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |

Table 572: Example: sRA LFEinfringementinfo

Example: sRA LFEinfringementinfo

| CoLa A | ASCII | <STX> sRA {SPC} LFEinfringementinfo {SPC} 4 {SPC} 7B2 {SPC} 1 {SPC} 1 {SPC} 2 {SPC} 32 {SPC} 1F {SPC} 3A1B0 {SPC} 1 {S PC} FFFFFDAE {SPC} D6 {SPC} 277 {SPC} 187083 {SPC} 4A <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 34 20 37 42 32 20 31 20 31 20 32 20 33 32 20 31 46 20 33 41 31 42 30 20 31 20 46 46 46 46 46 44 41 45 20 44 36 20 32 37 37 20 31 38 37 30 38 33 20 34 41 03 |
| CoLa B | Binary | 02 02 02 00 00 00 3A 73 52 41 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 00 04 07 B2 01 01 02 32 1F 00 03 A1 B0 01 FF FF FD AE 00 00 00 D6 00 00 02 77 00 18 70 83 00 00 00 4A F9 |

Table 573: Example: sRA LFEinfringementinfo

    - Request field infringement info continiously on event

Note

The command is used to request entries from the "field infringement logging” via telegram. Using this command, a telegram will be sent from the sensor on event (in case of a new field infringement). The answer telegram includes information regarding date and time of the infringement as well as the associated EVC and infringement position.

Please note: Only activated EVCs will be considered. An EVC is only valid, when an output is defined in the EVC. Using this telegram, an EVC without a defined output is not taken into considaration.

| Telegram structure: sEN LFEinfringementinfo |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Event | String | 3 | All | sEN | 73 45 4E |
| Command | Data telegram | String | 19 | All | LFEinfringementinfo | 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F |
| Reporting | Start/stop | Enum_8 | 1 | All | Stop: 0 | 00 |
|  |  |  |  |  | Start: 1 | 01 |

Table 574: Telegram structure: sEN LFEinfringementinfo

Example: sEN LFEinfringementinfo 1

| CoLa A | ASCII | <STX> sEN {SPC} LFEinfringementinfo {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 4E 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 19 73 45 4E 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 01 3C |

Table 575: Example: sEN LFEinfringementinfo 1

Info

| Telegram structure: sEA LFEinfringementinfo |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Event | String | 3 | All | sEA | 73 45 41 |
| Command | Data telegram | String | 19 | All | LFEinfringementinfo | 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F |
| Reporting | Start/stop | Enum_8 | 1 | All | Stop: 0 | 00 |
|  |  |  |  |  | Start: 1 | 01 |

Table 576: Telegram structure: sEA LFEinfringementinfo

Example: sEA LFEinfringementinfo 1

| CoLa A | ASCII | <STX> sEA {SPC} LFEinfringementinfo {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 45 41 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 19 73 45 41 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 01 33 |

Table 577: Example: sEA LFEinfringementinfo 1

NOTE

The answer to the telegram is followed by data that is sent on event.

The sensor only sends the following data if there is a new field infringement detected and noted in the field infringement logging.

| Telegram structure: sSN LFEinfringementinfo |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type |  | Read | String | 3 | All | sSN | 73 53 4E |
| Command |  | Data telegram | String | 19 | All | LFEinfringementinfo | 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F |
| Time | Counter | continious counter of infrigements | Uint_16 | 2 | All | 0... 9999d (0... 270Fh) | 00 00... 27 0F |

| Telegram structure: sSN LFEinfringementinfo |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | Year |  | Uint_16 | 2 | All | 0... 9999d (0... 270Fh) | 00 00... 27 0F |
|  | Month | 1 to 12 | Uint_8 | 1 | All | 0... 12d (0 &hellip; Ch) | 00... 0C |
|  | Day | Day of month 1 to 31 | Uint_8 | 1 | All | 0... 31d (0... 1Fh) | 00... 1F |
|  | Hour | 0 to 23 | Uint_8 | 1 | All | 0... 23d (0... 17h) | 00... 17 |
|  | Minute | 0 to 59 | Uint_8 | 1 | All | 0... 59d (0... 3Bh) | 00... 3B |
|  | Second | 0 to 59 | Uint_8 | 1 | All | 0... 59d (0... 3Bh) | 00... 3B |
|  | µSecond | 0 to 999999 | Uint_32 | 4 | All | 0... 999999d (0 &hellip; 000 F423Fh) | 00 00 00 00 00... 00 0F 42 3F |
| Infrigment-Info | EVC number | 1-10 | UInt_8 | 1 | All | 0... 10d (0 &hellip; Ah) | 00... 0A |
|  | X-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
|  | Y-Pos [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | Int_32 | 4 | LMS100, LMS111 | -20000d ... 20000d (FFFFB1E0h... 4 E20h) | FF FF B1 E0&hellip; 00 00 4E 20 |
|  |  |  |  |  | LMS151 | -50000d... 50000d (FFFF3CB0h&hellip;C350h) | FF FF 3C B0&hellip; 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | -80000d... 80000d (FFFEC780h &hellip; 13880h) | FF FE C7 80&hellip; 00 01 38 80 |
|  | Distance [mm] | LMS100, LMS111: 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |
|  |  |  |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |
|  | Angle [°] | LMS1xx: -45... 225 LMS5xx: -5 °... 185 ° | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  | Angle of Layer [°] | LMS1xx: -45... 225 LMS5xx: -5 °... 185 ° | Int_32 | 4 | LMS1xx | -450000d... +2250000d (FFF92230h &hellip; 225510h) | FF F9 22 30&hellip; 00 22 55 10 |
|  |  |  |  |  | LMS5xx | -50000d... + 1850000d (FFFF3CB0h... 1 C3A90h) | FF FF 3C B0&hellip; 00 1C 3A 90 |
|  | Object Size | LMS100, LMS111: | UInt_32 | 4 | LMS100, LMS111 | 0... 20000d (0... 4 E20h) | 00 00 00 00 00... 00 00 4E 20 |

| Telegram structure: sSN LFEinfringementinfo |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Telegram part |  | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
|  | [mm] | 0 m&hellip; 20 m LMS151: 0 m&hellip; 50 m LMS5xx: 0 m&hellip; 80 m |  |  | LMS151 | 0... 50000d (0... C350h) | 00 00 00 00 00... 00 00 C3 50 |
|  |  |  |  |  | LMS5xx | 0... 80000d (0 &hellip; 13880h) | 00 00 00 00 00... 00 01 38 80 |

Table 578: Example: sSN LFEinfringementinfo

Example: sSN LFEinfringementinfo

| CoLa A | ASCII | <STX> sSN {SPC} LFEinfringementinfo {SPC} 6 {SPC} 7B2 {SPC} 1 {SPC} 1 {SPC} 0 {SPC} 19 {SPC} 21 {SPC} 9D788 {SPC} 1 { SPC} 41 {SPC} C3 {SPC} 0 {SPC} CE {SPC} AE8F8 {SPC} 0 {SPC} A8 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 53 4E 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 36 20 37 42 32 20 31 20 31 20 30 20 31 39 20 32 31 20 39 44 37 38 38 20 31 20 34 31 20 43 33 20 30 20 43 45 20 41 45 38 46 38 20 30 20 41 38 03 |
| CoLa B | Binary | 02 02 02 00 00 00 40 73 53 4E 20 4C 46 45 69 6E 66 72 69 6E 67 65 6D 65 6E 74 69 6E 66 6F 20 00 06 07 B2 01 01 00 19 21 00 09 D7 88 01 00 00 00 41 00 00 00 C3 00 00 00 00 00 00 00 CE 00 0A E8 F8 00 00 00 00 00 A8 09 |

Table 579: Example: sSN LFEinfringementinfo

    - Fieldset selection method

| Telegram structure: sWN FieldSetSelectionMethod |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Fieldset selection Method | String | 23 | All | FieldsSetSelectionMethod | 46 69 65 6C 64 53 65 74 53 65 6C 65 63 74 69 6F 6E 4D 65 74 68 6F 64 |
|  |  |  | 1 | All | 0: digital inputs 1: telegram | 00 01 |

Table 580: Telegram structure: sWN FieldSetSelectionMethod

Example for TiM781: sWN FieldSetSelectionMethod

| CoLa A | ASCII | <STX> sWN {SPC} FieldSetSelection {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 46 69 65 6C 64 53 65 74 53 65 6C 65 63 74 69 6F 6E 4D 65 74 68 6F 64 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 4E 20 46 69 65 6C 64 53 65 74 53 65 6C 65 63 74 69 6F 6E 4D 65 74 68 6F 64 01 14 |

Table 581: Example: sWN FieldSetSelectionMethod

| Telegram structure: sWA FieldSetSelectionMethod |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Fieldset selection Method is active | String | 23 | All | FieldSetSelectionMethod | 46 69 65 6C 64 53 65 74 53 65 6C 65 63 74 69 6F 6E 4D 65 74 68 6F 64 |

Table 582: Telegram structure: sWA FieldSetSelectionMethod

Example: sWA FieldSetSelectionMethod

| CoLa A | ASCII | <STX> sWA {SPC} FieldSetSelectionMethod <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 46 69 65 6C 64 53 65 74 53 65 6C 65 63 74 69 6F 6E 4D 65 74 68 6F 64 03 |
| CoLa B | Binary | 02 02 02 00 00 00 15 73 52 41 20 46 69 65 6C 64 53 65 74 53 65 6C 65 63 74 69 6F 6E 4D 65 74 68 6F 64 20 1A |

Table 583: Example: sWA FieldSetSelectionMethod

    - Read active fieldset

| Telegram structure: sRN ActiveFieldSet |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Read | String | 3 | All | sRN | 73 52 4E |
| Command | Info of active fieldset | String | 14 | All | ActiveFieldSet | 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 |

Table 584: Telegram structure: sRN ActiveFieldSet

Example: sRN ActiveFieldSet

| CoLa A | ASCII | <STX> sRN {SPC} ActiveFieldSet <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 4E 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 52 4E 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 20 20 63 |

Table 585: Example: sRN ActiveFieldSet

| Telegram structure: sRA ActiveFieldSet |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sRA | 73 52 41 |
| Command | Active fieldset | String | 14 | All | ActiveFieldSet | 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 |
| Active Fieldset |  |  | 1 | All | 1h &hellip;10h | 00 01... 00 10 |

Table 586: Telegram structure: sRA ActiveFieldSet

Example: sRA ActiveFieldSet

| CoLa A | ASCII | <STX> sRA {SPC} ActiveFieldSet {SPC} F <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 52 41 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 46 03 |
| CoLa B | Binary | 02 02 02 00 00 15 73 52 41 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 20 00 0F 43 |

Table 587: Example: sRA ActiveFieldSet

    - Write active fieldset

| Telegram structure: sWN ActiveFieldSet |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Write | String | 3 | All | sWN | 73 57 4E |
| Command | Active fieldset | String | 14 | All | ActiveFieldSet | 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 |
|  |  |  | 2 | All | 0h... 10h | 00 01... 0010 |

Table 588: Telegram structure: sWN ActiveFieldSet

Example: sWN ActiveFieldSet

| CoLa A | ASCII | <STX> sWN {SPC} ActiveFieldSet {SPC} E <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 4E 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 20 45 03 |
| CoLa B | Binary | 02 02 02 00 00 13 73 57 4E 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 20 00 0E 63 |

Table 589: Example: sWN ActiveFieldSet

| Telegram structure: sWA ActiveFieldSet |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Telegram part | Description | Variable | Length | Sensor | Values CoLa A (ASCII) | Values CoLa B (Binary) |
| Command type | Answer | String | 3 | All | sWA | 73 57 41 |
| Command | Active Fieldset | String | 14 | All | ActiveFieldSet | 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 |

Table 590: Telegram structure: sWA ActiveFieldSet

Example: sWA ActiveFieldSet

| CoLa A | ASCII | <STX> sWA {SPC} ActiveFieldSet <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 57 41 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 03 |
| CoLa B | Binary | 02 02 02 00 00 00 13 73 57 41 20 41 63 74 69 76 65 46 69 65 6C 64 53 65 74 20 49 |

Table 591: Example: sWA ActiveFieldSet

- Diagnostics

  - SOPAS error codes

sFA ErrorCode

| Telegram structure: sFA ErrorCode |  |  |  |
| --- | --- | --- | --- |
| Error code | Description | Dec. | Hex. |
| Sopas_ OK | No error | 0 | 0 |
| Sopas_Error_METHODIN_ACCESSDENIED | Wrong userlevel, access to method not allowed | 1 | 1 |
| Sopas_Error_METHODIN_UNKNOWNINDEX | Trying to access a method with an unknown Sopas index | 2 | 2 |
| Sopas_Error_VARIABLE_UNKNOWNINDEX | Trying to access a variable with an unknown Sopas index | 3 | 3 |
| Sopas_Error_LOCALCONDITIONFAILED | Local condition violated, e.g. giving a value that exceeds the minimum or maximum allowed value for this variable | 4 | 4 |
| Sopas_Error_INVALID_DATA | Invalid data given for variable, this errorcode is deprecated (is not used anymore). | 5 | 5 |
| Sopas_Error_UNKNOWN_ERROR | An error with unknown reason occurred, this errorcode is deprecated. | 6 | 6 |
| Sopas_Error_BUFFER_OVERFLOW | The communication buffer was too small for the amount of data that should be serialised. | 7 | 7 |
| Sopas_Error_BUFFER_UNDERFLOW | More data was expected, the allocated buffer could not be filled. | 8 | 8 |
| Sopas_Error_ERROR_UNKNOWN_TYPE | The variable that shall be serialised has an un- known type. This can only happen when there are variables in the firmware of the device that do not exist in the released description of the device. This should never happen. | 9 | 9 |
| Sopas_Error_VARIABLE_WRITE_ACCESSDENIED | It is not allowed to write values to this variable. Probably the variable is defined as read-only. | 10 | A |
| Sopas_Error_UNKNOWN_CMD_FOR_NAMESERVER | When using names instead of indices, a command was issued that the nameserver does not under- stand. | 11 | B |
| Sopas_Error_UNKNOWN_COLA_COMMAND | The CoLa protocol specification does not define the given command, command is unknown. | 12 | C |
| Sopas_Error_METHODIN_SERVER_BUSY | It is not possible to issue more than one command at a time to an SRT device. | 13 | D |
| Sopas_Error_FLEX_OUT_OF_BOUNDS | An array was accessed over its maximum length. | 14 | E |

| Telegram structure: sFA ErrorCode |  |  |  |
| --- | --- | --- | --- |
| Error code | Description | Dec. | Hex. |
| Sopas_Error_EVENTREG_UNKNOWNINDEX | The event you wanted to register for does not exist, the index is unknown. | 15 | F |
| Sopas_Error_COLA_A_VALUE_OVERFLOW | The value does not fit into the value field, it is too large. | 16 | 10 |
| Sopas_Error_COLA_A_INVALID_CHARACTER | Character is unknown, probably not alphanumeric. | 17 | 11 |
| Sopas_Error_OSAI_NO_MESSAGE | Only when using SRTOS in the firmware and distri- buted variables this error can occur. It is an indica- tion that no operating system message could be created. This happens when trying to GET a vari- able. | 18 | 12 |
| Sopas_Error_OSAI_NO_ANSWER_MESSAGE | This is the same as Sopas_Error_OSAI_NO_MESSAGE with the difference that it is thrown when trying to PUT a variable. | 19 | 13 |
| Sopas_Error_INTERNAL | Internal error in the firmware, problably a pointer to a parameter was null. | 20 | 14 |
| Sopas_Error_HubAddressCorrupted | The Sopas Hubaddress is either too short or too long. | 21 | 15 |
| Sopas_Error_HubAddressDecoding | The Sopas Hubaddress is invalid, it can not be decoded (Syntax). | 22 | 16 |
| Sopas_Error_HubAddressAddressExceeded | Too many hubs in the address | 23 | 17 |
| Sopas_Error_HubAddressBlankExpected | When parsing a HubAddress an expected blank was not found. The HubAddress is not valid. | 24 | 18 |
| Sopas_Error_AsyncMethodsAreSuppressed | An asynchronous method call was made although the device was built with "AsyncMethodsSuppressed”. This is an internal error that should never happen in a released device. | 25 | 19 |
| Sopas_Error_ComplexArraysNotSupported | Device was built with „ComplexArraysSuppressed" because the compiler does not allow recursions. But now a complex array was found. This is an in- ternal error that should never happen in a released device. | 26 | 20 |

Table 592: SOPAS error codes

Example: sFA ErrorCode Wrong userlevel

| CoLa A | ASCII | <STX> sFA {SPC} 1 <ETX> |
| --- | --- | --- |
|  | Hex | 02 73 46 41 20 31 03 |
| CoLa B | Binary | 02 02 02 00 00 00 05 73 46 41 20 00 01 75 |

Table 593: Example: sFA ErrorCode Wrong userlevel

  - Additional information

SOPAS communication is a index based communication and can be identified with telegram beginning of: sRI,sWI,sMI, sAI, sEI,sSI. A parallel usage of one port can confuse. So we recommend a separat usage of different ports, if possible.

Every response telegram starts with a separat framed string:

<STX>sSI 2 1<ETX><STX>"Answer”<ETX>

- List of tables

Table 1: 	Example: Binary telegram 	8

Table 2: 	Example: ASCII telegram 	9

Table 3: 	Telegram structure: sMN SetAccessMode 	14

Table 4: 	Example: sMN SetAccessMode 	14

Table 5: 	Telegram structure: sAN SetAccessMode 	15

Table 6: 	Example for LMS100: sAN SetAccessMode 	15

Table 7: 	Telegram structure: sMN mLMPsetscancfg 	20

Table 8: 	Example: sMN mLMPsetscancfg for LMS1xx with 1 measurement sector of 270° 	21

Table 9: 	Example: sMN mLMPsetscancfg for LD-XXX###1 with 1 measurement sector of 360° 	22

Table 10: 	Example: sMN mLMPsetscancfg for LD-XXX###1 with

- measurement sector of 270° 	22

Table 11: 	Example: sMN mLMPsetscancfg for LD-XXX###1 with

- measurement sectors 	23

Table 12: 	Example: sMN mLMPsetscancfg for LD-XXX###1 with

4 measurement sectors 	23

Table 13: 	Example: sMN mLMPsetscancfg for LD-XXX###0 with 1 measurement sector of 360° 	24

Table 14: 	Example: sMN mLMPsetscancfg for LD-XXX###0 with 1 measurement sector of 270° 	24

Table 15: 	Example: sMN mLMPsetscancfg for LD-XXX###0 with 2 measurement sectors 	25

Table 16: 	Example: sMN mLMPsetscancfg for LD-XXX###0 with 4 measurement sectors 	25

Table 17: 	Telegram structure: sAN mLMPsetscancfg 	27

Table 18: 	Example: sAN mLMPsetscancfg 	28

Table 19: 	Telegram structure: sRN LMPscancfg 	28

Table 20: 	Example for LMS100: sRN LMPscancfg 	28

Table 21: 	Telegram structure: sRA LMPscancfg 	32

Table 22: 	Example: sRA LMPscancfg 	32

Table 23: 	Telegram structure: sWN MMAlignmentMode 	33

Table 24: 	Example: sWN MMAlignmentMode 	33

Table 25: 	Telegram structure: sWA MMAlignmentMode 	33

Table 26: 	Example: sWA MMAlignmentMode 	34

Table 27: 	Telegram structure: sMN mCLsetscancfglist 	34

Table 28: 	Interlace mode for sMN mCLsetscancfglist for LD series 	35

Table 29: 	Interlace mode for sMN mCLsetscancfglist for LMS5xx 	35

Table 30: 	Interlace mode for sMN mCLsetscancfglist for LRS4000 	36

Table 31: 	Example: Set scan configuration 1: sMN mCLsetscancfglist 1 	36

Table 32: 	Telegram structure: sAN mCLsetscancfglist 	37

Table 33: 	Example: sAN mCLsetscancfglist OK 	37

Table 34: 	Telegram structure: sMN mCLsetscancfglist 	38

Table 35: 	Example: sMN mCLsetscancfglist 5 	38

Table 36: Telegram structure: sAN mCLsetscancfglist 	39

Table 37: 	Example: sAN mCLsetscancfglist 0 	39

Table 38: 	Telegram structure: sMN mCLsetscancfglist 	40

Table 39: 	Example: Set scan configuration 1: sMN mCLsetscancfglist 1 	40

Table 40: 	Telegram structure: sAN mCLsetscancfglist 	41

Table 41: 	Example: sAN mCLsetscancfglist OK 	41

Table 42: 	Telegram structure: sMN LMCstandby 	42

Table 43: 	Example: sMN LMCstandby 	42

Table 44: 	Telegram structure: sAN LMCstandby 	42

Table 45: 	Example: sAN LMCstandby 	43

Table 46: 	Telegram structure: sMN LMCstartmeas 	44

Table 47: 	Example: sMN LMCstartmeas 	44

Table 48: 	Telegram structure: sAN LMCstartmeas 	45

Table 49: 	Example: sAN LMCstartmeas 	45

Table 50: 	Telegram structure: sMN LMCstopmeas 	46

Table 51: 	Example: sMN LMCstopmeas 	46

Table 52: 	Telegram structure: sAN LMCstopmeas 	46

Table 53: 	Example: sAN LMCstopmeas 	47

Table 54: 	Telegram structure: sWN LMPautostartmeas 	47

Table 55: 	Example: sWN LMPautostartmeas 1 	47

Table 56: 	Telegram structure: sWA LMDautostartmeas 	48

Table 57: 	Example: sWA LMPautostartmeas 	48

Table 58: 	Telegram structure: sWN IOlasc 	49

Table 59: 	Example: sWN IOlasc with Software Trigger and 0.5 s delay on

start. 	49

Table 60: 	Telegram structure: sWA IOlasc 	49

Table 61: 	Example: sWA IOlasc 	49

Table 62: 	Telegram structure: sWN CLApplication 	50

Table 63: 	Example: Activate the field application: sWN CLApplication 11 	50

Table 64: 	Telegram structure: sWA CLApplication 	50

Table 65: 	Example: sWA CLApplication correct and accepted 	51

Table 66: 	Telegram structure: sWN SetActiveApplications 	51

Table 67: 	Example: Activate the field application: : sWN

SetActiveApplications 1 FEVL 1 	51

Table 68: 	Telegram structure: sWA SetActiveApplications 	52

Table 69: 	Example: sWA SetActiveApplications correct and accepted 	52

Table 70: 	Telegram structure: sRN SetActiveApplications 	52

Table 71: 	Example for MRS1000: sRN SetActiveApplications 	52

Table 72: 	Telegram structure: sMN mSCloadfacdef 	53

Table 73: 	Example: sMN mSCloadfacdef 	54

Table 74: 	Telegram structure: sAN mSCloadfacdef 	54

Table 75: 	Example: sAN mSCloadfacdef 	54

Table 76: 	Telegram structure: sMN mSCloadappdef 	55

Table 77: 	Example: sMN mSCloadappdef 	55

Table 78: 	Telegram structure: sAN mSCloadappdef 	56

Table 79: 	Example: sAN mSCloadappdef 	56

Table 80: 	Telegram structure: sMN SetPassword 	57

Table 81: 	Example: sMN SetPassword 	57

Table 82: 	Telegram structure: sAN SetPassword 	58

Table 83: 	Example: sAN SetPassword 	58

Table 84: 	Telegram structure: sMN CheckPassword 	59

Table 85: 	Example: sMN CheckPassword 	59

Table 86: 	Telegram structure: sAN CheckPassword 	59

Table 87: 	Example: sAN CheckPassword 	60

Table 88: 	Telegram structure: sMN mSCreboot 	60

Table 89: 	Example: sMN mSCreboot 	60

Table 90: 	Telegram structure: sAN mSCreboot 	61

Table 91: 	Example: sAN mSCreboot 	61

Table 92: 	Telegram structure: sWN LCMcfg 	62

Table 93: 	Example: sWN LCMcfg 	62

Table 94: 	Telegram structure: sWA LCMcfg 	63

Table 95: 	Example: sWA LCMcfg 	63

Table 96: 	Telegram structure: sWN ContaminationConfig 	64

Table 97: 	Example: sWN ContaminationConfig 	65

Table 98: 	Telegram structure: sWA ContaminationConfig 	65

Table 99: 	Example: sWA ContaminationConfig 	65

Table 100: 	Telegram structure: sRN LCMcfg 	65

Table 101: 	Example: sRN LCMcfg 	66

Table 102: 	Telegram structure: sRA LCMcfg 	66

Table 103: 	Example: sRA LCMcfg 	66

Table 104: 	Telegram structure: sRN ContaminationConfig 	67

Table 105: 	Example: sRN ContaminationConfig 	67

Table 106: 	Telegram structure: sRA ContaminationConfig 	68

Table 107: 	Example: sRA ContaminationConfig 	68

Table 108: 	Telegram structure: sRN CMContLvlM 	69

Table 109: 	Example: sRN CMContLvlM 	69

Table 110: 	Telegram structure: sRA CMContLvlM 	69

Table 111: 	Example for LMS5xx: sRA CMContLvlM 	70

Table 112: 	Telegram structure: sRN ContaminationData 	70

Table 113: 	Example: sRN ContaminationData 	70

Table 114: 	Telegram structure: sRA ContaminationData 	71

Table 115: 	Example: sRA ContaminationData 	71

Table 116: 	Telegram structure: sRN ContaminationResult 	71

Table 117: 	Example: sRN ContaminationResult 	72

Table 118: 	Telegram structure: sRA ContaminationResult 	72

Table 119: 	Example: sRA ContaminationResult 	72

Table 120: 	Telegram structure: sMN mEEwriteall 	73

Table 121: 	Example: sMN mEEwriteall 	73

Table 122: 	Telegram structure: sAN mEEwriteall 	74

Table 123: 	Example: sAN mEEwriteall 	74

Table 124: 	Telegram structure: sMN Run 	75

Table 125: 	Example: sMN Run 	75

Table 126: 	Telegram structure: sAN Run 	76

Table 127: 	Example: sAN Run 	76

Table 128: 	Telegram structure: sWN LMDscandatacfg 	79

Table 129: 	Example 1: sWN LMDscandatacfg 	79

Table 130: 	Example 2: sWN LMDscandatacfg 	79

Table 131: 	Example3: sWN LMDscandatacfg 	80

Table 132: 	Telegram structure: sWA LMDscandatacfg 	80

Table 133: 	Example: sWA LMDscandatacfg 	80

Table 134: 	Telegram structure: sWN LMPoutputRange 	82

Table 135: 	Example: sWN LMPoutputRange 0,50°resolution, 0°-90° 	82

Table 136: 	Telegram structure: sWA LMPoutputRange 	83

Table 137: 	Example: sWA LMPoutputRange 	83

Table 138: 	Telegram structure: sRN LMPoutputRange 	84

Table 139: 	Example: sRN LMPoutputRange 	84

Table 140: 	Telegram structure: sRA LMPoutputRange 	86

Table 141: 	Example: sRA LMPoutputRange 	86

Table 142: 	Telegram structure: sWN LMDscandatascalefactor 	87

Table 143: 	Example: sWN LMDscandatascalefactor 1,0 	87

Table 144: 	Telegram structure: sWA LMDscandatascalefactor 	88

Table 145: 	Example: sWA LMDscandatascalefactor 	88

Table 146: 	Telegram structure: sRN LMDscandatascalefactor 	88

Table 147: 	Example: sRN LMPoutputRange 	88

Table 148: 	Telegram structure: sRA LMDscandatascalefactor 	89

Table 149: 	Example: sRA LMDscandatascalefactor 	89

Table 150: 	Telegram structure: sRN LMDscandata 	90

Table 151: 	Example: sRN LMDscandata 	90

Table 152: 	Telegram structure: sRA LMDscandata 	91

Table 153: 	Example: sRA LMDscandata 	91

Table 154: 	Telegram structure: sEN LMDscandata 	92

Table 155: 	Example: sEN LMDscandata 	92

Table 156: 	Telegram structure: sEA LMDscandata 	93

Table 157: 	Example: Confirmation of sEA LMDscandata 	93

Table 158: 	Telegram structure: Datastream of sRA LMDscandata/sSN LMDscandata 	105

Table 159: 	Example of one telegram stream 	111

Table 160: 	Telegram structure: sMN LSPsetdatetime 	113

Table 161: 	Example 1: sMN LSPsetdatetime 	113

Table 162: 	Example 2: sMN LSPsetdatetime 	113

Table 163: 	Telegram structure: sAN LSPsetdatetime 	113

Table 164: 	Example 1, 2: sAN LSPsetdatetime 	114

Table 165: 	Telegram structure: sRN STlms 	114

Table 166: 	Example: sRN STlms 	114

Table 167: 	Telegram structure: sRA STlms 	115

Table 168: 	Example: sRA STlms 	116

Table 169: 	Telegram structure: sRN DeviceTime 	116

Table 170: 	Example: sRN DeviceTime 	116

Table 171: 	Telegram structure: sRA DeviceTime 	117

Table 172: 	Example: sRA DeviceTime 0 	117

Table 173: 	Telegram structure: sWN TSCRole 	117

Table 174: 	Example: sWN TSCRole 	118

Table 175: 	Telegram structure: sWA TSCRole 	118

Table 176: 	Example: sWA TSCRole 	118

Table 177: 	Telegram structure: sWN TSCTCInterface 	119

Table 178: 	Example: sWN TSCTCInterface 	119

Table 179: 	Telegram structure: sWA TSCTCInterface 	119

Table 180: 	Example: sWA TSCTCInterface 	120

Table 181: 	Telegram structure: sWN TSCTCSrvAddr 	120

Table 182: 	Example: sWN TSCTCSrvAddr 192.168.0.11 	120

Table 183: 	Telegram structure: sWA TSCTCSrvAddr 	121

Table 184: 	Example: sWA TSCTCSrvAddr 	121

Table 185: 	Telegram structure: sWN TSCTCtimezone 	121

Table 186: 	Example: sWN TSCTCtimezone GMT +1 hour 	122

Table 187: 	Telegram structure: sWA TSCTCtimezone 	122

Table 188: 	Example: sWA TSCTCtimezone 	122

Table 189: 	Telegram structure: sWN TSCTCupdatetime 	123

Table 190: 	Example: sWN TSCTCupdatetime 600 s 	123

Table 191: 	Telegram structure: sWA TSCTCupdatetime 	123

Table 192: 	Example: sWA TSCTCupdatetime 	124

Table 193: 	Telegram structure: sRN TSCTCmaxoffset 	124

Table 194: 	Example: sRN TSCTCmaxoffset 	124

Table 195: 	Telegram structure: sRA TSCTCmaxoffset 	125

Table 196: 	Example: sRA TSCTCmaxoffset 18000 s 	125

Table 197: 	Telegram structure: sRN TSCTCdelay 	125

Table 198: 	Example: sRN TSCTCdelay 	126

Table 199: 	Telegram structure: sRA TSCTCdelay 	126

Table 200: 	Example: sRA TSCTCdelay 0.003 s 	126

Table 201: 	Telegram structure: sMN mResetMaxOff 	127

Table 202: 	Example: sMN mResetMaxOff 	127

Table 203: 	Telegram structure: sAN mResetMaxOff 	127

Table 204: 	Example: sAN mResetMaxOff 	127

Table 205: 	Telegram structure: sWN LFPparticle 	129

Table 206: 	Example: sWN LFPparticle 	129

Table 207: 	Telegram structure: sWA LFPparticle 	129

Table 208: 	Example: sWA LFPparticle 	130

Table 209: 	Telegram structure: sWN LFPmeanfilter 	130

Table 210: 	Example: sWN LFPmeanfilter 	130

Table 211: 	Telegram structure: sWA LFPmeanfilter 	131

Table 212: 	Example: sWA LFPmeanfilter 	131

Table 213: 	Telegram structure: sWN LFPnto1filter 	131

Table 214: 	Example: sWN LFPnto1filter 	132

Table 215: 	Telegram structure: sWA LFPnto1filter 	132

Table 216: 	Example: sWA LFPnto1filter 	132

Table 217: 	Telegram structure: sWN FREchoFilter 	133

Table 218: 	Example: sWN FREchoFilter 	133

Table 219: 	Telegram structure: sWA FREchoFilter 	133

Table 220: 	Example: sWa FREchoFilter 	133

Table 221: 	Telegram structure: sWN MSsuppmode 	134

Table 222: 	Example: sWN MSsuppmode 	134

Table 223: 	Telegram structure: sWA MSsuppmode 	134

Table 224: 	Example: sWA MSsuppmode 	135

Table 225: 	Telegram structure: sWN CLFogFilterEn 	135

Table 226: 	Example: sWN CLFogFilterEn 	135

Table 227: 	Telegram structure: sWA CLFogFilterEn 	136

Table 228: 	Example: sWA CLFogFilterEn 	136

Table 229: 	Telegram structure: sRN CLFogFilterEn 	136

Table 230: 	Example: sRN CLFogFilterEn 	136

Table 231: 	Telegram structure: sRA CLFogFilterEn 	137

Table 232: 	Example: sRA CLFogFilterEn 	137

Table 233: 	Telegram structure: sWN MCSenseLevel 	137

Table 234: 	Example: sWN MCSenseLevel 	138

Table 235: 	Telegram structure: sWA MCSenseLevel 	138

Table 236: 	Example: sWA MCSenseLevel 	138

Table 237: 	Telegram structure: sWN FogFilterMaxRangeEnable 	139

Table 238: 	Example: sWN FogFilterMaxRangeEnable 1 	139

Table 239: 	Telegram structure: sWA FogFilterMaxRangeEnable 	139

Table 240: 	Example: sWA FogFilterMaxRangeEnable 	139

Table 241: 	Telegram structure: sMN TeachFogFilterMaxRange 	140

Table 242: 	Example: sMN TeachFogFilterMaxRange 	140

Table 243: 	Telegram structure: sAN TeachFogFilterMaxRange 	140

Table 244: 	Example: sAN TeachFogFilterMaxRange 	141

Table 245: 	Telegram structure: sWN FogFilterMaxRange 	141

Table 246: 	Example: sWN FogFilterMaxRange 43FA0000 	141

Table 247: 	Telegram structure: sWA FogFilterMaxRange 	142

Table 248: 	Example: sWA FogFilterMaxRange 	142

Table 249: 	Telegram structure: sWN CLNFDigFilterEn 	142

Table 250: 	Example: sWN CLNFDigFilterEn 	143

Table 251: 	Telegram structure: sWA CLNFDigFilterEn 	143

Table 252: 	Example: sWA CLNFDigFilterEn 	143

Table 253: 	Telegram structure: sWN CLHWFilterSectEn 	144

Table 254: 	Example: sWN CLHWFilterSectEn 1 0 0 0 	144

Table 255: 	Telegram structure: sWA CLHWFilterSectEn 	144

Table 256: 	Example: sWA CLHWFilterSectEn 	145

Table 257: 	Telegram structure: sWN LFPmedianfilter 	145

Table 258: 	Example: sWN LFPmedianfilter 1 3 	145

Table 259: 	Telegram structure: sWA LFPmedianfilter 	146

Table 260: 	Example: sWA LFPmedianfilter 	146

Table 261: 	Telegram structure: sWN LFPedgefilter 	146

Table 262: 	Example: sWN LFPedgefilter 1 	147

Table 263: 	Telegram structure: sWA LFPmedianfilter 	147

Table 264: 	Example: sWA LFPedgefilter 	147

Table 265: 	Telegram structure: sWN LFPedgefilterMaxDist 	148

Table 266: 	Example: sWN LFPedgefilterMaxDist +5 	148

Table 267: 	Telegram structure: sWA LFPedgefilterMaxDist 	148

Table 268: 	Example: sWA LFPedgefilterMaxDist 	148

Table 269: 	Telegram structure: sWN LFPcubicareafilter 	149

Table 270: 	Example: sWN LFPcubicareafilter 1 +10000 +20000 -15000

+15000 ............................................................................................ 150

Table 271: 	Telegram structure: sWA LFPcubicareafilter 	150

Table 272: 	Example: sWA LFPedgefilterMaxDist 	150

Table 273: 	Telegram structure: sWN LFPglossfilter 	151

Table 274: 	Example: sWN LFPglossfilter 1 	151

Table 275: 	Telegram structure: sWA LFPglossfilter 	152

Table 276: 	Example: sWA LFPglossfilter 	152

Table 277: 	Telegram structure: sWN EnableLongRangeMode 	153

Table 278: 	Example: sWN EnableLongRangeMode 1 	153

Table 279: 	Telegram structure: sWA EnableLongeRangeModer 	154

Table 280: 	Example: sWA EnableLongeRangeMode 	154

Table 281: 	Telegram structure: sWN EnableLongRangeMode 	154

Table 282: 	Example: sRN EnableLongRangeMode 	155

Table 283: 	Telegram structure: sRA EnableLongRangeMode 	155

Table 284: 	Example: sWA EnableLongrangeMode 	155

Table 285: 	Telegram structure: sWN LICsrc 	156

Table 286: 	Example: sWN LICsrc 	156

Table 287: 	Telegram structure: sWA LICsrc 	157

Table 288: 	Example: sWA LICsrc 	157

Table 289: 	Telegram structure: sWN LICencset 	158

Table 290: 	Example: sWN LICencset 	158

Table 291: 	Telegram structure: sWA LICencset 	159

Table 292: 	Example: sWA LICencset 	159

Table 293: 	Telegram structure: sWN LICencres 	159

Table 294: 	Example: sWN LICencres 	160

Table 295: 	Telegram structure: sWA LICencres 	160

Table 296: 	Example: sWA LICencres 	160

Table 297: 	Telegram structure: sWN LICFixVel 	161

Table 298: 	Example: sWN LICFixVel 	161

Table 299: 	Telegram structure: sWA LICFixVel 	161

Table 300: 	Example: sWA LICFixVel 	162

Table 301: 	Telegram structure: sRN LICSpTh 	162

Table 302: 	Example: sRN LICSpTh 	162

Table 303: 	Telegram structure: sRA LICSpTh 	163

Table 304: 	Example: sRA LICSpTh 	163

Table 305: 	Telegram structure: sRN LICencsp 	163

Table 306: 	Example: sRN LICencsp 	164

Table 307: 	Telegram structure: sRA LICencsp 	164

Table 308: 	Example: sRA LICencsp 	164

Table 309: 	Telegram structure: sMN LIDrstencoderinc 	165

Table 310: 	Example: sMN LIDrstencoderinc 	165

Table 311: 	Telegram structure: sAN LIDrstencoderinc 	165

Table 312: 	Example: sAN LIDrstencoderinc 	165

Table 313: 	Telegram structure: sRN LIDportstate 	166

Table 314: 	Example: sRN LIDportstate 	167

Table 315: 	Telegram structure: sRA/sSN LIDportstate 	168

Table 316: 	Example: sRA LIDportstate 	169

Table 317: 	Telegram structure: sRN PortConfiguration 	169

Table 318: 	Example: sRN PortConfiguration 	169

Table 319: 	Telegram structure: sRN LIDinputstate 	172

Table 320: 	Example: sRN LIDinputstate 	172

Table 321: 	Telegram structure: sRA LIDinputstate 	173

Table 322: 	Example: sRA LIDinputstate 	173

Table 323: 	Telegram structure: sRN LIDoutputstate 	174

Table 324: 	Example: sRN LIDoutputstate 	174

Table 325: 	Telegram structure: sRA LIDoutputstate 	174

Table 326: 	Telegram structure: sEN LIDoutputstate 	175

Table 327: 	Example: sEN LIDoutputstate 	175

Table 328: 	Telegram structure: sRA/sSN LIDoutputstate 	179

Table 329: 	Example: sRA LIDoutputstate 	179

Table 330: 	Telegram structure: sMN mDOSetOutput 	180

Table 331: 	Example: sMN mDOSetOutput 	181

Table 332: 	Telegram structure: sAN mDOSetOutput 	181

Table 333: 	Example: sAN mDOSetOutput 	181

Table 334: 	Telegram structure PRO: sWN DO6Fnc/Lite: sWN DO3Fnc 	182

Table 335: 	Example: sWN DO6Fnc  Out6 to master sync 	182

Table 336: 	Telegram structure: PRO: sWN DO6Fnc/Lite: sWN DO3Fnc 	183

Table 337: 	Example: sWA DO6Fnc 	183

Table 338: 	Telegram structure: sWN DO1Fnc 	183

Table 339: 	Example: sWN DO1Fnc  Out1 to device ready 	184

Table 340: 	Telegram structure: sWA DO1Fnc 	184

Table 341: 	Example: sWA DO1Fnc 	184

Table 342: 	Telegram structure: sWN DO1Logic 	185

Table 343: 	Example: sWN DO1Logic  Active_Low 	185

Table 344: 	Telegram structure: sWA DO1Logic 	185

Table 345: 	Example: sWA DO1Logic 	186

Table 346: 	Telegram structure: sWN DO2Fnc 	186

Table 347: 	Example: sWN DO2Fnc  Out2 to device ready 	186

Table 348: 	Telegram structure: sWA DO2Fnc 	187

Table 349: 	Example: sWA DO2Fnc 	187

Table 350: 	Telegram structure: sWN DO2Logic 	187

Table 351: 	Example: sWN DO2Logic  Active_High 	188

Table 352: 	Telegram structure: sWA DO2Logic 	188

Table 353: 	Example: sWA DO2Logic 	188

Table 354: 	Telegram structure: sWN SYMode 	189

Table 355: 	Example: sWN SYMode 	189

Table 356: 	Telegram structure: sWA SYMode 	189

Table 357: 	Example: sWA SYMode 	189

Table 358: 	Telegram structure: sWN SYPhase 	190

Table 359: 	Example: sWN SYPhase +90 	190

Table 360: 	Telegram structure: sWA SYPhase 	190

Table 361: 	Example: sWA SYPhase 	191

Table 362: 	Telegram structure: sWN DO3And4Fnc 	191

Table 363: 	Example: sWN In4  In3+4 to slave sync 	191

Table 364: 	Telegram structure: sWA DO3And4Fnc 	192

Table 365: 	Example: sWA DO3And4Fnc 	192

Table 366: 	Telegram structure: sWN DI3DebTim 	192

Table 367: 	Example: sWN DI3DebTim 	193

Table 368: 	Telegram structure: sWA DI3DebTim 	193

Table 369: 	Example: sWA DI3DebTim 	193

Table 370: 	Telegram structure: sRN SYextmon 	193

Table 371: 	Example: sRN SYextmon 	194

Table 372: 	Telegram structure: sRA SYextmon 	194

Table 373: 	Example: sRA SYextmon 	194

Table 374: 	Telegram structure: sRN LCMstate 	195

Table 375: 	Example: sRN LCMstate 	196

Table 376: 	Telegram structure: sRA LCMstate 	196

Table 377: 	Example for LMS100: sRA LCMstate 	196

Table 378: 	Telegram structure: sRN DeviceIdent 	197

Table 379: 	Example: sRN DeviceIdent 	197

Table 380: 	Telegram structure: sRA DeviceIdent 	198

Table 381: 	Example: sRA DeviceIdent 	198

Table 382: 	Telegram structure: sRN SCdevicestate 	199

Table 383: 	Example: sRN SCdevicestate 	199

Table 384: 	Telegram structure: sRA SCdevicestate 	200

Table 385: 	Example: sRA SCdevicestate 	200

Table 386: 	Telegram structure: sRN LMCmeasstate 	202

Table 387: 	Example: sRN LMCmeasstate 	202

Table 388: 	Telegram structure: sRA LMCmeasstate 	202

Table 389: 	Example: sRA LMCmeasstate is Measure 2D 	203

Table 390: 	Telegram structure: sRN EMCustomerInfo 	203

Table 391: 	Example: sRN EMCustomerInfo 	203

Table 392: 	Telegram structure: sRA EMCustomerInfo 	204

Table 393: 	Example: sRA EMCustomerInfo = Device OK 	204

Table 394: 	Telegram structure: sRN LDMSenStat 	205

Table 395: 	Example: sRN LDMSenStat 	205

Table 396: 	Telegram structure: sRA LDMSenStat 	206

Table 397: 	Example: sRA LDMSenStat Device is in Idle mode 	206

Table 398: 	Telegram structure: sRN DIornr 	207

Table 399: 	Example: sRN DIornr 	207

Table 400: 	Telegram structure: sRA DIornr 	207

Table 401: 	Example for LMS511-20100: sRA DIornr 	208

Table 402: 	Example for TiM561-2050101: sRA DIornr 	208

Table 403: 	Telegram structure: sRN DItype 	208

Table 404: 	Example: sRN DItype 	208

Table 405: 	Telegram structure: sRA DItype 	209

Table 406: 	Example for LMS511-20100: sRA DItype 	209

Table 407: 	Example for TiM561-2050101: sRA DItype 	209

Table 408: 	Telegram structure: sRN ODoprh 	210

Table 409: 	Example: sRN ODoprh 	210

Table 410: 	Telegram structure: sRA ODoprh 	211

Table 411: 	Example: sRA ODoprh 	211

Table 412: 	Telegram structure: sRN ODpwrc 	212

Table 413: 	Example: sRN ODpwrc 	212

Table 414: 	Telegram structure: sRA ODpwrc 	213

Table 415: 	Example: sRA ODpwrc 	213

Table 416: 	Telegram structure: sRN OPcurtmpdev 	213

Table 417: 	Example: sRN OPcurtmpdev 	214

Table 418: 	Telegram structure: sRA OPcurtmpdev 	214

Table 419: 	Example: sRA OPcurtmpdev 	214

Table 420: 	Telegram structure: sWN LocationName 	215

Table 421: 	Example: sWN LocationName +13 OutdoorDevice 	215

Table 422: 	Telegram structure: sWA LocationName 	216

Table 423: 	Example: sWA LocationName 	216

Table 424: 	Telegram structure: sRN LocationName 	216

Table 425: 	Example: sRN LocationName 	217

Table 426: 	Telegram structure: sRA LocationName 	217

Table 427: 	Example: sRA LocationName 	217

Table 428: 	Telegram structure: sRN MCAngleCompSin 	218

Table 429: 	Example: sRN MCAngleCompSin 	218

Table 430: 	Telegram structure: sRA MCAngleCompSin 	218

Table 431: 	Example: sRA MCAngleCompSin 	219

Table 432: 	Telegram structure: sMN LIDrstoutpcnt 	219

Table 433: 	Example: sMN LIDrstoutpcnt 	220

Table 434: 	Telegram structure: sAN LIDrstoutpcnt 	220

Table 435: 	Example: sAN LIDrstoutpcnt 	220

Table 436: 	Telegram structure: sRN OPheatstateext 	221

Table 437: 	Example: sRN OPheatstateext 	221

Table 438: Telegram structure: sRA OPheatstateext 	222

Table 439: 	Example: sRA OPheatstateext 2 	222

Table 440: 	Telegram structure: sEN FieldEvalLoggingList 	222

Table 441: 	Example sEN FieldEvalLoggingList 	223

Table 442: 	Telegram structure: sEA FieldEvalLoggingList 	223

Table 443: 	Example: sEA FieldEvalLoggingList 1 	223

Table 444: 	Telegram structure: sSN Field EvalLogging List 	225

Table 445: 	Example: sSN Field EvalLogging List 	225

Table 446: 	Telegram structure: sRN FieldEvalLoggingList 	225

Table 447: 	Example sEN FieldEvalLoggingList 1 	226

Table 448: 	Telegram structure: sRA FieldEvalLoggingList 	226

Table 449: 	Example: sEA FieldEvalLoggingList 1 	226

Table 450: 	Telegram structure: sWN EIIpAddr 	228

Table 451: 	Example: sWN EIIpAddr 192.168.0.2 	228

Table 452: 	Telegram structure: sWA EIIpAddr 	229

Table 453: 	Example: sWA EIIpAddr 	229

Table 454: 	Telegram structure: sRN EIIpAddr 	230

Table 455: 	Example: srN EIIpAddr 	230

Table 456: 	Telegram structure: sRA EIIpAddr 	231

Table 457: 	Example: sRA EIIpAddr 192.168.0.2 	231

Table 458: 	Telegram structure: sWN Elgate 	232

Table 459: 	Example: sWN EIgate 192.168.0.1 	232

Table 460: 	Telegram structure: sWA Elgate 	232

Table 461: 	Example: sWA EIgate 	233

Table 462: 	Telegram structure: sRN Elgate 	233

Table 463: 	Example: sRN EIgate 	233

Table 464: 	Telegram structure: sRA Elgate 	234

Table 465: 	Example: sRA EIgate 192.168.0.1 	234

Table 466: 	Telegram structure: sWN EImask 	235

Table 467: 	Example: sWN EImask 255.255.254.0 	235

Table 468: 	Telegram structure: sWA EImask 	235

Table 469: 	Example: sWA EImask 	236

Table 470: 	Telegram structure: sRN Elmask 	236

Table 471: 	Example: sRN EImask 	236

Table 472: 	Telegram structure: sRA Elmask 	237

Table 473: 	Example: sRA EImask 255.255.254.0 	237

Table 474: 	Telegram structure: sWN SIHstBaud 	238

Table 475: 	Example: sWN SIHstBaud 	238

Table 476: 	Telegram structure: sWA SIHstBaud 	238

Table 477: 	Example: sWA SIHstBaud 	239

Table 478: 	Telegram structure: sRN SIHstBaud 	239

Table 479: 	Example: sRN SIHstBaud 	239

Table 480: 	Telegram structure: sRA SIHstBaud 	240

Table 481: 	Example: sRA SIHstBaud 	240

Table 482: 	Telegram structure: sWN SIHstHw 	241

Table 483: 	Example: sWN SIHstHw 	241

Table 484: 	Telegram structure: sWA SIHstHw 	241

Table 485: 	Example: sWA SIHstHw 	242

Table 486: 	Telegram structure: sRN SIHstHw 	242

Table 487: 	Example: sRN SIHstHw 	242

Table 488: 	Telegram structure: sRA SIHstHw 	243

Table 489: 	Example: sRA SIHstHw 	243

Table 490: 	Telegram structure: sWN EIHstPort 	244

Table 491: 	Example: sWN EIHstPort +2110 	244

Table 492: 	Example: sWN EIUDPPort +2214 	244

Table 493: 	Telegram structure: sWA EIHstPort 	245

Table 494: 	Example: sWA EIHstPort 	245

Table 495: 	Example: sWA EIUDPPort 	245

Table 496: 	Telegram structure: sWN EIHstCola 	246

Table 497: 	Example: sWN EIHstCola 1 	246

Table 498: 	Example: sWN EIUDPCola 0 	247

Table 499: 	Telegram structure: sWA EIHstCola 	247

Table 500: 	Example: sWA EIHstCola 	247

Table 501: 	Example: sWA EIUDPCola 	247

Table 502: 	Telegram structure: sWN SYtype 	248

Table 503: 	Example: sWN SYtype 2 	248

Table 504: 	Telegram structure: sWA SYtype 	248

Table 505: 	Example: sWA SYtype 	249

Table 506: 	Telegram structure: sWN SYphas 	249

Table 507: 	Example: sWN SYphas +100000 	249

Table 508: 	Telegram structure: sWA SYphas 	250

Table 509: 	Example: sWA SYphas 	250

Table 510: 	Telegram structure: sWN LMLfpen 	251

Table 511: 	Example: sWN LMLfpen 1 (Enable Front Panel) 	252

Table 512: 	Telegram structure: sWA LMLfpen 	252

Table 513: 	Example: sWA LMLfpen 	252

Table 514: 	Telegram structure: sWN LLMfpmode 	253

Table 515: 	Example: sWN LLMfpmode 0 (set to "Permanent”) 	253

Table 516: 	Telegram structure: sWA LLMfpmode 	253

Table 517: 	Example: sWA LLMfpmode 	253

Table 518: 	Telegram structure: sWN LMLfpFcn 	254

Table 519: 	Example: sWN LMLfpFcn 	254

Table 520: 	Telegram structure: sWA LMLfpFcn 	255

Table 521: 	Example: sWA LMLfpFcn 	255

Table 522: 	Telegram structure: sMN mLMLSetLed 	256

Table 523: 	Example: sMN mLMLSetLed 1 1 (Stop LED) 	256

Table 524: 	Telegram structure: sAN mLMLSetLed 	257

Table 525: 	Example: sAN mLMLSetLed 	257

Table 526: 	Telegram structure: sWN HMIfpFcn_Y1 	258

Table 527: 	Example: sWN HMIfpFcn_Y1 2 	258

Table 528: 	Telegram structure: sWA HMIfpFcn_Y1 	259

Table 529: 	Example: sWA HMIfpFcn_Y1 	259

Table 530: 	Telegram structure: sWN HMIfpFcn_Y2 	259

Table 531: 	Example: sWN HMIfpFcn_Y2 = Command 	260

Table 532: 	Telegram structure: sWA HMIfpFcn_Y2 	260

Table 533: 	Example: sWA HMIfpFcn_Y2 	260

Table 534: 	Telegram structure: sMN mHMISetLed 	261

Table 535: 	Example: sMN mHMISetLed 1 = On 	261

Table 536: 	Telegram structure: sAN mHMISetLed 	262

Table 537: 	Example: sAN mHMISetLed 01 	262

Table 538: 	Telegram structure: sMN mLMLSetDisp 	264

Table 539: 	Example: sMN mLMLSetDisp 07 (Showing the number "7” on the display) 	264

Table 540: Telegram structure: sAN mLMLSetDisp 	264

Table 541: 	Example: sAN mLMLSetDisp 1 	265

Table 542: 	Telegram structure: sEN ECRChangeArr 	267

Table 543: 	Example: sEN ECRChangeArr 1 	267

Table 544: Telegram structure: sEA ECRChangeArr 	268

Table 545: 	Example: sEA ECRChangeArr 1 	268

Table 546: Telegram structure: sSN ECRChangeArr 	269

Table 547: 	Example: sSN ECRChangeArr 1 	269

Table 548: 	Telegram structure: sRN ECRxy 	270

Table 549: 	Example: sRN ECR01 	271

Table 550: 	Telegram structure: sEN ECRxy 	271

Table 551: 	Example: sEN ECR01 1 	271

Table 552: 	Telegram structure: sEA ECRxy 	272

Table 553: 	Example: sEA ECR01 1 	272

Table 554: 	Telegram structure: sRA/sSN ECRxy 	274

Table 555: 	Example: sRA ECR01 	275

Table 556: 	Telegram structure: sMN mLFEgetField 	276

Table 557: 	Example: sMN mLFEgetField 1 	277

Table 558: 	Example: sAN mLFEgetField 	279

Table 559: 	Example: sAN mLFEgetField 	279

Table 560: 	Telegram structure: sRN LFEperpdistresult 	281

Table 561: 	Example: sRN LFEperpdistresult 	281

Table 562: Telegram structure: sRA LFEperpdistresult 	283

Table 563: 	Example: sRA LFEperpdistresult 	283

Table 564: 	Telegram structure: sEN LFEperpdistresult 	285

Table 565: 	Example: sEN LFEperpdistresult 1 	285

Table 566: Telegram structure: sEA LFEperpdistresult 	285

Table 567: 	Example: sEA LFEperpdistresult 1 	286

Table 568: Telegram structure: sSN LFEperpdistresult 	288

Table 569: 	Example: sSN LFEperpdistresult 	288

Table 570: 	Telegram structure: sRN LFEinfringementinfo 	289

Table 571: 	Example: sEN LFEinfringementinfo 	289

Table 572: 	Example: sRA LFEinfringementinfo 	291

Table 573: 	Example: sRA LFEinfringementinfo 	291

Table 574: 	Telegram structure: sEN LFEinfringementinfo 	292

Table 575: 	Example: sEN LFEinfringementinfo 1 	292

Table 576: Telegram structure: sEA LFEinfringementinfo 	293

Table 577: 	Example: sEA LFEinfringementinfo 1 	293

Table 578: 	Example: sSN LFEinfringementinfo 	295

Table 579: 	Example: sSN LFEinfringementinfo 	295

Table 580: 	Telegram structure: sWN FieldSetSelectionMethod 	295

Table 581: 	Example: sWN FieldSetSelectionMethod 	295

Table 582: 	Telegram structure: sWA FieldSetSelectionMethod 	296

Table 583: 	Example: sWA FieldSetSelectionMethod 	296

Table 584: 	Telegram structure: sRN ActiveFieldSet 	296

Table 585: 	Example: sRN ActiveFieldSet 	297

Table 586: 	Telegram structure: sRA ActiveFieldSet 	297

Table 587: 	Example: sRA ActiveFieldSet 	297

Table 588: 	Telegram structure: sWN ActiveFieldSet 	298

Table 589: 	Example: sWN ActiveFieldSet 	298

Table 590: 	Telegram structure: sWA ActiveFieldSet 	298

Table 591: 	Example: sWA ActiveFieldSet 	298

Table 592: 	SOPAS error codes 	301

Table 593: 	Example: sFA ErrorCode Wrong userlevel 	301

Australia

Phone +61 (3) 9457 0600

1800 33 48 02 - tollfree E-Mail sales@sick.com.au

Austria

Phone +43 (0) 2236 62288-0

E-Mail office@sick.at

Belgium/Luxembourg

Phone +32 (0) 2 466 55 66

E-Mail info@sick.be

Brazil

Phone +55 11 3215-4900

E-Mail comercial@sick.com.br

Canada

Phone +1 905.771.1444

E-Mail cs.canada@sick.com

Czech Republic

Phone +420 234 719 500

E-Mail sick@sick.cz

Chile

Phone +56 (2) 2274 7430

E-Mail chile@sick.com

China

Phone +86 20 2882 3600

E-Mail info.china@sick.net.cn

Denmark

Phone +45 45 82 64 00

E-Mail sick@sick.dk

Finland

Phone +358-9-25 15 800

E-Mail sick@sick.fi

France

Phone +33 1 64 62 35 00

E-Mail info@sick.fr

Germany

Phone +49 (0) 2 11 53 010

E-Mail info@sick.de

Greece

Phone +30 210 6825100

E-Mail office@sick.com.gr

Hong Kong

Phone +852 2153 6300

E-Mail ghk@sick.com.hk

Hungary

Phone +36 1 371 2680

E-Mail ertekesites@sick.hu

India

Phone +91-22-6119 8900

E-Mail info@sick-india.com

Israel

Phone +972 97110 11

E-Mail info@sick-sensors.com

Italy

Phone +39 02 27 43 41

E-Mail info@sick.it

Japan

Phone +81 3 5309 2112

E-Mail support@sick.jp

Malaysia

Phone +603-8080 7425

E-Mail enquiry.my@sick.com

Mexico

Phone +52 (472) 748 9451

E-Mail mexico@sick.com

Netherlands

Phone +31 (0) 30 229 25 44

E-Mail info@sick.nl

New Zealand

Phone +64 9 415 0459

0800 222 278 - tollfree

E-Mail sales@sick.co.nz

Norway

Phone +47 67 81 50 00

E-Mail sick@sick.no

Poland

Phone +48 22 539 41 00

E-Mail info@sick.pl

Romania

Phone +40 356-17 11 20

E-Mail office@sick.ro

Russia

Phone +7 495 283 09 90

E-Mail info@sick.ru

Singapore

Phone +65 6744 3732

E-Mail sales.gsg@sick.com

Slovakia

Phone +421 482 901 201

E-Mail mail@sick-sk.sk

Slovenia

Phone +386 591 78849

E-Mail office@sick.si

South Africa

Phone +27 10 060 0550

E-Mail info@sickautomation.co.za

South Korea

Phone +82 2 786 6321/4

E-Mail infokorea@sick.com

Spain

Phone +34 93 480 31 00

E-Mail info@sick.es

Sweden

Phone +46 10 110 10 00

E-Mail info@sick.se

Switzerland

Phone +41 41 619 29 39

E-Mail contact@sick.ch

Taiwan

Phone +886-2-2375-6288

E-Mail sales@sick.com.tw

Thailand

Phone +66 2 645 0009

E-Mail marcom.th@sick.com

Turkey

Phone +90 (216) 528 50 00

E-Mail info@sick.com.tr

United Arab Emirates

Phone +971 (0) 4 88 65 878

E-Mail contact@sick. AE

United Kingdom

Phone +44 (0)17278 31121

E-Mail info@sick.co.uk

USA

Phone +1 800.325.7425

E-Mail info@sick.com

Vietnam

Phone +65 6744 3732

E-Mail sales.gsg@sick.com

8014631/1FHD/2022-04-27 A4 4c int49

Detailed addresses and further locations at www.sick.com

SICK AG | Waldkirch | Germany | www.sick.com